| Dept. | CHEMISTRY | Policy #: |  | 04-004 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Policy/ Procedure Title | CALIBRATING BECKMAN SYNCHRON DxC800 | Effective Date: |  | 5/12/16 |  |  |
|  |  | Revised Date: |  |  |  |  |
|  |  | Authorized Approval: |  | Cindy Schwartz, CLS, Area Lab Manager Jana Pindur, M.D., Laboratory Director |  |  |
| Personnel Covered: | All Chemistry Staff |  |  |  |  |  |
|  |  | Page: | 1 |  | of | 5 |

CALIBRATORS NEEDED:

The following calibrators are used in this facility:

| Calibrator | Tests | Storage Temp | Stability |
| --- | --- | --- | --- |
| Drug Calibrators 1 (Level 1 - 6) | CAR, PHE, PHY, THE, VPA | 2-8 °C | Same as original expiration date |
| Drug Calibrators 2 (Level 1 - 6) | ACTM, DIGN | 2-8 °C | Same as original expiration date |
| Drug Calibrators 3 (Level 1 - 6) | GENT | 2-8 °C | Same as original expiration date |
| Calibrator 5 (Level 1) | CRPH | 2-8 °C | Same as original expiration date |
| Alcohol (ETOH) Calibrator (Level 1-2) | ETOH | 2-8 °C | Same as original expiration date |
| Microprotein Calibrator (Level 1) | M-TP | -15 to -20 °C | 60 days at 2-8 °C |
| Multi-Calibrator (Level 1) | Lactate, Mg, Uric | -15 to -20 °C | 20 days at 2-8 °C |
| Bilirubin Calibrator (Level 1) | TBIL and DBIL | -15 to -20 °C | No open vial stability |
| Protein Calibrators (Level 1-2) | ALB and TP | -15 to -20 °C | 60 days at 2-8 °C |
| AQUA Calibrators (Level 1 - 3) | Ca, CL, CO2, Creat, GLUm, K, Na, BUN, Phos | 2-8 °C | 1 month at Room Temp |
| Enzyme Calibrators (Level 1-2) | LIP (IFCC method) | -15 to -20 °C | 60 days at -15 to -20 °C |
| Vancomycin Calibrators (Level 1-6) | VANC | 2-8 °C | Same as original expiration date |
| Ammonia Calibrator (Level 1-2) | AMM | 2-8 °C | 60 days at 2-8 °C |
| Salicylate Calibrator (Level 1) | SALY | 2-8 °C | Same as original expiration date |

* Note: All Calibrator must be capped in original containers after used

Policy

The analyzer must have a valid calibration in memory before controls and patients can be processed.

CALIBRATION FREQUENCY:

Calibration is required:

- As indicated on the Rgts/Cal screen for tests.

- Loading a new reagent that require calibration as indicated by the analyzer.

- Resolving problems identified by out-of-range controls.

- When specific parts have been replaced, or maintenance has been performed.

LOADING CALIBRATOR DISKETTE:

Follow the steps below to load any calibrator diskette needed. Calibrators with diskettes are the following: Micro-protein, Mutli-Cal, Bilirubin Cal, Protein Cal and Enzyme Cal)

| Step | Action |
| --- | --- |
| 1. | Check calibrator disk with calibrator bottle to confirm lot number information is the same. |
| 2. | Select Rgts/Cal from the menu bar. |
| 3. | Select F6 Cal Options. |
| 4. | From Cal options dialog box, select <1> load Calibrator Diskette |
| 5. | Insert diskette in the drive. From the Load Calibrator Diskette dialog box, select OK. |
| 6. | Verify that the correct diskette was read from the response on the screen, indicating calibrator lot number loaded. |
| 7. | If no other disks to be loaded, select <Cancel> |
| 8. | Remove calibrator diskette from the disk drive |

ASSIGNING CALIBRATOR BARCODES:

Follow the steps below to assign barcodes to any calibrator.

| Step | Action |
| --- | --- |
| 1. | Select Rgts/Cal from the menu bar. |
| 2. | Verify that the reagent calibrator has an assigned bar code ID. Assign a calibrator ID and/or rack and position if the reagent has not been assigned. |
| 3 | Select F7 Assign. |
| 4. | Select <Next> or <Prev> to locate the calibrator of interest OR Select the Calibrator Name options button <> at the top of the Assign Bar Code/Rack dialog box to view the List of Calibrators. |
| 5. | Select <Avail Racks> to view available racks for programming or to view racks with existing calibrator assignments. The rack numbers displayed contain at least one empty position available for assignment. |

| 6. | Assign calibrator identification by entering calibrator ID information. NOTE: - If the rack contains calibrator ID assignments, they are displayed in the appropriate cup positions. - Cup positions which are empty or contain unreadable bar codes appear as blank input fields. - Calibrator assignments can be cleared using <Clear Samples> in the Available Racks dialog box. |
| --- | --- |
| 7. | Select <Close> to exit and save information. |

PERFORMING CALIBRATION:

Follow the steps below to calibrate the analyzer.

| Step | Action |  |  |  |
| --- | --- | --- | --- | --- |
| 1. | Verify calibrator information has been loaded into the computer for the lot number of calibrator being used. |  |  |  |
| 2. | Select Rgts/Cal from the menu bar. |  |  |  |
| 3. | Select chemistries to be calibrated. (Press F5 to see the list of calibrator required) |  |  |  |
| 4. | Obtain the rack holding the bar coded tubes for the calibration required. |  |  |  |
| 5. | Select F7 Assign. - Scroll through the list of calibrators until the one desired is found. - Make sure for each calibrator level the barcode ID is listed exactly as it appears on the actual barcode on the test tube in the rack. |  |  |  |
|  |  | If the ID listed is … | then you would … |  |
|  |  | the same | Select Close. |  |
|  |  | Different | Scan in the barcode on the tube for each level of calibrator, then select Close. |  |
| 6. | Select F4 Cal after all desired chemistries have been selected for calibration. |  |  |  |
| 7. | Place the rack with the filled calibrator tubes/cups (barcode facing out) on the analyzer. If a calibration requires 2 racks, they must be loaded sequentially. NOTE: Be sure there are no bubbles in the calibrator tubes/cups. |  |  |  |
| 8. | Press Run on the instrument to start the calibration when ready. |  |  |  |

NOTE: After a successful calibration, Quality Control material must be run to verify the analyzer is in control.

CANCELING A CALIBRATION REQUEST: Follow the steps below to cancel a calibration request.

| Step | Action |
| --- | --- |
| 1. | Select Rgts/Cal from the menu bar. |
| 2. | Select the appropriate requested and assigned chemistries to be canceled. |
| 3. | Select F8 No Cal to cancel calibration request. |

NOTE: - Chemistries currently selected and programmed for calibration may be canceled only if the calibration is not yet in progress. - Calibration reports for a given calibrator do not print until all calibrations are complete for each chemistry requested, unlike sample and control results, which print on a cup-by-cup basis.

CORRECTIVE ACTION:

If calibration fails refer to Synchron DxC Clinical Systems Diagnostics and Troubleshooting manual.

To reprint Calibration report:

| Step | Action |
| --- | --- |
| 1. | Select RGT/Cal from the menu bar |
| 2. | Select the chem. position(s) to print |
| 3. | Select options (F6) |
| 4. | From the Cal options dialog box select or enter 10 in the “Option Number field” Options: Print current Cal report Print pervious Cal report |

Refer to the instruction of use and package insert for the calibrator’s storage and stability information.

References

Unicel DxC Synchron Clinical Systems onboard IFU,

www.beckmancoulter.com

Document History Page

Effective Date:

| Change type: New, Major, Minor etc. | Changes Made to SOP - describe | Signature responsible person/date | Med. Director Authorized Reviewed/Date | Operational Director Authorized Reviewed/Date | Date change Implemented |
| --- | --- | --- | --- | --- | --- |
| Major | New | Cindy Schwartz 5/12/16 |  |  |  |