| Dept. | CHEMISTRY | Policy #: |  | 04-102 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Policy/ Procedure Title | ACCESS 2 STARTUP, SHUTDOWN AND REBOOT | Effective Date: |  | 05/24/16 |  |  |
|  |  | Revised Date: |  |  |  |  |
|  |  | Authorized Approval: |  | Cindy Schwartz, CLS Area Laboratory Manager Jana Pindur, M.D. Laboratory Director |  |  |
| Personnel Covered: | All Chemistry Staff |  |  |  |  |  |
|  |  | Page: | 1 |  | of | 4 |

Purpose

To understand the proper procedures for power down, start up and reboot of the Beckman access system.  This is necessary for the following conditions: unexpected power failure, computer down time and troubleshooting.

Procedure

## Turning the Power On

- If the PC was shut down, restart it.

- Locate the power button on the lower, right side near the back of the instrument.  Press the upper part of the switch to turn the power on (l position).

- Wait until the system is in the Ready mode and no message appears in the blue system mode area before you continue the procedure.

If initialization fails, review the Event Log and troubleshoot according to any error event with a similar date and time to the attempted initialization.

Reboot the Instrument

Use this procedure to reboot the Access 2 instrument.  There are two different ways to reboot the instrument:

- You can reboot using the reset button.  This is also called a warm boot.

- You can reboot using the power switch.  This is also called a cold boot.

NOTES

- Do not select any buttons or press any keys until rebooting is complete.  After you initiate a reboot, there is a pause of approximately 2 minutes as the software resets.  Then the system enters the Not Ready mode and system initialization begins.

- During system initialization, the system homes mechanical devices and displays a flashing message in the system mode area.  When most system devices complete initialization, the system changes to the Ready mode.

- The system continues to initialize the remaining devices and displays a flashing message in the system mode area.  When the message disappears, system initialization is complete.

- If the system does not successfully initialize, contact Technical Support.

Reboot Using the Reset Button

- Open the front panel of the instrument.

- Locate the reset button, just to the right of the pipettor gantry.

- Push and hold in the reset button for one second, then release.

- Close the front panel of the instrument immediately.

NOTE

If the front panel remains open too long, the system cannot home the mechanical devices.  If this occurs, press the reset button a second time, then close the front panel immediately.

5.	Wait until the system is in the Ready mode and no message appears in the blue system mode area before you continue normal operation.

Rebooting Using the Power Switch

- Be sure the front panel of the instrument is closed.

- Locate the power switch on the lower, right side near the back of the instrument.  Press the lower part of the switch to turn the power off (O position).

- Wait about 20 seconds, and then press the top part of the switch to turn the power on (l position).

4.   Wait until the system is in the Ready mode and no message appears in the blue system mode area before you continue normal operation.

Shutting Down the PC

- Go to the PC Admin screen.  To get to this screen from the Main Menu, select Configure F8 to display the

Configure menu, then select PC Admin F7.

- Select Shut Down PC F8.  A confirmation message is displayed.

- Select Yes F1 to shut down the PC.

If you are preparing to shut down the server, repeat Steps 1 – 3 for each of the client PCs first.  Then repeat: Steps 1 – 3 for the server.

- Wait about 20 seconds, and then perform the restarting procedure.

References

## Beckman Coulter Access 2.  Quick Reference Guide Sept. 2014

Document History Page

Effective Date:

| Change type: New, Major, Minor etc. | Changes Made to SOP - describe | Signature responsible person/date | Med. Director Authorized Reviewed/Date | Lab Manager Authorized Reviewed/Date | Date change Implemented |
| --- | --- | --- | --- | --- | --- |
| Major | Change in procedure format | Cindy Schwartz 5/25/16 |  |  |  |