| Dept. | CHEMISTRY | Policy #: |  | 04-105 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Policy/ Procedure Title | PROCESSING SAMPLES ON BECKMAN ACCESS 2 | Effective Date: |  | 05/25/16 |  |  |
|  |  | Revised Date: |  |  |  |  |
|  |  | Authorized Approval: |  | Cindy Schwartz, CLS Area Laboratory Manager Jana Pindur, M.D. Laboratory Director |  |  |
| Personnel Covered: | All Chemistry Staff |  |  |  |  |  |
|  |  | Page: | 1 |  | of | 3 |

Policy

Quality control must be processed and acceptable every 24 hours before patients can be resulted.

PREPARING SAMPLES:

Follow these steps to prepare samples for analysis on the Beckman Access 2.

| Step | Action |
| --- | --- |
| 1. | - Check specimen for the presence of fibrin clots or red cells in the plasma. Re-spin if necessary. |
| 2. | - Check specimen for sufficient volume. |
| 3. | Remove stoppers from the tubes just before loading into the sample sectors. |
| 4. | Load specimens into sample sectors appropriate for the sample tube size. |
| 5. | If the specimen is short, place an aliquot of the specimen in a Beckman Micro-tube. Place a barcode on the cup and place in the sample sector. |

PROCESSING BAR CODED SAMPLES:

Follow the steps to process samples which will be run in the bar coded bi-directional mode.

| Step | Action |
| --- | --- |
| 1. | Load tubes in racks verifying that the barcode labels are facing outward on the same side as the rack barcode. |
| 2. | Go to the Sample Manager screen. To get to this screen from the Main Menu, select Sample Manager F1. |
| 3. | Select Load a Rack F1. |
| 4. | Load a rack. |
| 5. | Select Done F1. |
| 6. | Close the sample carousel door. |
| 7. | To start sample processing, select Run. |

PROGRAMMING SAMPLES MANUALLY:

Follow the steps to program samples manually.

| Step | Action |
| --- | --- |
| 1. | Go to the Sample Manager screen. To get to this screen from the Main Menu, select Sample Manager F1. |
| 2. | Enter the rack ID in the Rack ID field and press [Enter]. |
| 3. | Select Test Request F3 to display the Test Request screen. |
| 4. | In the Sample ID field, enter the sample ID and press [Enter]. |
| 5. | Using the Test menu, select the tests you want to run on the sample. If the Test menu is not shown, select Show Test Menu F3 to display the menu. |
| 6. | Review the values for the Dilution, Sample Type, STAT, and Comment fields and edit them if necessary. |
| 7. | Place the samples in the rack. |
| 8. | Select Load Rack F1. Load the rack. |
| 9. | To start sample processing, select Run. |

CLEARING RACKS

Follow these steps to clear the racks with previous assigned samples.

| Step | Action |
| --- | --- |
| 1. | Go to the Sample Manager screen. To get to this screen from the Main Menu, select Sample Manager F1. |
| 2. | In the Off Board list, select the rack you want to clear. |
| 3. | Select Clear Selected Rack F7. The system deletes the rack from the Off Board list and clears all sample information for the rack from the Test Request or Maintenance Requests screen. |

RESULTING SPECIMENS:

Follow directions in Resulting Specimens in the Laboratory Information System.

References

Beckman Access 2 Quick Reference Guide Sept. 2014.

Document History Page

Effective Date:

| Change type: New, Major, Minor etc. | Changes Made to SOP - describe | Signature responsible person/date | Med. Director Authorized Reviewed/Date | Lab Manager Authorized Reviewed/Date | Date change Implemented |
| --- | --- | --- | --- | --- | --- |
| Major | Change in procedure format | Cindy Schwartz 5/25/16 |  |  |  |