| Dept. | CHEMISTRY | Policy #: |  | 04-106 |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Policy/ Procedure Title | QC BECKMAN ACCESS 2 | Effective Date: |  | 05/26/16 |  |  |
|  |  | Revised Date: |  |  |  |  |
|  |  | Authorized Approval: |  | Cindy Schwartz, CLS, Area Lab Manager Jana Pindur, M.D., Laboratory Director |  |  |
| Personnel Covered: | All Chemistry Staff |  |  |  |  |  |
|  |  | Page: | 1 |  | of | 3 |

Policy

Quality control must be processed and acceptable on the day patients samples are ran. In addition, QC should be performed:

- After a new reagent lot is loaded

- After calibration

- After major maintenance

Programming QC:

Follow the steps to program samples manually.

| Step | Action |
| --- | --- |
| 1. | Enter Rack ID from the Sample Manager. |
| 2. | Select F3 Test Request. |
| 3. | Select F5 Request QC. |
| 4. | Choose the QC to be run. |
| 5. | Select F3 Show Test Menu. |
| 6. | Select test(s) to run QC. |
| 7. | Place QC sample in the rack |
| 8. | Press F1 to load the rack on the analyzer and press run |

| Access 2 | [Cardiac Marker Plus] | [Immunoassay Plus] |
| --- | --- | --- |
| Sample Type | Serum | Serum |
| Storage Unopened Open | - 20 ˚C to - 70 ˚C 2 – 8 ˚C | - 20 ˚C to - 70 ˚C 2 – 8 ˚C |
| Stability (Thawed and Opened) | 10 days | 14 days |
| Level | 1 and 2 | 1 and 3 |
| Analytes | BNP, Troponin I | HCG, Estradiol |

| Access 2 | [Specialty Immunoassay] |  |
| --- | --- | --- |
| Sample Type | Serum |  |
| Storage Unopened Open | - 20 ˚C to - 70 ˚C 2 – 8 ˚C |  |
| Stability (Thawed and Opened) | 7 days |  |
| Level | 1 and 2 |  |
| Analytes | PTH I/O |  |
| All QC materials are good until expiration day if stored unopened at the recommended temperature. |  |  |

Document History Page

Effective Date:

| Change type: New, Major, Minor etc. | Changes Made to SOP - describe | Signature responsible person/date | Med. Director Authorized Reviewed/Date | Lab Manager Authorized Reviewed/Date | Date change Implemented |
| --- | --- | --- | --- | --- | --- |
| Major | Change in procedure format | Cindy Schwartz 5/25/16 |  |  |  |