| Attributes for the Access 2 Analyzer Effective Date: 5/15/16 |  |  |
| --- | --- | --- |
| Document #/version 04-107 |  |  |
| Access 2 | [BNP] | [Estradiol] |
| Methodology/Reaction | Two-site immunoenzymatic (“sandwich”)/ chemiluminescent | Two-site immunoenzymatic (“sandwich”)/ chemiluminescent |
| Clinical Utility | Diagnosis and assessment of severity of congestive heart failure | Monitoring ovulatory status, assessment of sexual development, causes of infertility and menopause |
| Sample Type | EDTA Plasma | Serum |
| Minimum Sample Volume | 55µL | 35µL |
| Prefer Sample Volume | 0.3 mL | 0.3 mL |
| Specimen Stability |  |  |
| Room Temp | 7 hrs | 8 hrs |
| Refrigerated | 24 hrs | 48 hrs |
| Frozen | Indefinite | indefinite |
| Patient/Sample Preparation | None | None |
| Reagents | No preparation required | No preparation required |
| Reagent Storage | 2 ˚C to 10 ˚C | 2 ˚C to 10 ˚C |
| Calculations | N/A | N/A |
| Interferences** | Hemolysis – NSI Icterus - NSI Lipemia – NSI * | Hemolysis – NSI Icterus - NSI Lipemia – NSI * |
| Method Performance Specifications | Linear range of detection: 1 – 5000 pg/mL | Linear range of detection: 20 – 4800 pg/mL |
| Reporting - Reference Range | < 100 pg/mL | Males; 20-75 pg/mL. Post-Menopausal Females; 20-88 pg/mL. Mid-follicular phase; 24-114 pg/mL. Periovulatory; 62-534 pg/mL. Mid-luteal phase; 80-273 pg/mL. |
| - Auto-Verification | Result; Action. <5.0; Manually verify after repeat. Numerical results >=5.0; Auto-verify. Result with ">"; Manually verify as >5000 (Enter 5001 in LIS to convert). | N/A |
| Reporting | N/A | N/A |
| -Alert Limits |  |  |
| Dilutions | Do not dilute, report as > 5000 pg/mL | Do not dilute, report as >4800pg/mL |
| NSI – No Significant Interference |  |  |
| Refer to the Beckman chemistry information sheet, operator tips and operating manual for more detailed information. *All samples with lipemia index >3 should be ultracentrifuged by airfuge. ** All immunoassay assays are affected by the present of heterophile antibodies. |  |  |

| Attributes for the Access 2 Analyzer Effective Date: 5/15/16 |  |  |
| --- | --- | --- |
| Document #/version 04-107 |  |  |
| Access 2 | [HCG] | [PTH-I/O] |
| Methodology/Reaction | Two-site immunoenzymatic (“sandwich”)/ chemiluminescent | Two-site immunoenzymatic (“sandwich”)/ chemiluminescent |
| Clinical Utility | Determining pregnancy, ectopic pregnancy or impending spontaneous abortion | Used intraoperatively during parathyroidectomy. The laboratory only performs intraoperative PTH. |
| Sample Type | Lithium Heparin Plasma Serum | EDTA Plasma |
| Minimum Sample Volume | 56µL | 55µL |
| Prefer Sample Volume | 0.3 mL | 0.3 mL |
| Specimen Stability |  |  |
| Room Temp | 8 hrs | 8 hrs |
| Refrigerated | 48 hrs | 48 hrs |
| Frozen | 6 months | Indefinite |
| Patient/Sample Preparation | None Post-menopausal patients can show slight elevation up to 12 mIU/mL | In parathyroidectomy, a PTH decrease > 50% after excision is a good indication that all hyperfunctioning parathyroid tissue was resected at the time of surgery |
| Reagents | No preparation required | No preparation required |
| Reagent Storage | 2 ˚C to 10 ˚C | 2 ˚C to 10 ˚C |
| Calculations | N/A | N/A |
| Interferences** | Hemolysis – NSI Icterus - NSI Lipemia – NSI * | Hemolysis – NSI Icterus - NSI Lipemia – NSI * |
| Method Performance Specifications | Linear range of detection: 0.6 – 1350 mIU/mL Auto-dilution (1:200) 1050–270,000 mIU/mL | Linear range of detection: 6 - 3500 pg/mL |
| Reporting - Reference Range | < 5 mIU/mL | 12 – 88 pg/mL |
| - Auto-Verification | Result; Action. All numerical results; Auto-verify. Result with ">"; Manually verify as >270,000 (Enter 270,001 in LIS to convert). | N/A |
| Reporting | N/A | All intraoperative PTH are STAT and the post excision result must be call to surgery. |
| -Alert Limits |  |  |
| Dilutions | Do not dilute beyond the auto-dilution by the analyzer, report as > 270,000 mIU/mL | Access Sample Diluent A – up to endpoint upon physician’s request |
| NSI – No Significant Interference |  |  |
| Refer to the Beckman chemistry information sheet, operator tips and operating manual for more detailed information. *All samples with lipemia index >3 should be ultracentrifuged by airfuge. ** All immunoassay assays are affected by the present of heterophile antibodies. |  |  |

| Attributes for the Access 2 Analyzer Effective Date: 5/15/16 |  |
| --- | --- |
| Document #/version 04-107 |  |
| Access 2 | [Troponin] |
| Methodology/Reaction | Two-site immunoenzymatic (“sandwich”)/chemiluminescent |
| Clinical Utility | Aids in the risk stratification of patients with unstable angina or non-ST segment elevation acute coronary syndromes |
| Sample Type | Lithium Heparin Plasma |
| Minimum Sample Volume | 50µL |
| Prefer Sample Volume | 0.3 mL |
| Specimen Stability |  |
| Room Temp | 2 hrs |
| Refrigerated | 24 hrs |
| Frozen | 6 months |
| Patient/Sample Preparation | None |
| Reagents | No preparation required |
| Reagent Storage | 2 ˚C to 10 ˚C |
| Calculations | N/A |
| Interferences** | Hemolysis – NSI Icterus - NSI Lipemia – NSI * |
| Method Performance Specifications | Linear range of detection: 0.02 – 100.0 ng/mL |
| Reporting - Reference Range | < 0.04 ng/mL |
| - Auto-Verification | Result; Action. 0.0-0.50; Auto-verify. >0.50; Manually verify follow critical value policy in documenting critical readback comment. Result with ">"; Manually verify as >100 and follow critical value policy in documenting critical readback comment (Enter 101 in LIS to convert). |
| Reporting | >0.5 ng/mL |
| -Alert Limits |  |
| Dilutions | Do not dilute, report as > 100 ng/mL |
| NSI – No Significant Interference |  |
| Refer to the Beckman chemistry information sheet, operator tips and operating manual for more detailed information. *All samples with lipemia index >3 should be ultracentrifuged by airfuge. ** All immunoassay assays are affected by the present of heterophile antibodies. |  |

Document History Page

Effective Date:

| Change type: New, Major, Minor etc. | Changes Made to SOP - describe | Signature responsible person/date | Med. Director Authorized Reviewed/Date | Lab Operational Director Authorized Reviewed/Date | Date change Implemented |
| --- | --- | --- | --- | --- | --- |
| Major | New procedure to define Access attributes for each tested analyte | Cindy Schwartz 5/15/16 |  |  |  |