#### COAGULATION SPECIMEN REQUIREMENTS

| ##### SPECIMEN TYPE | The only acceptable specimen is whole blood drawn in Blue top (plus plastic 3.2% buffered sodium citrate) vacutainer tubes. The tubes must be to full to maintain the proper anticoagulant to blood ratio. A minimum 90% fill is recommended |
| --- | --- |

| ##### SAFETY | Laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety, the safety of others and adhering to all departmental and medical center safety policies and procedures. Refer to Policy and Procedures Safety Manual Infection Control 11-085-01. |
| --- | --- |

| ##### TEST STABILITY | Protime If unopened specimens maybe left at room temperature, unspun for 24 hours. Testing must be completed within 24 hours after the draw time. aPTT - Must be tested within 4 hours of collection. - If testing cannot occur within 4 hours, then plasma must be separated into a plastic tube and placed in the refrigerator up to 4 hours prior to testing. D-dimer - Remove plasma if not run immediately. - Plasma storage: 8 hours at Room Temperature (20 ± 5ºC). NOTE: Samples for PT, aPTT can be stored frozen if sample has been double spunned. Double spun citrated plasma can be stored for 2 weeks at -20 °C or 6 months at -70 °C. |
| --- | --- |

#### COAGULATION SPECIMEN REQUIREMENTS

| ##### CENTRIFUGATION | Blue top tube for coagulation testing should be centrifuged in a refrigerated centrifuge if possible. Centrifuge for 5 minutes at an RPM high enough to create a platelet poor plasma (platelet count less than 10,000/uL). If freezing is needed, plasma must be double spunned as described below: - Separate the plasma in a plastic tube [being careful not to go near the bottom of the plasma and disturb the buffy coat] - Spin again for 5 minutes. - Separate the plasma again in a plastic tube [being careful not to go near the bottom of the plasma and disturb the buffy coat] - Record on the label that the specimen has been double spun. |
| --- | --- |

| ##### SPECIAL COAG STUDIES | Special Coagulation studies may require that the blood sample be placed on ice at the time it is drawn. See Kaiser Permanente Reference Manual For Laboratory Tests for specific requirements. |
| --- | --- |

#### COAGULATION SPECIMEN REQUIREMENTS

| ##### ANTICOAGULANT ADJUSTMENT FOR HEMATOCRIT SPECIMENS | This adjustment is for a 3 ml tube. The anticoagulant (sodium citrate) in blue top tube must be adjusted when Hematocrit is greater than 55%. The formula is as follows: Y (100 - HCT) X =__________________ HCT X = amount of anticoagulant needed Y = 0.3 (3 ml tube) Pipette out and discard the amount of anticoagulant from the 3 ml that you do not need and draw patient blood to the top of the tube. |
| --- | --- |

| ##### SPECIMEN REJECTION CRITERIA | The following conditions are not acceptable for coagulation specimens: - Unlabeled or improperly labeled. - Hemolyzed. - Inadequately filled blue top tube - Frozen whole blood. - Insufficient volume (QNS) - Wrong anticoagulant. - Clotted. - Improperly processed. |
| --- | --- |

#### COAGULATION SPECIMEN REQUIREMENTS

| ##### VALIDATION OF SPECIMEN PROCESSING FOR PLATELET-POOR PLASMA | NOTE: Platelet-poor plasma is particularly important when testing for the presence of a lupus anticoagulant, when measuring the level of unfractionated heparin, and in plasma samples to be frozen for later testing. Platelet-poor plasma should have a residual platelet concentration of less than 10 X 109/L. aPTT, prothrombin time/international normalized ratio (PT/NR) performed on fresh plasma samples are not affected by platelet counts of at least up to 200 x 109/L (200,000/μL).” Annually the lab will evaluate all centrifuges used for processing coagulation testing specimens. - Five specimens collected in sodium citrate tubes will be spun in each centrifuge to be evaluated for 5 minutes. - The residual platelet count of the plasma will be measured using the hematology analyzer. - If the average platelet count of the 5 specimens is less than 10x10 9 /L, specimens requiring platelet-poor plasma may be spun for 5 minutes in the centrifuge. - If the average platelet count exceeds 10x10 9 /L, repeat the test with 5 new specimens and centrifuge for 10 minutes. - See form for recording and evaluating data. |
| --- | --- |

#### COAGULATION SPECIMEN REQUIREMENTS

KAISER MEDICAL CARE PROGRAM

ORANGE COUNTY AREA

MEASUREMENT OF PLATELET - POOR PLASMA LOGSHEET

Centrifuge IR # ________________

Date Tested __________________

Platelet-poor plasma should have a residual platelet concentration of less than 10 X 109/L.

| Sample # | Platelet Count 5 minute Centrifugation | Acceptable/ Not Accepatable | Platelet Count 10 minute centrifugation | Acceptable/ Not Accepatable |
| --- | --- | --- | --- | --- |
| Average Platelet Count |  |  |  |  |

□PASSED:

(Specimens requiring platelet-poor plasma may be spun for 5 minutes in this centrifuge)

□NOT PASSED:

(Specimens requiring platelet-poor plasma must be spun for 10 minutes in this centrifuge)

CLS_________________________________________		Date_____________________

Document History Page

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Name of responsible person/date | Med. Dir. Reviewed/ date | Lab Manager reviewed/ date | Date change Imp. |
| --- | --- | --- | --- | --- | --- |
| New |  | Sam Ibe 01/21/2014 |  |  |  |
| Minor | - Template Change - Updated Index Number - Added 90% minimum fill requirement in specimen type - Added the following rejection criteria: - Inadequately filled blue top tube - Frozen whole blood. - Insufficient volume (QNS) | Yvette Lingat 10/10/19 |  |  | 11/20/19 |