#### REFERENCE, THERAPEUTIC & REPORTABLE RANGES

| ##### POLICY | Each analyzer has a reportable range set by the instrument manufacturer and/or established by the Clinical Laboratory. The reference ranges and therapeutic ranges are established by the Laboratory system. |
| --- | --- |

| ##### REFERENCE RANGES | The following are the current Reference Ranges for the Coagulation Department: Analyte; Range. APTT; 25-37 seconds. D-Dimer to R/O DVT & DIC; <0.50 FEU mcg/mL. D-Dimer to R/O PE. ≤50 years old:; <0.50 FEU mcg/mL. 51-79 years old:; <(age x 100) FEU mcg/mL. ≥80 years old:; <0.80 FEU mcg/mL. Protime; 12.0 - 15.4 Seconds (INR: 0.8-1.2). |
| --- | --- |

| ##### REPORTABLE RANGES | The following table lists the reportable ranges for the STAGO COMPACT MAX. For results outside the reportable range, see P&P COAG.MOB 03-0010 or COAG.MOB 03-0020*. Analyte; Reportable Range. Protime; 0.5-10.0 INR (8-150 seconds). APTT; 15 - 200 seconds. D-Dimer DVT/PE; 0.27- 4.00 FEU mcg/mL. D-Dimer DIC; 0.27- 20.00 FEU mcg/mL. *For values outside linearity limits (reportable range) a CLS must investigate to determine appropriate action in reporting. |
| --- | --- |

#### REFERENCE, THERAPEUTIC & REPORTABLE RANGES

Document History Page

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Name of responsible person/date | Med. Dir. Reviewed/ date | Lab Manager reviewed/ date | Date change Imp. |
| --- | --- | --- | --- | --- | --- |
| New |  | Sam Ibe 01/21/2014 |  |  |  |
| Major | Updated D-Dimer unit from ng/mL to mcg/mL | Yvette Lingat 3/6/19 | 3/6/19 | 3/6/19 | 3/26/19 |
| Major | - Template Change - Update Index Number - Added Reference Range for D-Dimer to R/O PE - Added Therapeutic range for D. Dimer DIC | Yvette Lingat 10/17/19 |  |  | 11/20/19 |

Imp. =Implemented