#### Quality Control (QC) & Quality Control Out of Range Action

| ##### POLICY | Before patient results can be reported the quality control must be run and be within acceptable limits. Controls must be run in the same manner as the patient samples. |
| --- | --- |

| ##### SAFETY | Laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety, the safety of others and adhering to all departmental and medical center safety policies and procedures. Refer to Policy and Procedures Safety Manual Infection Control SAF.01.0.60. |
| --- | --- |

| ##### FREQUENCY | Controls must be run every 8 hours at least once per shift. When Quality Control tolerance limits are exceeded, corrective actions must be performed and recorded before analyzing patient samples. Controls are also run: - When new reagents are placed on the coagulation analyzer, which include the following: - Neoplastin C Plus - CaCl - PTT Reagent - LIA Test Latex - LIA Test Buffer - Heparin Reagent - Heparin Buffer - When repairs have been completed on the analyzer. |
| --- | --- |

Continued on next page

#### Quality Control (QC) & Quality Control Out of Range Action, Continued

| ##### RULES & ACTION | Two levels of control must be run at the beginning of each shift, when there is reagent change or corrective maintenance performed. BEFORE ANY PATIENT RESULTS ARE REPORTED, control values must be within established limits. The following rule is used to determine if your control is OUT: - One value outside of posted limits. - If one control is outside posted limits, do one or more of the following, recording each step until controls are within limits: - Rerun the control that is out. - Reconstitute new controls using freshly poured reagent grade water. Pay special attention to pipetting and mixing and time of reconstitution and run with current reagents. - Reconstitute new reagents using freshly poured reagent grade water. Pay special attention to pipetting and mixing and time of reconstitution and run controls with new reagents. - Check temperature of system. - Perform preventive or corrective action on instrument. - If control is now in range or limits, indicate action taken on audit sheet. After thorough investigation, if controls are still “out of control", notify supervisor. |
| --- | --- |

Continued on next page

#### Quality Control (QC) & Quality Control Out of Range Action, Continued

| ##### MULTIPLE TEST OUT | Multiple tests are out IF; THEN. PT & PTT high; - Could be block in Needle 1-check needle; check syringe. - Could be block in Needle 3-check needle; check syringe. - Could be bubbles in the line; watch tubing above syringe and tubing in the area of the cleaner solution.. PT & PTT low; - Check and Bleach wash wells; especially 1 and 3. - Watch rinsing of needles; especially 1 and 3. - Check that Waste and Cleaner are in correct position. - May need to do a Back-Bleach of the system. - Residual bleach will affect results-run QC prior to patient testing. - Combination lowers FIB but raises PTT.. |
| --- | --- |

| ##### ERRATIC RESULTS ON ANY/ALL ASSAYS | Multiple tests are out - Watch processing of samples at all stages. - Watch for balls jumping out of cuvettes - Check wells for balls. Use a strong magnet. Use dental mirror to look in wells. - Watch measuring plate and see if the balls are moving back and forth in a consistent manner. - If they are not, check area around cuvette wheel. Strong magnets near wheel may magnetize cuvettes. - Put on a new cuvette wheel and watch measuring plate again. - Watch processing of samples at all stages. - Change Syringe tip and O ring for Needle 1. - Needle 1 tip appearance. See if the Teflon coating is intact/chipped. |
| --- | --- |

Continued on next page

#### Quality Control (QC) & Quality Control Out of Range Action, Continued

| ##### Daily QC Documentations | Step; Action. - 1.; Record Neoplastine CI+ lot number and expiration date on STAGO Quality Control and Action Log in space provided on each shift as needed.. Enter QC results daily in Clarity by STAGO Quality Assessment Program website. Steps; Action. 1; Login to the Clarity by STAGO website. 2; Select instrument model no.. 3; Select instrument ID.. 4; Select the month. 5; Select the correct QC product. 6; Select the correct QC Lot no.. 7; Click on “Raw Data” link under data entry. 8; Click on “Edit” link at the bottom of the window. 9; Select Analyte. 10; Enter QC value for the analyte selected. 11; Click “Next” to go to the next QC lot. 12; Click Submit at the bottom. 13; Select next analyte. 14; Repeat steps 10-12.. |
| --- | --- |

Continued on next page

#### Quality Control (QC) & Quality Control Out of Range Action, Continued

| ##### Monthly QC | Step; Action. - 1.; At the end of the month, printout a copy of the Levy-Jennings chart from STAGO COAG EXPERT. Review and file in Coag Quality Control book. To print Levy-Jennings in Stago Coag Expert: Steps; Action. 1; Go to Main Menu. 2; On Quality Control box, select “iQC”. 3; Select “By Test”. 4; Select test in the drop-down menu. 5; Filter date range on the right side of the screen. 6; Click printer icon in the right of the display to print L-J chart. 7; Click to go to the next analyte.. Clarity by STAGO Quality Control Summary Report are printed at the end of each month and reviewed by the Hematology/Coagulation Supervisor. Printing the Clarity Monthly Report Step; Action. 1; Login to the Clarity by STAGO website. 2; Select instrument model no.. 3; Select instrument ID.. 4; Select “Monthly Summary” under Reporting. 5; Print the Summary Report. 6; Select “Histogram” and “Youden” under Reporting. 7; Print both the Histogram and Youden report. The data is used for evaluation and comparison with other STAGO analyzers. The report provides interlaboratory comparison indicating precision and accuracy relative to peer data. Issues identified are noted and corrective action is taken as needed and documented. This review includes inspection of trends and small shifts that may indicate systematic or potential accuracy problems. Monthly SD and CV’s are examined, and corrective action is implemented and documented as needed.. |
| --- | --- |

#### Quality Control (QC) & Quality Control Out of Range Action,

Document History Page

| Change Type: New, Major, Minor, etc. | Changes Made to SOP-describe | Name of Responsible Person/date | Med.Dir. Reviewed/date | Lab Manager Reviewed/date | Date change Imp. |
| --- | --- | --- | --- | --- | --- |
| New |  | Sam Ibe 01/21/2014 |  |  |  |
| Major | - Updated format - Updated index number - Updated Safety Manual Index number - Added daily and monthly QC documentation instructions for clarity and Stago Coag Expert | Yvette Lingat 10/22/19 |  |  | 11/20/19 |

Imp. =Implemented