#### Assaying New Reagents, Controls, Population Means & Reference Ranges

| ##### POLICY | One month prior to the expiration of the current lot number of the controls and reagents, parallel testing and range establishment will occur. |
| --- | --- |

| ##### SAFETY | Laboratory employees are expected to maintain a safe working environment and an injury-free workplace. Laboratory employees are responsible for their own safety, the safety of others and adhering to all departmental and medical center safety policies and procedures. Refer to Policy and Procedures Safety Manual Infection Control SAF.01.0.60. |
| --- | --- |

| ##### ESTABLISHING CONTROL RANGES | To establish new ranges for controls when control lot or reagent lot changes do the following: Step; Action. 1; For each methodology (Protime, APTT, D-dimer), run a minimum of 20 values on different days.. 2; Calculate the Mean, SD, and CV.. 3; New ranges will be based on - Mean + 2 SD for D-dimer - Mean and CV for PT and APTT. |
| --- | --- |

| ##### REFERENCE RANGE | Run a minimum of 20 values on normal subjects to establish reference range for Protime and APTT by calculating the Mean, SD, and CV. The specimens used should not be patients receiving any anticoagulation therapy and/or have liver function tests ordered. |
| --- | --- |

Continued on next page

#### Assaying New Reagents, Controls, Population Means & Reference Ranges

#### , Continued

| ##### PATIENT POPULATION MEAN | The normal patient population mean must be determined with a change of lot numbers of the reagents used or a change of instrument. |
| --- | --- |

| ##### CALCULATING PATIENT POPULATION MEAN | Follow the steps below to calculate the normal patient population means for the Prothrombin Time and Activated Partial Thromboplastin Time. Step; Action. 1.; Run the normal patients with the new lot number of reagents for the PT and APTT.. 2.; Calculate the mean of the results in seconds for each analyte.. 3.; Enter the new PT mean and ISI value of the new Neoplastin CI+ reagent into the analyzer.. |
| --- | --- |

| ##### NEW REAGENT LOT NUMBERS | When reagent lots change follow these instructions before use for PT and APTT. Step; Action. 1; Run a minimum of 20 Normals, 20 Abnormals and 20 therapeutic samples on both the old and new reagents.. 2; Send data to STAGO technical specialist for calculation. - Correlation studies. 3; When data returns and is evaluated by CLS, make adjustments to ranges if needed in the instrument file.. For CaCl2, run 20 samples for APTT with the old and new lots. The percent difference should be < 5 %. |
| --- | --- |

Continued on next page

#### Assaying New Reagents, Controls, Population Means & Reference Ranges

#### Continued

Document History Page

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Signature responsible person/date | Med. Dir. Reviewed/ Date | Lab Manager reviewed/ date | Date change Imp. |
| --- | --- | --- | --- | --- | --- |
| New |  | Sam Ibe 01/21/2014 |  |  |  |
| Minor | - Updated format - Updated title and Index number - Updated Safety Manual index number | Yvette Lingat 11/5/19 |  |  | 11/20/19 |