#### Calculating INR (International Normalized Ratio) for Prothrombin time test

| ##### POLICY | The Prothrombin Time is reported in INR so that values will not vary depending on the laboratory performing the testing. The INR is calculated by the instrument before resulting. |
| --- | --- |

| ##### CALCULATING INR | Follow the steps below to manually calculate the INR if needed and the instrument is unable to calculate. Step; Action. 1.; Obtain the results of the patient's Prothrombin time in seconds.. 2.; Calculate the ratio using the following formula: Ratio = Patient’s result Normal patient population mean. 3.; Example: Patient seconds = 18.2 Normal Population mean = 13.0 ISI = 1.5 Ratio = 18.2/13.0 = 1.4 INR = 1.41.5 = 1.7. |
| --- | --- |
| ##### NOTES | - Each laboratory should establish its own patient population mean for each lot number of reagents used. The appropriate geometric mean of the PT reference interval must be used in the INR calculation. - The assigned ISI is established by the reagent manufacturer and is unique for each lot number of reagents. The ISI reflects the sensitivity of the PT reagent to decreased levels of the vitamin K-dependent coagulation factors. This change in sensitivity will affect the calculation of the INR value - It is critical to calculate and report appropriate INR values. Reporting erroneous INR values may lead to use of excessive or insufficient vitamin K antagonist medication, which may result in bleeding or thrombotic complications in patients - The ISI has been determined against a secondary standard of the RBT (rabbit brain thromboplastin) according to the method recommended by the WHO. This value may vary from one lot to another; check that the value taken into account corresponds exactly to that of the lot of reagents being used. |

Continued on next page

#### Calculating INR (International Normalized Ratio) for Prothrombin time test, Continued

| ##### INR Calculation Verification | Patient reports are checked once per year to verify correct INR calculation performed by the analyzer or by the laboratory information system. Results are entered in a excel spreadsheet and signed off by the medical director/designee. The document is filed in the Verification Binder. STAGO INSTRUMENT INR CALCULATION VERIFICATION. PATIENT ID; INSTRUMENT INR RESULT; PT (SEC); CERNER INR RESULT; PT (SEC); ISI; REFERENCE RANGE; MANUAL INR CALCULATION; CALCULATION ACCEPTABLE YES/NO. 218310002518b; 3.58; 36.9; 3.6; 36.9; 1.28; 13.6; 3.59; YES. 218310003467b; 1.44; 18.1; 1.4; 18.1; 1.28; 13.6; 1.44; YES. 218310002564a; 2.9; 31.3; 2.9; 31.3; 1.28; 13.6; 2.91; YES. 218310002595b; 2.61; 28.8; 2.6; 28.8; 1.28; 13.6; 2.61; YES. 218310002465b; 1.16; 15.2; 1.2; 15.2; 1.28; 13.6; 1.15; YES. |
| --- | --- |
| ##### Controlled Documents | The following controlled documents support this process. |

Reference

Stago Neoplastin C1 Plus 10 Product insert September 2007

CAP Coagulation Checklist 2018

Form

Stago Instrument INR Calculation Verification Form

Continued on next page

#### Calculating INR (International Normalized Ratio) for Prothrombin time test, Continued

Document History Page

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Name of responsible person/date | Med. Dir. Reviewed/ date | Lab Manager reviewed/ date | Date change Imp. |
| --- | --- | --- | --- | --- | --- |
| Major | - Updated format - Updated index number - Added ISI and geometric mean information in Notes - Added INR calculation verification | Yvette Lingat 11/13/19 |  |  | 11/20/19 |

Imp. =Implemented