#### Sysmex XN550 Hematology Analyzer Validation Plan

| ##### Purpose | Before reporting patient test results, the new Sysmex XN550 Hematology Analyzer must demonstrate that it can obtain performance specification comparable to those established by the manufacturer for the following performance characteristics:  Accuracy or Systematic Error Experiment.  Precision or Random Error Experiment.  Reportable Range of test results for the test system.  Verify that the current regionally established reference ranges are appropriate for the laboratory's patient population. |
| --- | --- |

| ##### Policy | All laboratories (per CLIA regulation) are required to verify or establish performance specifications for any test system introduced by the laboratory before reporting patient test results. |
| --- | --- |

| ##### Scope of the Study | The tests will be performed by the Clinical Laboratory Scientist (CLS) at Harbor Mac Arthur and Mission Viejo Urgent Care Laboratory Assays will be validated using the Sysmex XN Hematology Analyzers for the following measured parameters: - WBC - RBC - Hemoglobin - Hematocrit - Platelets Count - Automated Differential Count: - Neutrophil (%) - Lymphocyte (%) - Monocyte (%) - Eosinophils (%) - Basophils (%) - IG (%) Specimen Type: Whole Blood collected on BD EDTA Vacutainer Tube following Kaiser Permanente Laboratory-approved blood collection procedures. Validation Process: The laboratory will demonstrate that it can obtain performance specifications comparable to those established for the current methodology being replaced. The following performance characteristics will be used: |
| --- | --- |

#### Sysmex XN550 Hematology Analyzer Validation Plan

#### , Continued

| ##### Scope of the study continued | Precision - Within-Run Precision: Three levels of controls will be tested in 10 replicates for all Hematology parameters that will be evaluated. Between-Run Precision: For each parameter, three levels of controls will be tested. Each level will be tested over 20 different runs. Discrete runs must be separated by at least 4-hour intervals. Accuracy/Method Comparison: For Hemogram: Correlate Sysmex XN550 Hematology Analyzer with the existing Beckman Coulter HmX Hematology Analyzer. Test at least 40 split samples for both normal and abnormal patients. For WBC Differential Count: Compare Sysmex XN550 Diff Count with Manual Diff Count Method for at least 20 samples for both normal and abnormal patients. Linearity/Reportable Range Verification Reportable range identifies the range of patient values that can be reported. The laboratory must verify the range that is relevant for the laboratory's patient population. This range is established through calibration/linearity verification. A linearity verification material is used for this study, provided with 5 levels of specimens with known or assigned values and should be analyzed in 4 measurements on each specimen. Carryover: Run low and high levels of controls using the values for H3, L1, L3 and calculate % Carryover using the following formula: (L1-L3)/(H3-L3)x100. Acceptable Carryover for all parameter is ≤1.0%. Reference Range Verification Run at least 30 normal patient samples, adult patients >18 years old, to verify the current established reference ranges. Sensitivity and Interferences: Refer to manufacturer’s published data in the XN550 IFU document. |
| --- | --- |

#### Sysmex XN550 Hematology Analyzer Validation Plan

#### , Continued

| Acceptance Criteria | Study; Description; Goal. Precision; Within-run: 10 samples of each control level in a single run. Between-run: Each level controls, at least every four hours apart for twenty runs.; - RBC ± 1.5 % - HCT ± 1.5 % - HGB ± 1.0 % - WBC ± 3.0 % - Platelet ± 10.0 % - WBC Differential ± 3SD/± 25% - Other Parameters ± 25%. Accuracy/ Method Comparison; Perform split sample study on at least 40 samples (normal and abnormal) using specimens tested on the Sysmex XN550 and Beckman Coulter DxH; Acceptance criteria for the main CBC parameters: hemoglobin, hematocrit, RBC, WBC, platelets, MCV, RDW: R>0.900 Slope 0.900-1.100 Average bias <10% Absolute bias (%) at clinical decision points (LLN and ULN) <1.5% for the differential parameters. Linearity/ Reportable Range Verification; Use manufacturer recommended Linearity Material; R > 0.9 < 10% bias from any point on line, except for the lowest level < 15%. Carryover; Run with alternating sequences of low and high concentration samples. The percent carryover is calculated using the values for H3, L1, L3 and the following formula L1-L3)/(H3-L3) x100; Acceptable Carryover for all parameter is ≤1.0%.. Reference Range Verification; Run 30 normal patient samples, >18 years old.; > or = to 90% of results must be within the established Reference Range.. Interferences; Use manufacturer’s data or perform manufacturer’s recommendations.; Use manufacturer’s published criteria for acceptability.. |
| --- | --- |

#### Sysmex XN550 Hematology Analyzer Validation Plan

#### , Continued

| Study Outcome (refer to Validation binder) | Study; Goal; Outcome. Precision; - RBC ± 1.5 % - HCT ± 1.5 % - HGB ± 1.0 % - WBC ± 3.0 % - Platelet ± 10.0 % - WBC Differential ± 3SD/± 25% - Other Parameters ± 25%; Within-run: Pass Fail. Between-run: Pass Fail. Accuracy/ Method Comparison; Acceptance criteria for the main CBC parameters: Hemoglobin, Hematocrit, RBC, WBC, platelets, MCV, RDW: R>0.900 Slope 0.900-1.100 Average bias <10% Absolute bias (%) at clinical decision points (LLN and ULN) <1.5% for the differential parameters; Pass Fail. Linearity/ Reportable Range Verification; R > 0.9 < 10% bias from any point on line, except for the lowest level < 15%; Pass Fail. Carryover; Acceptable Carryover for all parameter is ≤1.0%.; Pass Fail. Reference Range Verification; > or = to 90% of results must be within the established Reference Range.; Pass Fail. Interferences; Use manufacturer’s published criteria for acceptability.; Based on Manufacturer’s data: WBC If any of the following conditions are present, the system may erroneously report a low white blood cell count. • White blood cell aggregation If any of the following conditions are present, the system may erroneously report a high white blood cell count.. |
| --- | --- |

#### Sysmex XN550 Hematology Analyzer Validation Plan

#### , Continued

| Study Outcome continued | Study; Outcome. Interferences Data from: Sysmex XN-L Series Flagging Guide Document Number: 1399-LSS, Rev. 2, February 2019; Based on Manufacturer’s data: • Platelet aggregation • Poor lysing of red blood cells during analysis • Erythroblasts • Red blood cell aggregation (cold agglutinin) • Chylemia • Cryoprotein • Cryoglobulin • Fibrin • Giant platelets RBC If any of the following conditions are present, the system may erroneously report a low red blood cell count. • Red blood cell aggregation (cold agglutinin) • Microcytic red blood cells • Fragmented red blood cells If any of the following conditions are present, the system may erroneously report a high red blood cell count. • Leukocytosis (>100,000/μL) • Giant platelets HGB If any of the following conditions are present, the system may erroneously report a high hemoglobin value. • Leukocytosis (>100,000/μL) • Lipemia • Abnormal protein HCT If any of the following conditions are present, the system may erroneously report a low hematocrit value. • Red blood cell aggregation (cold agglutinin) • Microcytic red blood cells • Fragmented red blood cells If any of the following conditions are present, the system may erroneously report a high hematocrit value. • Leukocytosis (> 100,000/μL) • Hyperglycemia • Uremia • Spherocytes PLT If any of the following conditions are present, the system may erroneously report a low platelet count. • Platelet aggregation • Pseudothrombocytopenia • Giant platelets If any of the following conditions are present, the system may erroneously report a high platelet count. • Microcytic red blood cells • Fragmented red blood cells • Fragmented white blood cells • Cryoprotein • Cryoglobulin. |
| --- | --- |

#### Sysmex XN550 Hematology Analyzer Validation Plan

#### , Continued

The verification data for precisions, accuracy, correlation, linearity, carryover and reference range verification studies on the Sysmex XN 550 Hematology analyzer have been reviewed, and the performance of the method is considered acceptable for patient testing.

| ##### Controlled Documents | The following controlled documents support this policy. |
| --- | --- |
|  | References |
|  | Sysmex XN-L Series Flagging Guide Document Number: 1399-LSS, Rev. 2, February 2019 |
|  | Sysmex XN-L Automated Hematology Analyzer Method Verification Guide Document Number: 1251-LSS, Rev 2 |
|  | Clinical and Laboratory Standards Institute (CLSI). Laboratory Documents: Development and Control; Approved Guideline; Fifth Edition. (GP2-A5, 2006). |

#### Sysmex XN550 Hematology Analyzer Validation Plan

#### , Continued

Document History Page

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Name of responsible person/date | Med. Dir. Reviewed/ Date | Director of Lab Ops. reviewed/ date | Date change Implemented |
| --- | --- | --- | --- | --- | --- |
| New | Procedure for new XN instruments. | Yvette Lingat 4/15/2020 |  | Mary Lou Beaumont | 4/28/2020 |