PROCEDURE

Corewell Health East - Downtime Emergency Issue - Blood Bank - Taylor, Trenton, Wayne

- Principle

This procedure provides instructions for dispensing blood components before a patient’s compatibility testing is complete without the use of the Blood Bank Information System.

- Responsibility

Personnel who have completed the competency requirements will perform these tasks.

- Definitions

  - BBIS: Blood Bank Information System

  - RBC: Red Blood Cell

  - MRN: Medical Record Number

  - P-Tag: Product Tag

- Reagent/Equipment Needed

  - Downtime Emergency Bucket

- Procedure

  - The policies and procedures in this document shall be applied only in the following situations:

    - At the blood bank’s discretion, when the technologist determines that taking the time required to dispense components in the BBIS would be detrimental to the patient.

    - When the BBIS is unavailable for any reason, and components are requested in an emergency (i.e., required compatibility testing is incomplete or emergency issue requests for neonates).

  - Advance Preparation of O Negative Units (Emergency Issue Bucket)

    - The blood bank will keep 2-4 units of O Negative Red Blood Cells ready for immediate issue, prepared in advance with an attached, partially completed crossmatch tag.

    - Verify that the expiration of the units picked for the Emergency Issue Bucket are in date and greater than 10 days from expiration.

      - If the expiration date is unacceptable, return the unit to inventory and obtain a suitable replacement unit from the inventory.

    - For each unit, make sure the following information is documented on both labels and the bottom half of the P-Tag:

      - Donor blood type

      - Donor unit number

      - Donor unit product code and component description

      - Leave space on P-Tags to document patient identifiers at time of issue, see Physical Dispense section below.

    - Initial the “Tagged by” section of the P-Tag and adhere one of the P-Tag labels to the unit.

    - Remove two segments from each unit, label them with a unit number sticker from the back of the unit and place them in a biohazard bag or test tube.

    - Make a copy of the face label of each unit.

      - If segments have been placed in bags, place the face label copies in the pocket of the bags and place them in the Emergency Issue Bucket.

      - If segments are in test tubes attached to units, place the face label copies directly into the Emergency Issue Bucket.

    - Place an Uncrossmatched Blood sticker on each of the units.

    - Fold the P-Tag(s) and wrap around applicable unit or attach with a rubber band, placing the unit in the Emergency Issue Bucket.

  - Physical Dispense of Downtime Emergency Issue Units

    - Depending on the patient’s age:

      - If the patient is greater than 4 months old, then select the necessary O Negative units from the Emergency Issue Bucket.

      - If the patient is less than 4 months old, select the necessary O Negative unit from the Baby Bucket.

        - Refer to Corewell Health East - Providing Emergency Blood Products for Neonatal and Infant Transfusion – Farmington Hills, Grosse Pointe, Trenton, Wayne for steps on preparing the Baby Bucket.

    - Refer to Corewell Health East - Emergency Issue of Blood Products for required blood bank notification and nursing documentation.

    - After selection of blood component, transcribe the following patient information:

      - Patient name – documented 3 times, on both labels and the bottom half of the P-Tag.

      - Patient MRN – documented 3 times, on both labels and the bottom half of the P-Tag.

      - Patient wristband number (if available) – documented 3 times, on both labels and the bottom half of the P-Tag.

      - Patient blood type (only if the ABO/Rh has been performed on a current sample) – documented 3 times, on both labels and the bottom half of the P-Tag.

    - Visually inspect the unit(s), verifying that the quality control standards are met.

      - Circle V.I. OK, document the date, time, and your initials.

    - On the P-Tag, document who the unit is being issued to, the cooler number if used, and the operating room number, if applicable.

    - Retain the bottom half of the completed P-Tag, the unit segments, and the copy of the face label.

    - Give the unit(s) to the runner. If a neonatal unit is issued, remember to also provide the syringes from the Baby Bucket.

  - Steps Post Issue of Downtime Emergency Units

    - Obtain a Type and Screen specimen if patient does not have a current sample.

      - Using the retained segments, perform crossmatch on all transfused units that were not crossmatched at issue.

      - Refer to Corewell Health East – Serologic Crossmatching of Red Blood Cells – All Beaumont Hospitals section IV. E. Crossmatching Red Blood Cells Post-Issue.

    - Once the BBIS is available, update the computer record.

      - Perform the Emergency Issue function in BBIS, using the bottom half of the P-Tag, dispense form, and the copies of the unit face labels.

      - Result the post-issue crossmatch results.

    - When time permits, replace the unit(s) in the Emergency Issue Bucket.

- Revisions

Corewell Health reserves the right to alter, amend, modify, or eliminate this document at any time without prior written notice.

| Signature | Date | Revision # |  |
| --- | --- | --- | --- |
| Prepared by: Danae Regan, MLS(ASCP)CM, Lead Medical Technologist |  | 01 |  |
| Approved by: Fatima Bazzi, MLS(ASCP)CM, Blood Bank Supervisor |  | 01 |  |
| Approved by: Jeremy T Powers, MD Transfusion Services Medical Director |  | 01 |  |
| Approved by: Muhammad Arshad, MD Laboratory Medical Director |  | 01 |  |
| Reviewed by: (Signature) | Date | Revision # | Modification |