#### DETERMINATION OF PROTHROMBIN TIME (PT)
STA-NEOPLASTINE CI PLUS ASSAY ON STAGO-COMPACT

| ##### Principle | STA-Neoplastine CI PLUS is used for Prothrombin times, and extrinsic Factor Assays on the STA Compact. A mixture of thromboplastin is added to citrated plasma and the time of clot formation is determined. The STA Compact is a fully automated coagulation instrument, which uses an electromagnetic mechanical clot detection system. The oscillation of a steel ball within the cuvette with the thromboplastin and plasma is monitored by the STA Compact. When the oscillation of the steel ball is stopped by clot formation, the sensor registers the time in seconds. The prothrombin time is a basic coagulation screening test for the assessment of congenital and acquired deficiencies of the extrinsic pathway (factors II, V, VII, X). The prothrombin time can be prolonged in certain clinical states, i.e. warfarin therapy, intestinal re-absorption disorders, liver failure, fibrinolysis and DIC. The prothrombin time is also used to monitor Warfarin therapy because of its sensitivity to variations in the concentration of the Vitamin-K dependent factors II, VII, and X. Because of the variations in the prothrombin time results with different thromboplastins and instruments, it is recommended that the prothrombin time results be converted to an INR. The INR corresponds to the value of the ratio of the patient’s PT and the geometric mean PT of the normal reference population raised to the ISI (International Sensitivity Index) power: INR=Patient'sPTGeometricMeanPTISI The ISI value of a given thromboplastin is determined by performing PT’s on normal plasmas and Coumadin-treated patient plasmas with the given thromboplastin and the WHO reference thromboplastin. The slope of this regression curve of the matched pairs is the ISI for the thromboplastin. The ISI of the WHO reference thromboplastin is 1.0. The ISI value for the STA-Neoplastine CI PLUS is provided by the manufacturer for each lot of reagent. The geometric mean is calculated from the data collected for the PT reference range. |
| --- | --- |

Continued on next page

#### DETERMINATION OF PROTHROMBIN TIME (PT)
STA-NEOPLASTINE CI PLUS ASSAY ON STAGO-COMPACT, Continued

| ##### Specimen | - Citrated blood 9:1 (blood to anticoagulant) 3.2% sodium citrate. Follow NCCLS guidelines H3-A3 and H21-A2. No other anticoagulant is acceptable. - Centrifugation: 11 Minutes at 4,500 g 3 Minutes at 8,500 g - PT samples are stable for 24 hours if left unspun and stoppered at room temperature or refrigerated. Plasma is stable at room temperature for 8 hours. Plasma may be frozen immediately at a temperature -20ºC for up to two weeks if testing is delayed. - Unacceptable Specimens: Samples that are drawn short, over draws, clotted, or hemolyzed may yield incorrect results. |
| --- | --- |

| Reagents/ Equipments | STAGO STA-COMPACT Centrifuge STA Desorb U Reagent Grade Water STA Cleaner Solution Cuvette Roll STA Maxi Reducer Magnetic Stir Bars (Cat. No. 27425) STA Microtainer Pipette |
| --- | --- |

STA Microcups			     Pipette Tips

| ##### Reagents | - Reagent 1: STA Neoplastine CI Plus, Freeze-dried thromboplastin prepared from rabbit cerebral tissues. Reagents contain a specific inhibitor of heparin, thus, times obtained will not be affected by heparin at therapeutic levels. Transfer the entire contents of one vial of Reagent 2 into one vial of Reagent 1 of the same lot. Let sit 30 minutes at room temperature. Swirl gently. Add a stirring-bar (stir bars are to be reused. Wash thoroughly with reagent grade water and pat try before reuse) to the vial and place STA – Reducer and install the perforated plastic caps. Request the reagent or product drawer to open through MAIN MENU under LOADING and bar code the reagent. Place the reagent into a stirring position in the product drawer on the STA Compact. |
| --- | --- |

Continued on next page

#### DETERMINATION OF PROTHROMBIN TIME (PT)
STA-NEOPLASTINE CI PLUS ASSAY ON STAGO-COMPACT, Continued

| ##### Reagents, Cont.. | Reconstituted Stability - With the stirring bar, STA – Reducer and perforated plastic cap in place: 48 hours for 5 ml (Ref 00606 and 00605) and 10 ml (Ref 00667 and 00666) vials. - In its original capped vial (remove the STA – Reducer): 8 days at 2-8ºC for 5 ml vial ( 00605 and 00606) ONLY . Allow reagent to stand at room temperature for 30 minutes before use. - Reagent 2: 10 ml solvent. Ready to use. - STA – Coag Control Ⓝ+ABN citrated control plasmas. Normal and abnormal levels; or STA – System Control Ⓝ+Ⓟ citrated control plasmas, normal and abnormal levels; both freeze-dried. Reconstitute each vial with 1.0 ml reagent grade water. Let sit 30 minutes at room temperature. Swirl gently. Reconstituted stability on the STA Compact in 8 hours. |
| --- | --- |

| ##### Calibration | No calibration of the system is necessary for performing a PT if results are reported in seconds. If PT-INRs are reported, both the ISI and the geometric mean must be entered into the instrument for INR calculation. The ISI is automatically entered into the analyzer when the new lot of reagent is first barcoded and loaded onto the instrument. This can be verified by selecting the CALIB/CONTROL option on the MAIN MENU and highlighting PT from the list of tests and pressing Enter. The ISI is seen at the lower right hand corner of the screen, and should match the ISI listed on the barcode sheet that accompanies each box of Neoplastine reagent; it is always lot number specific. With each new lot of PT reagent, the geometric mean must be manually entered into the analyzer in order for correct calculation of the INR. Validation of the normal range geometric mean must be confirmed with each new lot of reagent. The geometric mean or Reference Time is entered in CALIB/CONTROL screen in the MAIN MENU. (See NOTES at the end of this protocol). |
| --- | --- |
|  | _____________________________________________________________ |

Continued on next page

#### DETERMINATION OF PROTHROMBIN TIME (PT)
STA-NEOPLASTINE CI PLUS ASSAY ON STAGO-COMPACT, Continued

| ##### Quality Control | - STA® -Coag Control Ⓝ+ABN or STA® -System ControlⓃ+Ⓟ After the reconstitution period, request the product drawer (STA Compact® to open through the MAIN MENU under LOADING and bar code the controls. Place the controls into the appropriate drawer. - QC must be run every 8 hours by manually programming from Quality Control menu. Analyzer will give a “blocked pipetting” error if patients are run after the 8 hour period expires without rerunning the QC. - All control ranges are monitored automatically by the STA Compact®. If any controls are outside the  2 SD range, the instrument will audibly and visually alarm the operator. Otherwise, the results can be found in the individual QC files. Control results are automatically filed in the STA Compact® QC file. All results for a 24-hour period will be reduced to a “mean” value at midnight. This mean is used in the statistical data and is plotted on the Levy-Jennings chart as a daily mean. - The CLS running the QC for the PM shift will be responsible for printing the QC for the day. - To print all the QC data points for the PT test, perform the following procedure prior to midnight. From the MAIN MENU under CAL/CONTROL select QUALITY CONTROL. Cursor to the PT test and press Enter  to view the Levy Jennings chart. Press F1 to view the results in tabular form. Press F6 and Select Execute then press Enter  to print the individual values under current controls. Press ESC key to exit (back to graph). Press F2 or F3 to view other levels and continue with F1 to view the result list. |
| --- | --- |

Continued on next page

#### DETERMINATION OF PROTHROMBIN TIME (PT)
STA-NEOPLASTINE CI PLUS ASSAY ON STAGO-COMPACT, Continued

| ##### Procedure | - Refer to START-UP Procedure for STA Compact before running patient specimens on the instrument at the start of each shift. - Request quality control. Under MAIN MENU under CALIB/CONTROL select QUALITY CONTROL and press Enter  Cursor to the PT test. Select PT by pressing F1 and then F10, Type in your Access Code (CQ) to run the QC. - Load patients’ samples: Access the sample drawer(s) through the MAIN MENU, under LOADING, Select Sample, Press Enter. After the drawer opens, identify the type of specimen, such as microsample (press F8), or stat (press F12). Identify the sample by bar-coding or manually entering on the keyboard the patient identification number and then placing the specimen into the drawer. - In MANUAL MODE, the operator must order the test(s) from the Selection menu, or from the Recorded Profile(s). Cursor to the test and press Enter  to select. When all tests are ordered, press F10 to save. - In AUTO MODE, the STA Compact will automatically order the test(s) selected in the AUTO MODE profile. If TELELOADING is selected as the AUTO MODE profile, the STA Compact will query the host computer and download the test(s) as well as assign the status (i.e., stat). |
| --- | --- |

Continued on next page

#### DETERMINATION OF PROTHROMBIN TIME (PT)
STA-NEOPLASTINE CI PLUS ASSAY ON STAGO-COMPACT, Continued

| ##### Procedure, Cont.. | - As soon as the sample drawer closes, the TEST STATUS screen will appear. If there is not enough reagent(s) to run the test(s), the suspect reagent(s) will appear in red with the amount of deficiency. This deficiency will BLOCK the SAMPLE PIPETTING. When this occurs, add the necessary reagent(s) to run the samples by responding N (NO) to the warning message ‘NEW TESTS ARE DELAYED – REACTIVATE?. Reagents can then be loaded in the drawer. By responding Y (YES) to the warning message ‘NEWTESTS ARE DELAYED – REACTIVE?’, the instrument will continue to perform all tests for which there is sufficient reagent (i.e. while waiting for reagents to stabilize after reconstitution). - All patient results are displayed on the TEST PANEL screen and automatically transmitted if selected on the SYSTEM STATUS menu. - For results in question that need operator intervention, cursor to the identification number in the TEST PANEL screen and press enter. This will display the FILE PROCESSING screen. Follow the options on the left-hand side of the screen ( i.e. F3 – rerun test) |
| --- | --- |

Continued on next page

#### DETERMINATION OF PROTHROMBIN TIME (PT)
STA-NEOPLASTINE CI PLUS ASSAY ON STAGO-COMPACT, Continued

| ##### Calculations | - All instrument measurements and calculations are performed by the STA Compact. Results are printed in a report form. - The International Normalized Ratio (INR) is automatically calculated by the STA Compact. INR=Patient'sPTGeometricMeanPTISI It is a derived correction factor (see formula) used to standardize PT values around the world to the World Health Organization (WHO) reference thromboplastin. - The International Sensitivity Index (ISI) is the value determined by the manufacturer for a specific lot number of thromboplastin as compared to the WHO thromboplastin standard. It is used as an exponential power for calculating INR by the laboratory performing a PT using that specific thromboplastin lot number with a specific instrument. It is stored in the CALIB/CONTROL screen for PT. - Geometric Mean PT (Reference Time) is the statistically calculated value based on the reference population for PT in seconds for the reagent. It must be manually entered in the CALIB/CONTROL screen for PT (see NOTES at the end of this protocol) - Patient’s PT is the measured PT in seconds for the reagent. - If no INR is read by the instrument, due to high PT, manually calculate the INR using the formula provided. Use a scientific calculator Example: PT reads 100 sec in the analyzer with no INR reading 90 sec (PT) 13.1 sec (Geometric mean) then press (xy) key in scientific calculator = 12.9 INR |
| --- | --- |

Continued on next page

#### DETERMINATION OF PROTHROMBIN TIME (PT)
STA-NEOPLASTINE CI PLUS ASSAY ON STAGO-COMPACT, Continued

| ##### Reference Range | 1. INR (Normal Range): 0.8-1.2 (Therapeutic Range): 2.0-4.5 (Critical Value): Greater than 5.0. Anything greater than 10.0 should be reported as “greater than 10.0”. Critical INR values must be phoned into the patient’s nursing unit, physician, the coumadin clinic or the in-patient pharmacist immediately. - Geometric mean of reference range: Varies per lot number, refer to instrument under PT calibration. - PT normal range = 11.5 – 14.7 sec. |
| --- | --- |

| ##### Reporting Results | - The INR is reported out to the 0.01 number (Example: 1.63) by the primary coagulation instrument. - The INR reported out to the 0.1 number (Example 1.6) by the lab information system to the clinicians. |
| --- | --- |

________________________________________________________________

| ##### Notes | - New lot of Thromboplastin: With each new lot of thromboplastin, the operator must enter the geometric mean time (Reference Time) before the STA Compact will allow QC to be run. Through the MAIN MENU select CALIB/CONTROL. Select CALIBRATION, press ENTER  Cursor to the PT test and press Enter  to select. Press ESC key for options. Select MODIFY REFERENCE TIME/RANGE, press ENTER Type in the Geometric Mean Time and save with F10. This screen also stores the ISI value, as downloaded from the reagent bar code sheet. - The PT reference range and geometric mean are validated with each change of PT reagent lot number. - A schedule for running the QC every 8 hours has to be followed. If QC is run too early, it needs to be rerun again after 2 to 3 hours. If run too late, the CLS has to put a comment/reason on the QC sheet for run delay. |
| --- | --- |

_______________________________________________________________

Continued on next page

#### DETERMINATION OF PROTHROMBIN TIME (PT)
STA-NEOPLASTINE CI PLUS ASSAY ON STAGO-COMPACT, Continued

| ##### Limitations of the Procedure | - Many commonly administered drugs affect the results obtained in PT testing. (Example: coumadin, heparin, and direct thrombin inhibitors). - STA Neoplastine CI+ contains a specific inhibitor of heparin. Therefore, only levels of heparin outside of the therapeutic range will affect the PT results. |
| --- | --- |

| ##### References | - STA-Neoplastine CI PLUS, (Cat # 00606) & STA-Neoplastine CI PLUS (Cat. #0667), are used for Prothrombin times Determination of Prothrombin Time (PT) by STA analyzers. Package insert for use in PT determinations. Revised February 2007 - STA® -Coag Control Ⓝ+ABN (Cat. No. 00676): citrated control plasmas normal and abnormal levels; or STA® -System ControlⓃ+Ⓟ (Cat. No. 00678): Control Plasmas for Assays of Coagulation Parameters on STA. Analyzers. Package insert 26379 - Revised November 2004; Revised February 2005, respectively. - STA Compact Operator’s Manual - Stago Manufacturer’s package inserts. - STA-Compact Software Version: 106.06; 106.07.Vx; 106.08; 106.10 - Stago procedures, DSI-TSD-US, 08/09/2007 - cd |
| --- | --- |

Continued on next page

#### DETERMINATION OF PROTHROMBIN TIME (PT)
STA-NEOPLASTINE CI PLUS ASSAY ON STAGO-COMPACT, Continued

| ##### Author | Carirose Aquisap, CLS Richard Ulep, CLS, MT (AMT) |
| --- | --- |

| ##### Distributions | Kaiser Permanente Riverside Medical Center Laboratory |
| --- | --- |

End

#### DETERMINATION OF PROTHROMBIN TIME (PT)
STA-NEOPLASTINE CI PLUS ASSAY ON STAGO-COMPACT, Continued

Reviewed and approved by:

| SIGNATURE | DATE |
| --- | --- |
| Denise Topliff, MHA, CLS Director, Clinical Laboratory – Riverside Medical Center |  |
| Dong Quach, M.D. Medical Director, Clinical Laboratory – Riverside Medical Center |  |

#### DETERMINATION OF PROTHROMBIN TIME (PT)
STA-NEOPLASTINE CI PLUS ASSAY ON STAGO-COMPACT, Continued

HISTORY PAGE

#### Effective Date: 12-01-2009

| Change type: New, major, minor | Changes made to SOP - describe | Signature responsible person/date | Medical Director review/date | Laboratory Director review/date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| Minor | Replaced policy and procedure number LIV to LXII | Carirose Aquisap 11-21-09 | Dr. Dong Quach 11-25-09 | Denise Topliff 11-22-09 | 12-01-09 |
| Minor | Added #3 requirement under – Specimen | Carirose Aquisap 11-21-09 | Dr. Dong Quach 11-25-09 | Denise Topliff 11-22-09 | 12-01-09 |
| Minor | Revised #2 requirement and added #4 Under – Quality Control Section | Carirose Aquisap 11-21-09 | Dr. Dong Quach 11-25-09 | Denise Topliff 11-22-09 | 12-01-09 |
| Minor | Added Notes #3 | Carirose Aquisap 11-21-09 | Dr. Dong Quach 11-25-09 | Denise Topliff 11-22-09 | 12-01-09 |
| Minor | Revised #1 and added #2 requirement of Reporting Results | Carirose Aquisap 11-21-09 | Dr. Dong Quach 11-25-09 | Denise Topliff 11-22-09 | 12-01-09 |
| Minor | Added #5 under Calculation | Carirose Aquisap 11-21-09 | Dr. Dong Quach 11-25-09 | Denise Topliff 11-22-09 | 12-01-09 |
| Minor | Established Reference Ranged for the KP Riverside Lab | Carirose Aquisap 11-21-09 | Dr. Dong Quach 11-25-09 | Denise Topliff 11-22-09 | 12-01-09 |
| Major | Added sample stability under Specimen. Revised Geometric Mean under Reference Range. Revised sample centrifugation speed under Specimen. | Richard Ulep | Dr. Dong Quach | Denise Topliff | 04/06/2013 |
| Minor | Added items under Reagents/ Equipments. Template Change. | Richard Ulep | Dr. Dong Quach | Denise Topliff | 04/06/2013 |