#### DETERMINATION OF ACTIVATED PARTIAL THROMBOPLASTIN TIME (APTT)

| ##### Principle | STA®-PTT A⑤, involves the re-calcification time of plasma in the presence of a standardized amount of platelet substitute and a specific activator. This procedure minimizes test variables by standardizing the contact activation and optimizes the concentration of platelet-like phospholipids. The time of clot formation is measured on the STA-Compact®. The STA-Compact® is a fully automated coagulation instrument, which uses an electromagnetic mechanical clot detection system. The oscillation of a steel ball within the cuvette with the reagents and plasma is monitored by the STA-Compact®. When the oscillation of the steel ball is stopped by clot formation, the sensor registers the time. The activated partial thromboplastin time (APTT) is a basic screening test for the assessment of congenital and acquired deficiencies of the intrinsic pathway (factors XII, XI, IX, and VIII). It is also the test most often used to monitor heparin therapy. |
| --- | --- |

| ##### Specimen | - Citrated blood 9:1 (blood to anticoagulant) 3.2% sodium citrate. Follow NCCLS guidelines H3-A3 and H21-A2. No other anticoagulant is acceptable. - Centrifugation: 11 minutes at 4,500g 3 minutes at 8,500g - Plasma Storage: 4 hours at 20 +/- 5ºC 2 hours at 20 +/- 5ºC if on heparin therapy. Plasma may be frozen at a temperature -20ºC for up to two weeks if testing is delayed. - Unacceptable Specimens: Samples that are short, over draws, clotted or hemolyzed may yield incorrect results. |
| --- | --- |

| Reagents/ Equipments | STAGO STA-COMPACT Centrifuge STA Desorb U Reagent Grade Water STA Cleaner Solution Cuvette Roll (contains 1000 cuvettes) STA Microtainer Pipette STA Microcups Pipette Tips ________________________________________________________ |
| --- | --- |

Continued on next page

#### DETERMINATION OF ACTIVATED PARTIAL THROMBOPLASTIN TIME (APTT), Continued

| ##### Reagents | 1. Reagent 1: PTT-Automate: The reagent is a preparation of brain phospholipids in a buffered suspension of particulate activators with added stabilizers. Reconstitute each vial with 5.0 ml of reagent grade water. Let sit 30 minutes at room temperature. Shake vigorously for 10 seconds with rubber stopper on. Remove rubber stopper and replace the perforated plastic cap. Request the product drawer (STA COMPACT) to open through the MAIN MENU under LOADING and bar code and reagent. Place the reagent into the drawer: Reconstituted Stability: -24 hours with the perforated cap in place on the analyzer. -7 days at 2-8ºC in its original capped vial. Allow the reagents to stand at room temperature for 30 minutes before use. - Reagent 2: 0.025 M CaCl2 Ready to use. Request the product (STA COMPACT) drawer to open through MAIN MENU under LOADING and bar code reagent. Place the reagent into drawer. Stability is 72 hours. STA-Coag Control Ⓝ+ABN citrated control plasmas normal and abnormal levels; both freeze-dried. Reconstitute each vial with 1.0 ml reagent grade water. Let sit 30 minutes at room temperature. Swirl gently. Reconstituted stability on the STA-Compact is 8 hours. |
| --- | --- |

| ##### Calibration | No calibration of the system is necessary for performing an APTT |
| --- | --- |

Continued on next page

#### DETERMINATION OF ACTIVATED PARTIAL THROMBOPLASTIN TIME (APTT), Continued

| ##### Quality Control | - STA-Coag Control Ⓝ+ABN: After the reconstitution period, request the product drawer (STA Compact) to open through the MAIN MENU under LOADING and bar code the controls. Place the controls into the appropriate drawer. - QC must be rerun every 8 hours by manually programming from Quality Control Menu. Analyzer will give a “blocked pipetting” error if patients are run after the 8 hour period expires without rerunning the QC. - All control ranges are monitored automatically by the STA-Compact. If any controls are outside the +/- 2SD range, the instrument will audibly and visually alarm the operator. Otherwise, the results can be found in the individual QC files. Control results are automatically filed in the STA-Compact QC file. All results for a 24 hour period are converted to a “mean” value at midnight. This mean is used in the statistical data and is plotted on the Levy-Jennings chart as a daily mean. - To print all the QC data points for the APTT test, perform the following procedure prior to midnight. From the MAIN MENU under CAL/CONTROL, select QUALITY CONTROL press Enter  Cursor to the APTT test and press Enter  to view the Levy-Jennings chart. Press F1 to view the results in tabular form. Press F6, select Execute then press Enter  to print the individual values under current controls. Press ESC key to exit (back to graph). Press F2 or F3 to view other levels and continue with F1 to view the result list. The CLS running the QC for the PM shift will be responsible for printing the QC for the day. |
| --- | --- |

Continued on next page

#### DETERMINATION OF ACTIVATED PARTIAL THROMBOPLASTIN TIME (APTT), Continued

| ##### Procedure | Allow reagents and controls to equilibrate “on-board” for 7 minutes, before running. Refer to the START-UP procedure for the STA-Compact before running patient specimens on the STA-Compact at the start of each shift. Request Quality Control: Through MAIN MANU under CAL./CONTROL, select QUALITY CONTROL, and press Enter  Cursor to the APTT test. Select APTT by pressing F1 and then F10 then enter the Access code Enter  to run QC. Load patient’s samples: Access the sample drawer(s) through the MAIN MENU under LOADING select Sample then press Enter  After the drawer opens, identify the type of specimen, such as microsample, (press F8), or stat, (press F12). Identify the sample by bar-coding or manually entering on the keyboard the patient identification number and then placing the specimen into the drawer. In MANUAL MODE, the operator must order the test(s) from the selection menu or from the recorded profile(s). Cursor to the test or Recorded Profile and press Enter  to select. When all tests are ordered, press F10 to save. In AUTO MODE, the STA-Compact will automatically order the test(s) selected in the AUTO MODE profile. If TELELOADING is selected as the AUTO MODE profile, the STA Compact will query the host computer and download the test(s) as well as assign the status (i.e. stat). As soon as the sample drawer closes, the TEST STATUS screen will appear. If there is not enough reagent(s) to run the test(s), the suspect reagent(s) will appear in red with the amount of the deficiency. This deficiency will BLOCK the SAMPLE PIPETTING. When this occurs, add the necessary reagent(s) to run the samples by responding N (NO) to the warning message ‘NEW TESTS ARE DELAYED – REACTIVATE?’ Reagents can then be loaded in the drawer. By responding Y (YES) to the warning message ‘NEW TESTS ARE DELAYED – REACTIVATE?’ the instrument will continue to perform all tests for which there is sufficient reagent (i.e. while waiting for reagents to stabilize after reconstitution). All patient results are displayed on the TEST PANEL screen and automatically transmitted. For results in question that need operator intervention, cursor to the identification number in the TEST PANEL screen and press enter. This will display the FILE PROCESSING screen. Follow the options on the left hand side of the screen, (i.e. F3 – rerun test). |
| --- | --- |

#### DETERMINATION OF ACTIVATED PARTIAL THROMBOPLASTIN TIME (APTT), Continued

| ##### Calculations | No calculations are required for the APTT |
| --- | --- |

| ##### Reference Range | Normal range: PTT: 25-37 Seconds Critical Value: a) Non medicated / non heparin therapy: > 68 sec b) Heparin therapy: > 140 sec Instrument Range: Min 20 sec Max 200 sec |
| --- | --- |

| ##### Reporting Results | The results for an APPT are reported out in seconds to the nearest 0.1 sec (Example: 31.0 Seconds) |
| --- | --- |

| ##### Notes | - New lot of PPT-Automate: If RATIO is selected as CALIBRATION MODE in page 2 of the test setup, the operator must enter the REFERENCE TIME before the STA-Compact will allow QC to be run. Through the MAIN MENU; select CAL/CONTROL. Select CALIBRATION, press Enter  Cursor to the APTT test, press Enter  Press ESC key for options. Select MODIFY REFERENCE TIME/RANGE and press Enter  Type in your Access Code press Enter  Enter reference time and save with F10. This is NOT necessary if the calibration is set to RAW mode in test set-up (page-2) - With the introduction of STA-Coag Control Ⓝ+ABN (Cat. No. 00676), it is no longer necessary to use a stir bar with PTT Automate Reagent. - A schedule for running the QC every 8 hours has to be followed. If QC is run too early, it needs to be run again after 2 to 3 hours. If run too late, the CLS has to put the reason for delay on the comment section of the QC sheet. |
| --- | --- |

Continued on next page

#### DETERMINATION OF ACTIVATED PARTIAL THROMBOPLASTIN TIME (APTT), Continued

| ##### Limitations of the Procedure | - Many commonly administered drugs may affect the results obtained in APTT testing. (Example: heparin and coumadin). - When monitoring heparin therapy, any release of platelet factor 4 (PF4) which is a potent inhibitor of heparin represents a potential source of error. - A congenital or acquired inhibitor may prolong the APTT. |
| --- | --- |

| ##### References | - STA® PTT-A⑤, Determination of Activated Partial Thromboplastin Time (APTT) Package insert for use in APTT determinations. 24528 - Revised February 2007. - STA® -Coag Control Ⓝ+ABN (Cat. No. 00676): citrated control plasmas, normal and abnormal levels; or STA® -System ControlⓃ+Ⓟ (Cat. No. 00678): Control Plasmas for Assays of Coagulation Parameters on STA. Analyzers. Package insert 26379 - Revised Control Plasma for Coagulation Tests on STA®. Package insert 26378-01-Revised October 2001 and 26384 -Revised June 2000, respectively. - STA-Compact® Operators Manual. - STA-Compact® Software Version: 106.06; 106.07.Vx; 106.08 - Stago DSI-TSD-US, Revised 07/02, 08/09/2007 – cd - Kaiser Riverside Chemistry Procedure Manual revised 10/10/06. |
| --- | --- |

Continued on next page

#### DETERMINATION OF ACTIVATED PARTIAL THROMBOPLASTIN TIME (APTT), Continued

| ##### Author | Carirose Aquisap, CLS Richard Ulep, CLS, MT (AMT) |
| --- | --- |

| ##### Distributions | Kaiser Permanente Riverside Medical Center Laboratory |
| --- | --- |

End

#### DETERMINATION OF ACTIVATED PARTIAL THROMBOPLASTIN TIME (APTT), Continued

Reviewed and approved by:

| SIGNATURE | DATE |
| --- | --- |
| Denise Topliff, MHA, CLS Director, Clinical Laboratory – Riverside Medical Center |  |
| Dong Quach, M.D. Medical Director, Clinical Laboratory – Riverside Medical Center |  |

#### DETERMINATION OF ACTIVATED PARTIAL THROMBOPLASTIN TIME (APTT), Continued

HISTORY PAGE

#### Effective Date: 12-1-2009

| Change type: New, major, minor | Changes made to SOP - describe | Signature responsible person/date | Medical Director review/date | Laboratory Director review/date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| Minor | Replaced Policy and Procedure # LV to LXIII | Carirose Aquisap 11/21/09 | Dr. Dong Quach 11/25/09 | Denise Topliff 11/22/09 | 12/1/09 |
| Minor | Added #2 requirement under QC Revised #4 requirement under QC | Carirose Aquisap 11/21/09 | Dr. Dong Quach 11/25/09 | Denise Topliff 11/22/09 | 12/1/09 |
| Minor | Added Notes #3 | Carirose Aquisap 11/21/09 | Dr. Dong Quach 11/25/09 | Denise Topliff 11/22/09 | 12/1/09 |
| Minor | Added Reference #6 | Carirose Aquisap 11/21/09 | Dr. Dong Quach 11/25/09 | Denise Topliff 11/22/09 | 12/1/09 |
| Major | Established Reference Ranges for KP Riverside Laboratory | Carirose Aquisap 11/21/09 | Dr. Dong Quach 11/25/09 | Denise Topliff 11/22/09 | 12/1/09 |
| Major | Incorporated Correction Procedure for aPTT | Carirose Aquisap 11/21/09 | Dr. Dong Quach 11/25/09 | Denise Topliff 11/22/09 | 12/1/09 |
| Major | Reference range changed under Reference Range. Revised sample centrifugation speed under Specimen. | Richard Ulep | Dr Dong Quach | Denise Topliff | 04/06/2013 |
| Minor | Added items under Reagents/ Equipments. Template change. | Richard Ulep | Dr Dong Quach | Denise Topliff | 04/06/2013 |