#### CRITERIA FOR APTT MIXING STUDY

| ##### Purpose | - This document describes the criteria for performing APTT Mixing Study. |
| --- | --- |

| ##### Policy | - In order to perform the APTT mixing study, the baseline APTT result should exceed the upper limit of the laboratory’s defined reference range by 5 seconds or more. |
| --- | --- |

| ##### Reference Range | - APTT Reference Range: 25-37 Seconds |
| --- | --- |
| Criteria | _____________________________________________________________ - Criteria for performing APTT Mixing Study: - 1.) Verify the APTT Mixing Study order to the RN if the patient is on - Heparin or an anticoagulant. |

A.) Patient on Heparin therapy.

i.) If the patient is on Heparin, and the order is for both APTT and APTT

Mixing Study, perform the APTT. Report out the APTT and APTT baseline

result on the APTT Mixing Study. Under the comment section of the APTT

Mixing Study, enter a free text comment “Mixing Study is not indicated

because the patient is on Heparin”.

ii.) If the order is only for APTT Mixing Study, and the patient is on

Heparin, perform the APTT. Report out the baseline APTT, and enter a free

text comment “ Mixing Study is not indicated because the patient is on

Heparin”.

B.) Patient not receiving anticoagulant therapy or on anticoagulant other

Heparin.

i.) Perform the test as ordered.

ii.) APTT baseline result should exceed the upper limit of the laboratory`s

defined reference range by 5 seconds or more. If APTT mixing study is

ordered, and the baseline result is within the reference range or doesn`t

exceed the upper limit of the reference range by 5 seconds or more, a

canned text comment will be entered indicating that the test is not

performed or indicated. Refer to APTT MIXING STUDY P& P for

reporting.

Continued on next page

| #### CRITERIA FOR APTT MIXING STUDY, Continued ##### ___________________________________________________________________ The following controlled document supports this procedure: ##### Controlled ##### Documents |  |
| --- | --- |

| Document Number | Title |
| --- | --- |
| LCOS8010 | Determination of Activated Partial Thromboplastin Time (aPTT) aPTT Assay on Stago - Compact |
| COAG.SCPMG.002 | APTT Mixing Study |

| ##### Author | Richard Ulep, CLS, MT (AMT) |
| --- | --- |
| ##### Distributions | ______________________________________________________________ Kaiser Permanente Riverside Medical Center Laboratory |

End

#### CRITERIA FOR APTT MIXING STUDY, Continued

Reviewed and approved by:

| SIGNATURE | DATE |
| --- | --- |
| Denise Topliff, MHA, CLS Director, Clinical Laboratory – Riverside Medical Center |  |
| Dong Quach, M.D. Medical Director, Clinical Laboratory – Riverside Medical Center |  |

#### CRITERIA FOR APTT MIXING STUDY, Continued

HISTORY PAGE

#### Effective Date: ___________

| Change type: New, major, minor | Changes made to SOP - describe | Signature responsible person/date | Medical Director review/date | Laboratory Director review/date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| New |  | Richard Ulep | Dr Dong Quach | Denise Topliff |  |