#### SPECIMEN HANDLING WITH ELEVATED HEMATOCRIT FOR COAGULATION TESTS

| ##### Purpose | - This document describes the procedure for correcting the amount of anticoagulation present in the tube when the hematocrit is greatly elevated. - Patients with elevated hematocrit value of >55% require special handling procedure for coagulation tests done at the medical center. |
| --- | --- |

| ##### Principle | - Elevated hematocrit can be seen in newborn, patients with polycythemia vera and other conditions. - When a patient has a high hematocrit, the plasma volume will be too low for the amount of sodium citrate solution in the tube. Thus, coagulation test results like PT, aPTT, may be falsely prolonged because of the excess anticoagulation in the plasma, as well as erroneous results for other calcium-dependent clotting tests such as clottable protein C/protein S and factor assays. |
| --- | --- |

| ##### Reagents / Equipment | BD Vacutainer Buffered Na Citrate 0.109M, 3.2% Pipettes Pipette tips |
| --- | --- |

Continued on next page

#### SPECIMEN HANDLING WITH ELEVATED HEMATOCRIT FOR COAGULATION TESTS, Continued

| ##### Procedure | Step; Action. 1.; The medical center uses BD Citrate Tube (2.7 ml) which contains 0.3 ml of 0.109M, 3.2% buffered Na Citrate.. 2.; To correct for the increase anticoagulant volume, the amount of sodium citrate must be reduced in the tube before drawing the blood sample. The formula to calculate the appropriate sodium citrate volume is: X = (100-PCV Vol) (595-PCV) X = Volume of sodium citrate required for unit volume of blood. PCV = Packed Cell Volume in Percent (%) (Hematocrit) Vol. = Blood draw volume required in the blood collection tube. Example: Patient has a hematocrit of 60% and blood is to be drawn into a 2.7 ml blue-top tube. Adjusted citrate volume is calculated as follows: (100-60) x 2.7 ml = 0.2 ml of sodium citrate anticoagulant (595-60). 3.; Once the corrected volume was determined, aspirate the Na citrate using Finn pipette the difference of 0.3 ml (original container) and the corrected amount leaving only the adjusted citrate in the tube. Example: Na Citrate original Volume : 0.3 ml - Calculated Citrate Volume : 0.2ml 0.1ml (100 ul) = amount of Na citrate to be taken out of original tube.. 4.; Draw the patient using the calculated Na Citrate Tube.. |
| --- | --- |

Continued on next page

#### SPECIMEN HANDLING WITH ELEVATED HEMATOCRIT FOR COAGULATION TESTS, Continued

| ##### Procedural Notes | - Using the above formula will help to ensure that the anticoagulant volume is correct when drawing patients with elevated hematocrit. - There are no current data to support a recommendation for adjusting the citrate concentration in the presence of severe anemia (hematocrit<20%). |
| --- | --- |

| ##### References | - NCCLS Guideline “ Collection, Transport and Processing of Blood Specimens for Testing Plasma-Based Coagulation Assays; Approved Guideline – Fourth Edition” H21-A4 Vol. 23 No. 35 (Refer to Appendix, Amount of Anticoagulant Solution/Volume of Blood at Different Packed Cell Volume Values, page 9). - CAP Checklist Question HEM.22830 Phase I, College of American Pathologists Lab Accreditation Program, 325 Waukegan Road, Northfield, IL Sept. 30, 2004. - Workfolk, Laura “Pre-analytical Variables of Coagulation Testing,” Advance for Administrators of the Laboratory, Aug. 2003. |
| --- | --- |

| ##### Author | Richard Ulep, CLS, MT (AMT) |
| --- | --- |

| ##### Distributions | Kaiser Permanente Riverside Medical Center Laboratory |
| --- | --- |

End

#### SPECIMEN HANDLING WITH ELEVATED HEMATOCRIT FOR COAGULATION TESTS, Continued

Reviewed and approved by:

| SIGNATURE | DATE |
| --- | --- |
| Denise Topliff, MHA, CLS Director, Clinical Laboratory – Riverside Medical Center |  |
| Dong Quach, M.D. Medical Director, Clinical Laboratory – Riverside Medical Center |  |

#### SPECIMEN HANDLING WITH ELEVATED HEMATOCRIT FOR COAGULATION TESTS, Continued

HISTORY PAGE

#### Effective Date: ___________

| Change type: New, major, minor | Changes made to SOP - describe | Signature responsible person/date | Medical Director review/date | Laboratory Director review/date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| NEW |  | Richard Ulep | Dr Dong Quach | Denise Topliff |  |
| Minor | Template change | Richard Ulep | Dr Dong Quach | Denise Topliff |  |