#### CALIBRATION VERIFICATION of PRE-CALIBRATED TESTS

| ##### Purpose | - This document describes the procedure for performing calibration verification of pre-calibrated tests on the Stago Compact analyzer. The Clinical Laboratory Improvement Act (CLIA) states calibration verification is only required for “modified” moderately complex tests. |
| --- | --- |

| ##### Policy | - Fibrinogen and D-Dimer calibration verifications must be performed semi-annually, and when major maintenance is done on the analyzer as required. Calibration Verification materials can be purchase or made by diluting a specimen with an elevated value and using diluent done automatically by the instrument. |
| --- | --- |

| ##### Procedure | - 1. FIBRINOGEN: - a.) Find a patient with a Fibrinogen between 700-800 mg/dl. Possible source for high Fibrinogen will be L&D or post-partum patients, and patients with a positive blood culture. |
| --- | --- |
|  | - b.) Run the Fibrinogen test. Once this test is complete, and the result - is in between 700-800mg/dl, add on the following dependent tests: - i.) Fib 1:15 (Fib x15) - ii.) Fib 1:20 (Fib- repeats the primary test) - iii.) Fib 1:40 (Fib x40) - iv.) Fib 1:100 (Fib x100) - c.) Print the test results. To print the results, select the patient on the test - panel screen. Press Enter to access the File Processing Screen. Once the - patient results are displayed, select (F6) to print. - d.) Submit the data to Stago for evaluation. Provide the company with the - Calibration data indicated on the barcode calibration sheet specific for that - reagent lot number. - e.) File the calibration verification in the appropriate binder for our records. ______________________________________________________________ - Continued on next page |

#### CALIBRATION VERIFICATION of PRE-CALIBRATED TESTS, Continued

##### Procedure, cont…

- 2. D-DIMER:

- a.)  Find a patient with a D-Dimer between 3.0-3.8 ug/ml FEU (3,000-

- 3,800 ng/ml FEU). Screen patients to find one with a suitable value.

- b.) Run the D-Dimer test. Once the test is complete, and the result is in

- between 3.0-3.8 ug/ml FEU (3,000-3,800 ng/ml FEU), add the following

- dependent tests:

- i.) 	D-Di 1:2 (D-DI X2)

- ii.) 	D-Di 1:8 (D-DI X8)

- iii.)	D-Di 1:15 (D-DI X15)

- c.) Print the test results. To print the results, select the patient on the test

- panel screen. Press Enter to access the File Processing Screen. Once the

- patient results are displayed, select (F6) to print.

- d.) Submit the data to Stago for evaluation. Provide the company with the

- Calibration data indicated on the barcode calibration sheet specific for that

- reagent lot number.

- e.) File the calibration verification in the appropriate binder for our records.

| ##### Target Recovery | - Acceptable Recovery Limit: - 1.) Fibrinogen: + 20% - 2.) D-Dimer: |
| --- | --- |
|  | - a.) 0.0-1.0 ug/ml FEU (0-1,000 ng/ml): + 0.15 ug/ml FEU ( 150ng/ml) |
|  | - b.) 1.0-4.0 ug/ml FEU (1.0-4.0 ng/ml): + 0.50 ug/ml FEU (500ng/ml) |

Continued on next page

| #### CALIBRATION VERIFICATION of PRE-CALIBRATED TESTS, Continued ##### ___________________________________________________________________ Stago Compact (corrector), Rev. 4/08, 1/09, 6/11 ##### Reference |  |
| --- | --- |

| ##### Author | Richard Ulep, CLS, MT (AMT) |
| --- | --- |
| ##### Distributions | ______________________________________________________________ Kaiser Permanente Riverside Medical Center Laboratory |

End

#### CALIBRATION VERIFICATION of PRE-CALIBRATED TESTS, Continued

Reviewed and approved by:

| SIGNATURE | DATE |
| --- | --- |
| Denise Topliff, MHA, CLS Director, Clinical Laboratory – Riverside Medical Center |  |
| Dong Quach, M.D. Medical Director, Clinical Laboratory – Riverside Medical Center |  |

#### CALIBRATION VERIFICATION of PRE-CALIBRATED TESTS, Continued

HISTORY PAGE

#### Effective Date: ___________

| Change type: New, major, minor | Changes made to SOP - describe | Signature responsible person/date | Medical Director review/date | Laboratory Director review/date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| New |  | Richard Ulep | Dr Dong Quach | Denise Topliff |  |