| # Purpose | This document describes how to perform ABO testing by the tube method. This procedures covers full ABO typing (red cells and serum/plasma testing), red cell only testing and serum/plasma testing only testing |
| --- | --- |

| # Policy | - Clotted specimens and samples drawn into EDTA may be tested up to 14 days after collection. - Donor blood may be tested up to the unit expiration date. - Routine tests on patients must include both red cell and plasma/serum tests - It is permissible to test red blood cells only when: - Confirming ABO of donor blood - Testing specimens from infants less than 4 months of age. - If a patient's red cells test AB and also Rh positive, a negative control MUST be run, e.g testing the patient's red cells with 6% albumin. |
| --- | --- |
|  | - ABO DISCREPANCIES MUST BE RESOLVED BEFORE UNITS CAN BE RELEASED FOR TRANSFUSION. - IF THE CLINICAL CONDITION OF THE PATIENT MAKES IT NECESSARY TO RELEASE BLOOD BEFORE THE DISCREPANCY IS RESOLVED, OBTAIN ENOUGH PRETRANSFUSION SAMPLE TO PERMIT ADDITIONAL TESTING AND RELEASE GROUP O RED BLOOD CELLS OF THE APPROPRIATE Rh TYPE ONLY. |
|  | - Each reagent must be tested on each day of use with appropriate positive control and negative control. - Red cell suspensions do not need to be washed, but they must be suspended in saline, not the patient's or donor's own plasma. |
|  | Continued on next page |

| EQUIPMENT AND REAGENTS | - Properly labeled blood sample. - 10 X 75 or 12 X 75mm test tubes - Test Tube rack - Uniform drop pipettes or transfer pipettes - Serologic centrifuge - Marker pen - Blood grouping reagents Anti-A, Anti-B, Anti-A,B - Reagent red blood cells: group A1 and group B - Isotonic saline, 0.85-0.9% sodium chloride/phosphate buffered saline - Note: anti-A,B is not used routinely on patients. It is mainly used for confirming group O donor units received from the blood center and in resolving ABO discrepancies. |
| --- | --- |

| # Procedure | Follow the steps below for patient testing and complete ABO test |  |
| --- | --- | --- |
|  | Step | Action |
| # Red Cell and serum Testing | 1. | Prepare a 3% to 5% suspension of patient or donor red blood cells in isotonic saline. Note: It is not necessary to wash cells before preparing the suspension, but they must be suspended in saline, NOT the patient's or donor's own plasma. |
|  | 2. | Label 4 test tubes; - one set of 2 tubes for cell testing; one tube labeled "A", and one tube labeled "B" and enough information to uniquely identify the patient or donor unit. - One set of 2 tubes for serum/plasma testing; one tube labeled "a" and one with "b" and enough information to uniquely identify the patient or donor unit. Note: If using the accession number on the test tube to identify the patient, there must also be enough letters of the patient's name to uniquely identify that patient. The accession number only is not sufficient. |
|  |  | Continued on next page |
| # Procedure | Follow the steps below for patient testing and complete ABO test |  |
|  | Step | Action |

|  | 3. | - Add one drop of each blood-grouping reagent to the appropriately tubes labeled "A" and "B". - Add two drops of the patient's serum or plasma to the appropriately test tube labeled "a" and "b". |
| --- | --- | --- |
|  | 4. | Add one drop of a 3-5% saline suspension of patient red cells to the tubes labeled "A" and "B". |
|  | 5. | Gently mix the reagent red cells. Add 1 drop of reagent red cells to the appropriate tube labeled "a" or "b". |
|  | 6. | Mix the contents of each tube gently but thoroughly and centrifuge on high speed for the appropriate time for the centrifuge being used. (Optimal times are posted on each centrifuge, usually about 15 seconds). |
|  | 7. | Gently agitate each tube to resuspend the red cell buttons. - Examine for agglutination using agglutination viewer. - Grade reaction and record results immediately, before the tubes are discarded. Note: Before entering results in the computer or on a manual sheet check the sample tube against the patient record (work card or computer) for name and Medical Record number, or donor unit number if confirming a donor unit. |

|  | Follow the steps below for Red Cell only Typing; Donor Unit confirmation, Cord blood testing, etc… |  |
| --- | --- | --- |
|  | Step | Action |
| # Red Cell only Testing | 1. | Prepare a 3% to 5% suspension of patient, donor or cord red blood cells in isotonic saline. Note: It is not necessary to wash cells before preparing the suspension, but they must be suspended in saline, NOT the patient's or donor's own plasma. |
|  |  | Continued on next page |
|  | Follow the steps below for Red Cell only Typing; Donor Unit confirmation, Cord blood testing, etc… |  |
|  | Step | Action |

| # Red Cell Only Testing, Cont’d | 2. | Label 2 test tubes; one with "A", one with "B" and enough information to identify the patient or donor unit. - If performing an ABO confirmation test on group O donor units, label one tube with “A,B” for each group O unit to be confirmed. It is not necessary to test each group O unit with anti-A and Anti-B, just anti-A,B. |
| --- | --- | --- |
|  | 3. | Add one drop of each blood-grouping reagent to the appropriately tubes labeled "A","B" or "AB". |
|  | 4. | Mix the contents of each tube thoroughly and centrifuge on high speed for the appropriate time for the centrifuge being used. (Optimal times are posted on each centrifuge-usually around 15 sec). |
|  | 5. | Gently agitate each tube to resuspend the red cell buttons. - Examine for agglutination using agglutination viewer. - Grade reaction and record results immediately, before the tubes are discarded. Note: Before entering results in the computer or on a manual sheet check the unit label against the unit record (down time record or computer) for unit ABO type, and the donor unit number before confirming a donor unit. |

|  | Follow the following steps for Serum/Plasma testing only; confirmation of FFP or FP, other as directed: |  |
| --- | --- | --- |
|  | Step | Action |
| # Serum/ Plasma Testing | 1. | Label two test tubes, one as "a" and one tube "b" and enough information to identify the patient. |
|  | 2. | Place 2 drops of the patient’s/donor plasma or serum to be tested into each of the tubes. |
|  |  | Continued on next page |

|  | Follow the following steps for Serum/Plasma testing only; confirmation of FFP or FP, other as directed: |  |
| --- | --- | --- |
|  | Step | Action |

| # Serum/ Plasma Testing, Cont’d | 3. |  | - Gently mix the reagent red cells. - Add 1 drop of reagent red cells to the appropriate tube labeled "a" or "b". - Mix contents of each tube thoroughly. |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 4. |  | Centrifuge on high speed for appropriate time (usually 15 sec.- see posted optimum time). |  |  |  |  |  |  |
|  | 5. |  | Gently agitate each tube to resuspend the red cell buttons. - Examine for agglutination using agglutination viewer. - Grade reaction and record results immediately, before the tubes are discarded. Note: These results cannot be entered in the computer so the results need to be recorded on the downtime form and/or the transfusion reaction form. |  |  |  |  |  |  |
| Interpretation Of Patient Testing |  | Observed Reactions of Cells Tested With |  |  |  | Observed Reaction of Serum or Plasma Tested Against |  | Interpretation |  |
|  |  | Anti-A |  | Anti-B | Anti-A,B | A1 Cells | B Cells |  |  |
|  |  | + |  | 0 | + | 0 | + | A |  |
|  |  | 0 |  | + | + | + | 0 | B |  |
|  |  | 0 |  | 0 | 0 | + | + | O |  |
|  |  | + |  | + | + | 0 | 0 | AB* |  |
|  |  | Positive test (+): agglutination of red cells and/or hemolysis Negative test (0): no agglutination of red cells |  |  |  |  |  |  |  |
|  |  | * if AB Rh+, must also have a negative control, e.g patient's red cells tested with 6% albumin |  |  |  |  |  |  |  |

| Interpretation of Donor Confirmation Testing- ABO | Observed Reactions of Cells Tested With |  |  | Interpretation |
| --- | --- | --- | --- | --- |
|  | Anti-A | Anti-B | Anti-A,B |  |
|  | + | 0 |  | A |
|  | 0 | + |  | B |
|  |  |  | 0 | O |
|  | + | + |  | AB* |
|  | Positive test (+): agglutination of red cells and/or hemolysis Negative test (0): no agglutination of red cells |  |  |  |
|  | * if AB Rh+, must also have a negative control, e.g patient's red cells tested with 6% albumin |  |  |  |

### .

| Interpretation of plasma/serum testing-ABO | Observed Reactions of Plasma/Serum Tested With |  | Interpretation |
| --- | --- | --- | --- |
|  | A1 Cells | B Cells |  |
|  | 0 | + | A |
|  | + | 0 | B |
|  | + | + | O |
|  | 0 | 0 | AB |
|  | Positive test (+): agglutination of red cells and/or hemolysis Negative test (0): no agglutination of red cells |  |  |

| Authors: | Jane P. Byrne, CLS | Karen Lavine, MT(ASCP) |
| --- | --- | --- |
|  | Helen To, MT(ASCP) | Virgie Eugenio, MT(ASCP)SBB |
|  | Helen Noriega, MT(ASCP)SBB | Rosalinda Catamisan, CLS, SBB(ASCP)CM, CQA(ASQ) |
|  | Roger AngLee, CLS | Howard D. Ammerman, MT(ASCP) |
|  | Harvey Shulman, MT(ASCP)SBB | Cynthia Calderon, MT(ASCP)SBB |
|  | Angela Varela, MT(ASCP) | Mona Yu, MS,MT(ASCP)SBB |
|  | Billy Leung, CLS |  |

| Distribution | All SCPMG Transfusion Services |
| --- | --- |

| Controlled Documents | Attachment A: ABO Discrepancy Attachment B: Computer Comparison of Historic Results |
| --- | --- |

| Uncontrolled Documents | - Brecher, ME, ed. Technical Manual, American Association of Blood Banks, 14th Edition, 2002 - Manufacturer's Directions.Ortho Diagnostic Systems, Inc. |
| --- | --- |

| Reviewed and approved by: |  |  |
| --- | --- | --- |
| Electronic Signature |  | November 18, 2004 |
| Virginia Vengelen-Tyler, MBA, MT,ASCP(SBB), CQA(ASQ) Regional Blood Bank Compliance Officer |  | Date |
| Signature Collected Electronically |  | January 5, 2011 |
| Adriana A. Bedoya, M.D. FCAP, FASCP Medical Director- San Diego –SA |  | ## Date |
| Signature Collected Electronically |  | ## September 16, 2004 |
| Gary Gochman, MD, Medical Director –Tri-Central SA |  | ## Date |
| Signature Collected Electronically |  | ## September 14, 2010 |
| Jeffrey D. Shiffer, MD. Medical Director –San Fernando Valley SA |  | ## Date |
| Signature Collected Electronically |  | ## September 30, 2004 |
| Joseph Thompson, MD. Medical Director –Metropolitan SA |  | ## Date |
| Signature Collected Electronically |  | ## October 11, 2006 |
| David R. Huebner-Chan, MD. Medical Director –Orange County SA |  | ## Date |
| Signature Collected Electronically |  | ## November 9, 2004 |
| Dong Quach, MD. Medical Director –Inland Empire SA |  | ## Date |
| Signature Collected Electronically |  | ## November 5, 2004 |
| Ramesch Doshi, MD. Medical Director- Tri-Central SA |  |  |
| Signature Collected Electronically |  | ## November 15, 2004 |
| Brian Platz, MD, Medical Director- West Los Angeles |  | ## Date |

DOCUMENT HISTORY PAGE

Effective Date: 	November 18, 2004

| Change type: new, major, minor etc. | Changes Made to Document – Describe | Signature responsible person/Date | Med. Dir. Reviewed/ Date | Lab Manager reviewed/ Date | Date change Imp. |
| --- | --- | --- | --- | --- | --- |
| New | Released 11-19-04, then rescinded 11-20-04 to add new manufacturer recommendation. |  |  |  |  |
| Minor | - Changed procedure to only allow saline suspended cells per manufacturer’s recommendation. - Added cell and serum/plasma testing together. - Did not allow only the accession number on the tube to identify the patient. - Added table for interpretation of donor testing. | 03/16/05 Ginny Tyler | NA | NA |  |
| Minor | Added Irvine and Work Place Safety. No new version needed. | N.A. | N.A. |  |  |
| IMP = Implemented |  |  |  |  |  |

| MasterControl History of Change: |  |  |
| --- | --- | --- |
| Change type: new, major, minor etc. | Version # | Description of Change |
|  |  | - Reformatted - Added table for serum typing - Added phosphate buffered saline |