#### Automated vs Manual Differential Correlation

| Policy | To perform automated versus manual differential correlation semi-annually. Two CLS each to perform a 200 cell manual differential on two different slides. |
| --- | --- |

| ##### Materials and supplies | 5 CBC with differential samples from each instrument for analysis. At minimum, samples should cover the range of values normally encountered. |
| --- | --- |

| ##### Safety or Special Safety Precautions | Refer to the safety manual for general safety requirements. |
| --- | --- |

| ##### Procedure: | - Follow the steps below to …. |  |
| --- | --- | --- |
|  | Step | Action |
|  | 1 | Run 5 automated CBC with differential on each instrument. Print results. |
|  | 2 | Make 3 smears. (2 for CLS plus extra smear) |
|  | 2 | Two CLS to perform manual differential using 200 cells. |

| Method Performance Specifications | The automated WBC differential subsets were rounded off to whole numbers. CLSI (H20)ᶟ recommends that two technologists each perform a 200 cell manual differential on two different slides (total of 400 cells analyzed by each technologist). More dells counted results in a more precise and accurate reference against which to judge the accuracy of the automated method. Automated differential systems analyze thousands of cells. The Rumke Binomial Distribution Table, published by C.L. Rumke in 1978, illustrates that the statistical uncertainty of reference values is a direct function of the number of cells counted. The imprecision of the manual differential is especially pronounced with low numbers (e.g. monocyte, eosinophil, and basophil percent). |  |
| --- | --- | --- |

- Continued on next page

#### Automated vs Manual Differential Correlation

| Expected Values or Reference Range |  |  |
| --- | --- | --- |

#### Continued on next page

#### Automated vs Manual Differential Correlation

| ##### Non-Controlled Documents | The following non-controlled documents support this procedure. Hematology Performance Verification Manual (Aug 2009). Beckman Coulter Inc. Fullerton, California. Retrieved on July 28, 2016. |
| --- | --- |

| ##### Controlled Documents | The following controlled documents support this procedure. |  |
| --- | --- | --- |
|  | Procedure | Number |
|  | Method Validation and Correlation Plan at KP Moreno Valley Lab | MVGP-0560 |
|  | Form | Number |
|  | Automated vs Manual Differential Correlation Form | HEM-0901 |

| ##### Author |  |
| --- | --- |
| - Patricia Jasper, MHA CLS |  |

| ##### Distribution |  |
| --- | --- |
| - Kaiser Permanente Moreno Valley Medical Center Laboratory |  |

End

#### Automated vs Manual Differential Correlation

Reviewed and approved by:

| SIGNATURE | DATE |
| --- | --- |
| Denise Topliff, MHA, CLS Director, Clinical Laboratory – Riverside Medical Center and Moreno Valley Medical Center |  |
| Mark Taira, M.D. Medical Director, Clinical Laboratory – Riverside Medical Center and Moreno Valley Medical Center |  |

Automated vs Manual Differential Correlation

HISTORY PAGE

- Effective Date: __________

| Change type: New, Major, Minor etc. | Changes Made to SOP – describe | Signature responsible person/date | Med. Dir. Reviewed/ Date | Lab Director reviewed/ date | Date change Imp. |
| --- | --- | --- | --- | --- | --- |
| NEW |  |  |  |  |  |