QUALITY CONTROL PROCESS IN HEMATOLOGY DEPARTMENT

| ##### Policy | - All unacceptable control results are resolved before testing patient specimens. - When controls are out of range, a step by step approach must be taken to resolve the out of control results on commercial controls. - It is the responsibility of the Clinical Laboratory Scientist assigned in the department to ensure that quality controls are run and meet the acceptable criteria prior to releasing patient results as outlined in this policy. |
| --- | --- |

| ##### Purpose | To provide a general summary of the quality control process in Hematology Department. Please refer to each specific policy and procedure for testing quality controls mentioned in this policy |
| --- | --- |

| ##### Safety ##### Message | All Quality Control reagent products contain biological source materials from human, avian, reptile and ungulate. Always use appropriate universal precautions while using them to perform a test. |
| --- | --- |

| ##### Definition | "Out-of-Control" - If a result outside of these limits is obtained. "Acceptable limit of error" - the permissible range by which the obtained results may vary from the specified result. Repeated analysis of the control or an unknown specimen should give results within this limit. “Parallel Testing” – Run samples of old and new lot numbers concurrently. Refer to the table. “Abbreviated parallel testing”- When reagents are not available or adequate to complete parallel testing, test the new lot number and verify its range with the manufacturer’s range provided. Run at least 10 samples of old and new lot numbers concurrently. Note: Abbreviated parallel testing must only be practiced in the event old or current lot numbers are unavailable and must not be routinely applied. |
| --- | --- |

________________________________________________________________

Continued on next page

QUALITY CONTROL PROCESS IN HEMATOLOGY DEPARTMENT, Continued

| ##### Interpreting results and course of action: | New lot numbers/shipments of reagents and controls must be verified before placed in operation for patient testing. Packet inserts (if applicable) must be reviewed for any changes. New reagents are tested in parallel with the current reagents in use and results for both must be within the established acceptable range. If values are outside the acceptable range, the data must be carefully evaluated and resolved before any patient results are reported. If the source of error cannot be found, a new control specimen and/or with several of the unknowns chosen at random is re-assayed. If the new control gives an appropriate value and the unknowns check with the previous tests performed the same day, the values are reported; if not, the data is held back and the reagents procedure and instrument are checked one by one until the problem is found and resolved. |
| --- | --- |

| ##### Equipment ##### And ##### Reagents | ### Coulter DxH 800 ### Coulter 6C tri-level control ### Coulter Retic-X tri-level control ### Coulter Latron-X Control #### Coulter Calibrator #### Coulter Reagents (Lyse, Erythrolyse, Retic Stain, Cleaner) - Coulter SMS- stains checked daily with every patient run ### iSED ESR ### iSED ESR bi-level control ### Body Fluid Count ### Hemacytometer /C- Chip and cover-slip ### Streck Cell Chex ### Malaria Binax Now ### Binax Now Malaria Positive Control Kit ### Negative Control (pool of equal volume of EDTA blood samples from 3-5 presume malaria negative individuals. See policy MICRO.SCPMG.001 .v02 BinaxNow Malaria Test-Med Centers.) #### E. WBC Stool- Hematek 3000 stains checked daily with every patient run. |
| --- | --- |

Continued on next page

QUALITY CONTROL PROCESS IN HEMATOLOGY DEPARTMENT, Continued

| Control Run Table: |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Control |  | Instrument |  | Frequency |  | Performed by: |  | Reviewed |
| DxH800 |  |  |  |  |  |  |  |  |
| Latron Control |  | DxH 800 |  | - Latron Control is run after instrument start up on a daily basis. |  | CLS |  | Daily by CLS Monthly by supervisor/designee. |
| Daily Checks |  | SMS |  | - Daily Check is run daily after shutdown/startup |  | CLS |  | Daily by CLS Monthly by supervisor/designee. |
| 6C Control |  | DxH 800 |  | - When analyzer is in use, 3 levels are run at 400 and 2 levels are run at 1200 and 2000. |  | CLS |  | Daily by CLS Monthly by supervisor/designee. |
| Retic C Controls |  | DxH 800 |  | - When analyzer is in use, 3 levels are run at 400 and every 8 hours of patient run. |  | CLS |  | Daily by CLS Monthly by supervisor/designee |
| XB |  | DxH 800 |  | - Monitor results with each batch. |  | CLS |  | As needed by CLS As needed by supervisor/designee |
| Coulter Reagents |  | DxH 800 |  | - Run 1 level of control (preferably abnormal QC) |  | CLS |  | As needed by CLS As needed by supervisor/designee |
| Calibration |  | DxH 800 |  | - Biannual Calibration |  | CLS |  | Biannually by supervisor/designee |
| iSED ESR |  |  |  |  |  |  |  |  |
| iSED ESR I &II | iSED |  | - Daily with every patient run |  | CLS |  | Daily by CLS Monthly by supervisor/designee. |  |
| Body Fluid Count |  |  |  |  |  |  |  |  |
| Streck Cell Chex | Hemacytometer/ C chip |  | - Alternating 1 control of the 2 levels on every 8 hours of patient testing |  | CLS |  | As needed by CLS Monthly by supervisor/designee. |  |
| Diluting Fluid | Hemacytometer/ C chip |  | - Day of use |  | CLS |  | As needed by CLS Monthly by supervisor/designee |  |
| Malaria BinaxNOW |  |  |  |  |  |  |  |  |
| Positive Control Kit | Malaria Binax NOW |  | - Internal Procedural control- prior to release of patient results. - External Control- each day of testing are parallel tested with each new shipment and each new lot of BinaxNOw Malaria Test Kit |  | CLS |  | As needed by CLS Monthly by supervisor/designee |  |
| Negative Control (Pooled) | Malaria Binax NOW |  |  |  | CLS |  | As need by CLS Monthly by supervisor/designee |  |
| WBC Stool |  |  |  |  |  |  |  |  |
| Stains | Hematek 3000 |  | - Quality of stains are checked daily with every patient run |  | CLS |  | As needed by CLS Monthly by supervisor/designee |  |

Continued on next page

QUALITY CONTROL PROCESS IN HEMATOLOGY DEPARTMENT, Continued

| I. Schedule |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Control | Time | Instrument or Equipment |  |
|  | 6C (Normal, Abnormal I, Abnormal II) | 0400 | DxH800-1 & DxH800-2 |  |
|  | 6C (Normal, Abnormal I) | 1200 | DxH800-1 & DxH800-2 |  |
|  | 6C (Normal, Abnormal II) | 2000 | DxH800-1 & DxH800-2 |  |
|  | Retic-X (Level 1, 2, and 3) | 0400 and every 8 hours of patient testing | DxH800-1 & DxH800-2 |  |
|  | Latron CP-X | Daily after shutdown | DxH800-1 & DxH800-2 |  |
|  | iSED I & II | Every 8 hours of patient testing | iSED |  |
|  | Streck Cell-Chex | Every 8 hours of patient testing | Hemacytometer/ C Chip |  |
|  | Stains | Daily with every patient run | Hematek 3000 and DxH SMS |  |

| - II. UNSCHEDULED: |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Parallel Testing: New Lot # or Shipment vs Current Lot #. |  |  |  |  |  |  |  |
| New Lot# | Instrument/Equipment |  |  | Parallel | Abbreviated/ Establish Range |  |  |
| Coulter 6C | DxH800-1 & DxH800-2 |  |  | At least 5 runs | At least 10-20 runs on different days and shifts |  |  |
| Coulter Retic-X | DxH800-1 & DxH800-2 |  |  |  |  |  |  |
| iSED L1 and L2 | iSED ESR |  |  |  |  |  |  |
| Streck Cell Chex | Hemacytometer/ C Chip |  |  |  |  |  |  |
| Coulter Reagents | DxH 800-1 & DxH 800-2 |  |  | 1 level of QC. Preferably abnormal QC | 1 level of QC. Preferably abnormal QC |  |  |
| New Lot# | Test with Reagent/Kit |  |  |  |  |  |  |
| *BinaxNow Pos Control | BinaxNow Malaria Test Kit |  |  | 1 | 1 |  |  |
| BinaxNow Malaria Test Kit | -BinaxNow Positive Control -Freshly prepared negative control |  |  | 1 | 1 |  |  |
| *Parallel testing is required for both new lot and new shipment. * Review change in package insert |  |  |  |  |  |  |  |
| Note: Ideally to establish/verify new lot number control range, parallel testing is conducted over several days. i.e. 5-10 days. Verification run must be within assay range of insert. If reagents are not available to parallel test or if the new control material arrives late, perform abbreviated parallel testing. Test the new lot # and use the manufacturer’s range to establish/verify range. Please see definition of abbreviated parallel testing at the Definition section of this policy. |  |  |  |  |  |  |  |

Continued on next page

QUALITY CONTROL PROCESS IN HEMATOLOGY DEPARTMENT, Continued

|  | III. Evaluating/Troubleshooting Control Run: |
| --- | --- |

|  |  |  |  | - Run Controls |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | - NO |  |  | - Within Acceptable Limits |  |  |  |  |  |  |  |  | - YES |  |  |
|  | - Document results in corrective action log/instrument. Check reagents for volume, outdate (age/expiration date), correct lot number, and signs of contamination. Run controls again. |  |  |  |  |  |  |  |  | - Proceed to testing patient samples. |  |  |  |  |  |
|  | - NO |  | - Reading within acceptable limits |  |  |  |  |  |  |  |  |  |  |  | - YES |
|  | - Document results in corrective action log/instrument. Open a new vial of control. Run controls again |  |  |  |  |  |  |  |  | - Proceed to testing patient samples. |  |  |  |  |  |
|  | - NO |  | - Reading within acceptable limits |  |  |  |  |  |  |  |  |  |  |  | - YES |
|  | - Document results in corrective action log/instrument. Troubleshoot instrument. Document on corrective action log. Run controls again. |  |  |  |  |  |  |  |  | - Proceed to testing patient samples. |  |  |  |  |  |
|  | - NO |  | - Reading within acceptable limits |  |  |  |  |  |  |  |  |  |  |  | - YES |
|  | - Document results in corrective action log/instrument. Get assistance from Regional Instrument Repair department/vendor. Notify supervisor. Use back-up instruments. Run controls on backup instruments if needed. Document results. |  |  |  |  |  |  |  |  | - Proceed to testing patient samples. |  |  |  |  |  |
|  | - NO | - Reading within acceptable limits |  |  |  |  |  |  |  |  |  |  |  |  | - YES |
|  | - Notify lab supervisor. Call instrument repair (8-397-7000) or vendor. Notify house supervisor if there will be a delay in results. |  |  |  |  |  |  |  |  | - Proceed to testing patient samples. |  |  |  |  |  |

Continued on next page

QUALITY CONTROL PROCESS IN HEMATOLOGY DEPARTMENT, Continued

Beckman DXH 800 Accuracy and Precision Studies

|  |  | Policies | Reviewed |
| --- | --- | --- | --- |
| Previous Lot | Print and submit to IQAP | HEM 5000: Beckman IQAP Process | Monthly by supervisor/designee |
| New Lot | Print summary data monthly | HEM 1000: Changing Lot# of Controls - New mean value should be within 5% of the expected range of the package insert - Apply established historical SD range - Range must fall within package insert | Monthly by supervisor/designee |

Instrument Bi Annual Correlation

|  | Analytes | Policy | Reviewed |
| --- | --- | --- | --- |
| DXH | WBC RBC HGB MCV PLATELET MPV Neutrophils Lymphocytes Monocyte Eosinophil Basophil | Run 10 samples that covers reportable range Refer to policy: MVGP-0560 Method Validation and Correlation Plan | Bi annual by supervisor/designee |
| iSED | ESR | Run 10 samples that covers reportable range Refer to policy: MVGP-0560 Method Validation and Correlation Plan | Bi annual by supervisor/designee |
| Auto Diff vs Manual Diff | Neutrophils Lymphocytes Monocytes Eosinophils Basophil | Run 5 samples Refer to policy Hem-900 Auto vs Manual Correlation | Bi annual by supervisor/designee |

Continued on next page

QUALITY CONTROL PROCESS IN HEMATOLOGY DEPARTMENT, Continued

Accuracy and Precision Studies Maximum Variations

(MVGP-0560 Method Validation and Correlation Plan at KP Moreno Valley Lab)

| HEMATOLOGY | Maximum Variations |
| --- | --- |
| WBC | 15% |
| RBC | 6% |
| Hematocrit | 6% |
| Hemoglobin | 7% |
| Platelet Count | 25% |
| Retic | 30% |
| Others | 20% |

| ##### Controlled Policies | QM 5.5.2.100 Validation Policy for FDA Approved Non-waived test MVGP-0560 Method Validation and Correlation Plan Hem-900 Auto vs Manual Correlation HEM 5000: Beckman IQAP Process HEM 1000: Changing Lot# of Controls HEM 2150: ESR using iSED MICRO.SCPMG.001 .v02 BinaxNow Malaria Test-Med Centers |
| --- | --- |

| ##### References | College of American Pathologists: Laboratory Accreditation Manual 2007 July Edition by Francis Sharkey, MD, FCAP |
| --- | --- |

| ##### Author | Carlo E. Punu, MT, CLS Patricia Jasper, MHA CLS |
| --- | --- |

| ##### Distributions | Kaiser Permanente Moreno Valley Medical Center Laboratory |
| --- | --- |

End

QUALITY CONTROL PROCESS IN HEMATOLOGY DEPARTMENT, Continued

Reviewed and approved by:

| SIGNATURE | DATE |
| --- | --- |
| Denise Topliff, MHA, CLS Director, Clinical Laboratory – Riverside Medical Center and Moreno Valley Medical Center |  |
| Mark Taira, M.D. Medical Director, Clinical Laboratory – Riverside Medical Center and Moreno Valley Medical Center |  |

HISTORY PAGE

#### Effective Date: ___________

| Change type: New, major, minor | Changes made to SOP - describe | Signature responsible person/date | Medical Director review/date | Laboratory Director review/date | Date change implemented |
| --- | --- | --- | --- | --- | --- |
| New |  |  |  |  |  |