## Slide 1: Nitrazine pH Test for Amniotic Fluid
Audience: Labor and Deliver & OB/GYN Providers

## Slide 2: Nitrazine pH Test for Amniotic Fluid
Nitrazine paper is used to detect small quantities of amniotic fluid in vaginal secretions. It is used in conjunction with the Fern Test to help detect ruptured membranes. Premature rupture of the membranes before onset of labor may lead to fetal infection and subsequent mortality. The risk is largely eliminated by induction of labor

## Slide 3: Nitrazine pH Test for Amniotic Fluid
Specimen Requirements
Interferences
Quality Control
Procedure
Results
Procedure Notes

## Slide 4: Specimen Requirements
Vaginal secretions from the posterior vaginal pool
Do not touch the pH nitrazine paper to the mucus plug in the cervix
Test sample immediately after collection
Note: Bloody specimens are not acceptable

## Slide 5: Interferences
False positive results may occur from specimen contamination due to heavy vaginal discharge, blood, cervical mucus, semen alkaline urine, and soap.
False negative results may be produced by prolonged rupture of membranes (longer than 24 hours) or when a small volume of fluid has leaked.
Specimen contamination will result in erroneous pH results.

## Slide 6: Quality Control
Quality Control should be performed once a day, on days when patient testing is performed.
Quality Control consists of using two pH buffers
- PH 6.0 Buffer (acidic)
- pH 8.0 Buffer (alkaline)
Record the Quality Control test results on the QC log
Acceptable Criteria: results are within + 0.5 of their designated pH

## Slide 7: Procedure
Tear off a piece of pH Nitrazine paper of the desired length
Apply patient sample to the Nitrazine paper
Immediately match the strip color with the closest color on the dispenser color chart

## Slide 8: Results
Reference Range
A pH equal or greater than 7.0 is considered indicative of the presence of amniotic fluid
A pH of less than 7.0 is considered negative or indeterminate

## Slide 9: Procedural Notes
The Nitrazine test is highly sensitive but not specific
If the Nitrazine and Fern Test are positive, probable membrane rupture has occurred
If the Nitrazine Test is negative, but the Fern Test is positive, there is probable rupture of the membranes due to the Fern Test’s greater specificity
If the Nitrazine Test is positive, but the Fern Test is negative, a second specimen should be collected and tested

## Slide 10: Nitrazine pH Test for Amniotic Fluid
This is the end of the Nitrazine Paper Competency
Please click the “Take Test Button” on the left to complete the Nitrazine Paper competency test and to receive credit