## Slide 1
CELL-DYN Sapphire™ Basic Maintenance
Maintenance Part 2 and Troubleshooting

## Slide 2
CELL-DYN Sapphire™ Maintenance
Dilution Cup Area Cleaning
Use the following procedure to clean the Dilution Cup when needed.
1. Select Analyzer>
2. Select Special Protocols....
3. From the Vacuum and Pressure System area of the SPECIAL PROTOCOLS window, select the Disable for Service button.
4. Wait until a message in the Status Panel area indicates that the system is disabled and prepared for service. (The following status message will appear: [Select Enable for Service to Return System to Ready.])
5. Open the Right and Left Flow Panel Covers and remove the Processor Cover for access to the Sample Processing area.

## Slide 3
CELL-DYN Sapphire™ Maintenance
Dilution Cup Area Cleaning continues…
6. Clean the area around each of the Dilution Cups with a cotton swab moistened with a 10% sodium hypochlorite solution. (See the following figure.)
7. Rinse the area with a water-moistened swab.

## Slide 4
CELL-DYN Sapphire™ Maintenance
Dilution Cup Area Cleaning continues…
8. Replace the Processor Cover, ensure the Open Tube Mode Aspiration Switch wire has been reconnected, and close the Right and Left Flow Panel Covers.
9. From the Vacuum and Pressure System area of the SPECIAL PROTOCOLS window, select the Enable for Service button.
10. Select Close.
11. When the System returns to Ready status, ensure that background results are within specifications before running controls and patient specimens.
12. Verify System performance by running controls before running patient specimens.

## Slide 5
CELL-DYN Sapphire™ Maintenance
Reagent Container Replacement
Reagent Container Replacement in Response to a Reagent Not Detected Message
CAUTION: Do not pour the remaining contents of the old reagent container into the new reagent container. This mixing could contaminate the new reagent.
1. Obtain the new reagent container and confirm that the lot number is same and the expiration date is acceptable. Puncture the cubitainer seal and unscrew the cap on the new reagent.
WARNING: Do not touch or blow on crystallization that can be around the new reagent cap. Crystallized reagent can cause skin irritation.
2. Remove the cap and tubing from the reagent container to be replaced.
3. Insert tubing into the new reagent container and secure the cap.

## Slide 6
CELL-DYN Sapphire™ Maintenance
Reagent Container Replacement continues…
4. Select the Clear Fault button to remove the [Reagent Not
Detected] message.
5. Select Analyzer>
6. Select Reagent Log... from the menu. Select the button corresponding to the reagent being replaced.
7. Select the New Entry... button to display the REAGENT LOG ENTRY window. Enter information required for each blank field.
Lot number and expiration date may be entered either manually or with the Hand-held Bar Code Reader.

## Slide 7
CELL-DYN Sapphire™ Maintenance
Reagent Container Replacement continues…
Entries must be completed for all fields (except <Remarks>). The current date and time, as well as package size, are entered automatically.
8. Select OK to save the entry and close the REAGENT LOG ENTRY window.
9. Select the Close button to exit the REAGENT LOGS window.
10. For all reagents except the Reticulocyte Reagent, run a background count in addition to the auto background already run and ensure that results are within specifications before running specimens.
11. If the reagent lot number has changed, run controls before processing any patient specimens.

## Slide 8
CELL-DYN Sapphire™ Maintenance
Syringe Cleaning/Replacement
Use the following procedures to clean or replace syringes as part of troubleshooting or as directed by Abbott personnel.
Clean and replace syringes one at a time to ensure that each is
replaced in the correct position.
This procedure is used for the following syringes only:
• Reticulocyte Reagent Syringe, on Syringe Drive A1
• Impedance Injection/Metering Syringe, on Syringe Drive C1
• Optical Injection/Metering Syringe, on Syringe Drive C2

## Slide 9
CELL-DYN Sapphire™ Maintenance
Removing Syringe
1. Select the Analyzer>.
2. Select Special Protocols....
3. From the Sample Processor and Autoloader area of the SPECIAL PROTOCOLS window, select the Syringe Remove button.
4. Wait until a message in the Status Panel area indicates that the System is disabled and prepared for syringe removal.
5. Open the Right Flow Panel Cover to gain access to the Syringe Assembly.
6. Remove the syringe by pulling it straight out.
7. To detach the tubing from the syringe, grasp the Luer-Lok fitting at the top of the syringe. Turn the syringe clockwise to release it.
8. Push the syringe plunger in all the way to dispense the residual reagent into an appropriate waste container.

## Slide 10
CELL-DYN Sapphire™ Maintenance
Cleaning Syringe
1. Submerge the syringe in the container of deionized water to remove any crystalline deposits. Aspirate deionized water into the syringe until it is full, but do not remove the plunger from the barrel. Alternately, aspirate and expel the deionized water several times.
CAUTION: Pushing or pulling on the plunger when the syringe is dry
could damage the plunger.
2. Dry the outside of the syringe with a paper towel.

## Slide 11
CELL-DYN Sapphire™ Maintenance
Reinstalling Syringe
1. To reattach the tubing to the syringe, place the Luer-Lok fitting on the syringe and rotate the syringe counter-clockwise until it is finger-tight.
CAUTION: Be careful not to overtighten the fitting or crimp the associated tubing.
2. Reinstall the syringe, making sure that the flat sides of the flange are facing to the sides and the tubing is not twisted.
3. Adjust the plunger depth as necessary. Ensure that the collar at the bottom of the syringe is properly seated in the drive mechanism.
4. Position the flange on the syringe in the slot at the back of the bracket.
5. Repeat the procedure for any other syringes that need to be cleaned.

## Slide 12
CELL-DYN Sapphire™ Maintenance
Reinstalling Syringe continues…
6. From the Sample and Autoloader area of the SPECIAL PROTOCOLS window, select the Syringe Install... button.
7. Ensure that the syringe is installed properly and select the OK button.
8. Close the Flow Panel Cover.
9. Select the Close button to exit the SPECIAL PROTOCOLS window.
10. When the System returns to Ready status, ensure that background results are within specifications before running controls and patient specimens.
11. Verify System performance by running controls before running
patient specimens

## Slide 13
CELL-DYN Sapphire™ Maintenance
Diluent/Sheath Reagent Syringe (included in weekly mait.)
This procedure is used for the following syringes only:
• Diluent/Sheath Syringe, on Syringe Drive B3 Removing Syringe
1. Select the Analyzer>.
2. Select Special Protocols....
3. From the Sample Processor and Autoloader area of the SPECIAL
PROTOCOLS window, select the Syringe Remove button.
4. Wait until a message in the Status Panel area indicates that the System is disabled and prepared for syringe removal.
5. Open the Right Flow Panel Cover to gain access to the Syringe Assembly.

## Slide 14
CELL-DYN Sapphire™ Maintenance
Diluent/Sheath Reagent Syringe continues…
6. To detach the tubing from the syringe, grasp the Luer-Lok fitting at the top of the syringe. Turn the Luer-Lok nut clockwise.
7. The Diluent/Sheath Reagent Syringe is mounted with a swing style clamp. Note the location of the horizontal and vertical ribs on the syringe barrel in the mount. (refer to following figure)

## Slide 15
CELL-DYN Sapphire™ Maintenance
Diluent/Sheath Reagent Syringe continues…
8. To remove the syringe from the mount, gently release the pull tab on the right side of the clamp. The clamp will swing open away from the syringe, allowing removal of the syringe.
9. Grasp the syringe barrel and gently pull the syringe forward.
10. Dispense the residual reagent into an appropriate waste container by pushing the syringe plunger in all the way.

## Slide 16
CELL-DYN Sapphire™ Maintenance
Diluent/Sheath Reagent Syringe continues…
Cleaning Syringe
1. Submerge the syringe in the container of deionized water to remove any crystalline deposits. Aspirate deionized water into the syringe until it is full, but do not remove the plunger from the barrel. Alternately, aspirate and expel the deionized water several times.
CAUTION: Pushing or pulling on the plunger when the syringe is dry
could damage the plunger.
2. Dry the outside of the syringe with a paper towel.

## Slide 17
CELL-DYN Sapphire™ Maintenance
Diluent/Sheath Reagent Syringe continues…
Reinstalling Syringe
1. Adjust the syringe plunger depth as necessary. Ensure that the collar at the bottom of the syringe is properly seated in the drive mechanism.
2. To reattach the tubing to the syringe, place the Luer-Lok fitting in the nut and turn the nut counter-clockwise until it is finger-tight.
CAUTION: Be careful not to overtighten the fitting or crimp the associated tubing.
3. From the Sample and Autoloader area of the SPECIAL PROTOCOLS window, select the Syringe Install... button.
4. Ensure that the syringe is installed properly.
5. Select the OK button.

## Slide 18
CELL-DYN Sapphire™ Maintenance
Diluent/Sheath Reagent Syringe continues…
6. Observe the syringe during the Prime Cycle and Background Count; ensure that the syringe is installed properly and not leaking (liquid dripping from the syringe).
7. Close the Flow Panel Cover.
8. Select the Close button to exit the SPECIAL PROTOCOLS window.
9. When the System returns to Ready status, ensure that background results are within specifications before running controls and patient specimens.
10. Verify System performance by running controls before running patient specimens.

## Slide 19
CELL-DYN Sapphire™ Maintenance
Vent Needle Cleaning
Use the following procedure to automatically clean the Vent Needle and Centering Cone.
1. Select Analyzer>.
2. Select Special Protocols….
3. From Sample Processor and Autoloader area of the SPECIAL PROTOCOLS window, select the Vent Needle Clean button.
4. Select the Close button to exit the SPECIAL PROTOCOLS window.
5. When the System returns to Ready status, run and verify background results are within specifications before running controls and patient specimens.
6. Verify System performance by running controls before running patient specimens.

## Slide 20
CELL-DYN Sapphire™ Maintenance
Vent Assembly Unit Replacement
Abbott recommends replacing the Vent Assembly Unit at least every 5,000 Autoloader cycles to maintain optimal performance. Each laboratory should determine its replacement schedule for the Vent Assembly Unit based on the estimated average number of Autoloader cycles run per day. The following table provides guidelines for establishing such a schedule.
Daily Autoloader Cycles Vent Assembly Unit Schedule | Vent Assembly Unit Schedule
725 or more | *at least every week
350 or more | at least every two weeks
250 or more | at least every three weeks
175 or more | at least every four weeks
* Vent needle replacement is performed with weekly maintenance.

## Slide 21
CELL-DYN Sapphire™ Maintenance
Vent Assembly Unit Replacement
Use the following procedure to replace the Vent Assembly.
1. Confirm that the System is in Autoloader mode. If the Autoloader is processing specimens, allow specimens to complete processing or interrupt it.
2. Select the Analyzer>.
3. Select Special Protocols….
4. From Sample Processor and Autoloader area of the SPECIAL PROTOCOLS window, select the Vent Needle Remove button.
5. Wait until a message in the Status Panel area indicates that the System is disabled. Open the Right and Left Flow Panel Covers and remove the Processor Cover for access to the Sample Processing area.
6. Disconnect the tubing from the Vent Head Drain Connector on the back side of the Vent Head Unit above the Centering Cone.
Be careful not to bend or break the connector when removing tubing. (See the following figure.)

## Slide 22
CELL-DYN Sapphire™ Maintenance
Vent Assembly Unit Replacement
WARNING: Potential Biohazard. The Vent Needle is sharp and potentially contaminated with infectious materials. Exercise caution when handling the Vent Assembly Unit. Do not grasp, lift, or handle the Vent Assembly Unit by the Centering Cone. Do not push up on the Centering Cone when handling the Vent Assembly Unit. This will expose the Vent Needle.

## Slide 23
CELL-DYN Sapphire™ Maintenance
Vent Assembly Unit Replacement
7. Release the Vent Assembly Unit from the Vent Assembly Carriage Pivot. Using your thumb and index finger, depress and hold the Vent Assembly Release Button with your thumb while pressing down on the rear end of the Vent Assembly Unit to release the rear pin from the Pivot with your index finger. (See the following figure.)

## Slide 24
CELL-DYN Sapphire™ Maintenance
Vent Assembly Unit Replacement
8. Unhook the front pin and remove the Vent Assembly Unit from the Vent Assembly Carriage hook by slightly tipping the Vent Assembly Unit down and lifting the unit towards the rear of the Analyzer.(See the following figure.)

## Slide 25
CELL-DYN Sapphire™ Maintenance
Vent Assembly Unit Replacement
9. The Vent Assembly Unit can now be removed from the Sample Processor area and discarded in an appropriately marked, puncture resistant container before treatment and disposal. Do not make any attempts to remove the Vent Needle from the Vent Assembly Unit.
10. Reconnect the tubing to the Vent Head Drain Connector on the back side of the Vent Head Unit above the Centering Cone. Be careful not to bend or break the connector when connecting the tubing. (See the following figure.)

## Slide 26
CELL-DYN Sapphire™ Maintenance
Vent Assembly Unit Replacement
11. Hook the Vent Assembly Unit onto the Vent Assembly Carriage by placing the front pin of the Vent Assembly Unit onto the Vent Assembly Carriage Hook and guide the rear pin up and into the Vent Assembly Carriage Pivot to lock the unit in place. (See the following figure.)

## Slide 27
CELL-DYN Sapphire™ Troubleshooting

## Slide 28
Troubleshooting
Troubleshooting Procedures
Draining/Filling Impedance Transducer (MCV QC out)
Debris and air bubbles can be cleared from the Impedance Transducer by draining and filling it as follows:
1. From the Transducers area of the SPECIAL PROTOCOLS window, select the Aperture Plate Remove button to drain the Impedance Transducer.
2. From the Transducers area of the SPECIAL PROTOCOLS window, select the Aperture Plate Install... Button to fill the Impedance Transducer.
3. If the Aperture Plate was not removed, select the OK button at the warning message appearing on the screen: [Check that component is installed and then select the OK button].
4. Record any appropriate comments in the MAINTENANCE LOG ENTRY window.

## Slide 29
Troubleshooting
Performing Background Counts
Performing Normal and RETC Background Counts
Perform Normal and RETC background counts in the following manner:
1. Select the Run Open Tube... button to open the NEXT OPEN TUBE SETUP window.
2. Select the Specimen Type> button.
3. Select Background from the menu.
4. Select Normal from the Sub Type> menu to check the background counts for the CBC subsystem.
5. To check the background count for the RETC subsystem, select RETC from the Sub Type> menu. CBC and RETC background counts are done when RETC is selected.
6. Select the Close button to exit the NEXT OPEN TUBE SETUP window.
7. Press the aspiration switch to start the background counting cycle.
8. Ensure the background count is within specification before running controls and patient specimens.

## Slide 30
Troubleshooting
Observable Problems
Observable problems are easily observed by the operator but are not detected by the instrument’s sensing mechanisms. The most common observable problems are grouped as follows:
1. Autoloader and Mixing Assembly Problems
• Rack fails to advance
• Mixing Assembly Dropping Tubes or Not Picking up Tubes
• Mixing Assembly Not Depositing Tubes in Rack after mixing
2. Sampling Problems
• Aspiration Probe Bent
• Vent Needle Not Retracting
3. Processing Problems
• Loose Tubing
• Leaks
• Syringe Plunger Showing Salt Deposits
• Obstruction in a Staging Pathway
4. Bar Code Reader Problems
• Autoloader Bar Code Reader Failing to Recognize Tubes in Racks

## Slide 31
Troubleshooting
The System can enter Standby automatically or on demand. After two hours of inactivity, the System runs an Autobackground check. After an additional two hours of inactivity, the System automatically goes into Standby.
Perform the following procedure to manually place the System in Standby.
1. Select Analyzer>
2. Select Special Protocols...
3. Select To Standby. The System will then enter Standby.
  Standby is logged in the System Log with date, time, and operator ID. Once in Standby, the System exercises pinch valves every two hours to unpinch tubing.
The System can be Primed on demand by the operator as described in the
following procedure.
1. Select Analyzer>
2. Select Special Protocols...
3. Select the Prime button to start the Prime cycle.

## Slide 32
Troubleshooting
Prime continues…
4. Select the Close button to close the Special Protocols... window.
The Prime cycle takes approximately 8 minutes. An
autobackground is run at the end of the Prime cycle. When the
cycle is complete, the [Ready] message is displayed in the Status
Panel area and the Ready status light is green.
Check the Data Log to make sure an autobackground was run. If a
Record for an autobackground is not present, the Prime did not
complete correctly, and the Analyzer must be reprimed.
5. Ensure that background counts are within acceptable limits
before running controls and patient specimens

## Slide 33
Troubleshooting
Prime the Analyzer through one of three ways:
    Select Run Loader when the Analyzer status is Not Ready: Standby.
    Select Run Open Tube when the Analyzer status is Not Ready: Standby.
    Access the Special Protocols window and select Prime.
Ensure that background count results are within acceptable limits.
    Check Data Log.
Enter, check, or change operator ID
Check waste container level
    If applicable, empty waste and replace container if necessary.
Check reagent levels in each reagent container
    Replace reagent(s) as needed.

## Slide 34
Troubleshooting
Background too high at the Hgb analyte- put a few drops of methanol into hgb dilution up and run a background.
Background too high PLT optical- disable for service to remove all fluids and clear any bubbles or particulates in the system.
Hgb &Hct do not match on patient samples- ensure auto tube mixhead is functioning and mixing samples well before sampling.
Moving Averages Troubleshooting
One or two batches out, determine if non-randomized specimens have just been analyzed (example: oncology patients).
Three batches out (stop patient testing), review current batch for specimen-related reasons. Run controls. If control results are within acceptable limits, check reagents, maintenance, calibration, and system logs for events that may cause problems.
Three batches out with commercial control also outside limits begin troubleshooting. Document problem and take corrective action. Call service when necessary.

## Slide 35
Troubleshooting
Troubleshooting for Short sample error
Review the following parameters:
HGB, MCHC, RBCi
RBC < 1.00M/uL
HGB < 1.00g/dL
MCHC < 27g/dL or > 40g/dL
• Check sample tube for clots
• Check for clots in sample aspiration tubing and connectors
• Clean RBC impedance aperture plate
• Ensure sample volume at or above 500uL
• If isolated to the closed mode only, check Vent Assembly Unit, Replace if indicated. Perform Zero Park in Special Protocols.
• Verify Probe condition, Replace if bent
Note: Ensure sample tube labels are not too thick on tube or have fallen off the tube this can cause bent probes

## Slide 36
Troubleshooting
Troubleshooting for Short sample error
Check Diluent Syringe, for signs of leaks. Replace syringe and clean up spill with DI water.
  Check peri-pump tubing #2, replace if is tubing is split or you if see the tubing is getting flat.
Check Diluent Syringe for air bubbles and check tubing connection to syringe.
Good

## Slide 37
Troubleshooting
Current Autoloader Rack Failed to Advance
Verify rack is not being blocked. Remove and replace the rack from right-hand side.
Verify the rack position bar code labels are readable.
Verify no obstruction is preventing the mix head assembly from picking up or depositing tubes in rack.
Clean the racks and Autoloader platen if racks are not advancing properly.
Look for signs of blockage here on both sides of the Autoloader (Right Side of Autoloader)