Pathology and Laboratory Medicine Service Department of veterans affairs lexington, ky

Pathology and Laboratory Medicine Safety Manual

3-26-2014

P&LMS Laboratory Safety Manual

| Prepared by: | Lee Ann Speaks |  | Authorized by: | Dr. Bonnie Mitchell |
| --- | --- | --- | --- | --- |
| Title: | Lab Safety Officer |  | Title: | Chief PLMS |
| Date Prepared: | 11/29/06 |  | Date Authorized: | 11/26/06 |

| Review Authority: | Chief PLMS |
| --- | --- |
| Review Authority: | Lab Safety Officer |

| Reviewed (Date) | ###### By |  | Revised (Date) | ###### By |
| --- | --- | --- | --- | --- |
| 11/29/06 | Lee Ann Speaks |  | 11/29/06 | Lee Ann Speaks |
| 11/29/06 | Dr. Bonnie Mitchell |  | 12/14/07 | Lee Ann Speaks |
| 12/14/07 | Lee Ann Speaks |  | 01/16/09 | Lee Ann Speaks |
| 12/21/07 | Dr. Bonnie Mitchell |  | 12/14/09 | Vicki B Richardson |
| 12/9/08 | Lee Ann Speaks |  | 11/18/10 | Vicki B Richardson |
| 12/9/08 | Dr. Bonnie Mitchell |  | 3/4/11 | Vicki B Richardson |
| 01/16/09 | Lee Ann Speaks |  | 3/21/2012 | Vicki B. Richardson |
| 01/16/09 | Dr. Bonnie Mitchell |  | 11/19/2012 | Vicki B. Richardson |
| 12/14/09 | Vicki B Richardson |  | 1/16/2014 | Vicki B. Richardson |
| 12/18/09 | Dr. Bonnie Mitchell |  | 3/26/2014 | Vicki B. Richardson |
| 11/18/10 | Vicki B Richardson |  |  |  |
| 11/18/10 | Dr. Bonnie Mitchell |  |  |  |
| 3/4/11 | Vicki B Richardson |  |  |  |
| 3/4/11 | Dr. Bonnie Mitchell |  |  |  |
| 3/21/2012 | Vicki B. Richardson |  |  |  |
| 3/21/2012 | Dr. Luis Samayoa |  |  |  |
| 8/27/2012 | Dr. David Hunt |  |  |  |
| 11/19/2012 | Dr. David Hunt |  |  |  |
| 1/16/2014 | Dr. David Hunt |  |  |  |
| 3/26/2014 | Dr. David Hunt |  |  |  |

Table of Contents

I.	Responsibilities	5

II.	Procedures	7

A.	General Laboratory Safety Requirements	7

B.	Protective Apparel and Equipment	8

C.	Biohazardous Waste Disposal	10

D.	Bloodborne Pathogen Exposure Control Plan	11

E.	Specimen Transport	12

F.	Biological Spills and Accidents	13

G.	Viral Exposure	13

H.	Latex Allergy Protection	14

I.	Ergonomics	19

J.	Noise	23

K.	Housekeeping	23

L.	Maintenance	24

M.	Inspections	24

N.	Passageways	24

O.	Recordkeeping	24

P.	Medical Recordkeeping	25

Q.	Signs	26

R.	Employee information and training	26

S.	Electrical safety	27

T.	Fire Safety	28

U.	Ultraviolet Light Safety	29

V.	Guidelines for Handling Patient Emergencies in Outpatient Phlebotomy Areas	30

W.	Laboratory Safety Committee	30

III.	References	31

Attachment A:   BLOODBORNE PATHOGEN EXPOSURE PREVENTION PLAN   Memorandum No 111-01	32

Attachment B:  BLOODBORNE PATHOGEN ECPOSURE MANAGEMENT PLAN Memorandum No. 111-02	62

Attachment C:  ERGONOMICS PROGRAM MEMORANDUM NO. 138-01	68

Attachment D:  MEDICAL EQUIPMENT MANAGEMENT PROGRAM Memorandum No. 138-09	81

Attachment E:  FORMALDEHYDE SOP-ENG-138-12	90

Attachment F:  NOISE CONTROL AND HEARING CONSERVATION PROGRAM  SOP-ENG-138-02	97

Attachment G:  Biomedical Engineering Annual Checklist	112

Attachment H:  BIOMEDICAL NEW EQUIPMENT CHECK IN SHEET	114

Attachment I:  SERVICE SAFETY TEAM INSPECTION PROGRAM CHECKLIST	115

Attachment J:  HAND HYGIENE FOR HEALTH CARE WORKERS Memorandum No. 111-09	122

Attachment K: Ergonomic Symptoms Survey	129

Attachment L: Instructions for ASISTS Incident Reporting	130

Attachment M: P&LMS New Employee Safety Walkthrough Checklist	133

# Responsibilities

#### The Chief Pathologist, Pathology & Laboratory Medicine Service and the Laboratory Manager are responsible for:

- providing a safe workplace environment for the employees and volunteers within the service

- enforcing the medical center/hospital policies and all relevant procedures regarding employee safety, and volunteers performing tasks within the PLMS

- ensuring that the requirements of various accrediting agencies are fulfilled

- providing the designated laboratory safety officer with necessary support in obtaining resources and personnel to implement the Infection Control and Chemical Hygiene Plan

- ensuring that corrective action is taken for any problems/deficiencies identified

- monitoring safety compliance within the laboratory

#### The designated Laboratory Safety Officer is responsible for:

- developing policies and procedures for operations that involve hazardous chemicals

- reviewing and evaluating the Infection Control Plan and Laboratory Safety Manual on an annual basis

- educating and training employees in procedures relating laboratory safety

- monitoring and enforcing compliance with PLMS and hospital/medical center policies

- documenting compliance with safe work practices via the Quality Improvement Program of PLMS

- reviewing all accident reports and ensures proper documentation

- serving as Chairman of the Laboratory Safety Subcommittee.  Performs annual evaluation of the effectiveness of the Laboratory Safety Infection Control and Laboratory Safety Plan by reviewing accidents, checking for patterns of accidents, prevention efforts, and presenting this information in the form of graphical analysis as related to the employee’s work section and type of accident.

- thoroughly investigating incidents or unsafe conditions concerning laboratory safety and assuring that prompt action is taken to prevent recurrence

- ensuring training of employees with regard to laboratory safety in general and observation of Medical Center safety policies and procedures

- performing monthly inspection of the Laboratory Service and assures compliance of the Safety Program

#### The Section Supervisor/Lead Technologist is responsible for:

- ensuring that all laboratory work practices comply with service specific policies, as detailed in this plan

- enforcing safe work practices and the use of personal protective equipment within the assigned section

- displaying the appropriate signage

- assisting the designated Laboratory Safety Officer in ensuring that all employees receive adequate training and that this training is documented in the employee training record

- implementing appropriate corrective action when problems or deficiencies are identified

- advising employees of safe medical treatment or first aid should the need arise

- investigating all reported accidents within the assigned section

- ensuring completion of VA Form 2162, "Report of Accident" and the CA-1 U.S. Department of Labor form for each employee reported accident (if employee chooses)

- coordinating the entry of employee reported accident data into the ASISTS program

#### Each employee is responsible for:

- following established policies and procedures designed to protect employees

- recognizing potential safety hazards and reporting them to appropriate supervisors in a timely manner

- reporting all incidents or accidents so that steps may be taken to prevent their recurrence

- completing mandatory training within the established time frame

- reviewing the P&LMS's Infection Control and Laboratory Safety Plan on an annual basis.

# Procedures

General Laboratory Safety Requirements

- Eating, drinking, application of cosmetics (including lip balm), or manipulation of contact lenses are strictly prohibited in the technical work areas (including phlebotomy) and are only permitted in designated clean areas.

- Smoking or use of any tobacco substance is strictly prohibited in the technical work areas (including phlebotomy) and is only permitted in designated smoking areas.

- Food and beverages are not permitted in analytical/laboratory area refrigerators.

- Hands should be thoroughly washed before leaving laboratory technical area.

- Mouth pipetting is prohibited.

- The use of glass capillary tubes has been discontinued and is strictly prohibited for specimen collection and specimen handling.

- Personal Protective Equipment (PPE), such as a fluid resistant lab coat/apron, closed toed shoes, and gloves are required when working in the analytical areas.  Eye protection (safety goggles, face shields) will be used when a splash hazard exists.

- Hair should be secured back and off the shoulders in such a manner to prevent contact with contaminated materials or surfaces and moving instrumentation.

- Avoid practical jokes or other behavior which might confuse, startle, or distract other laboratory personnel.

- Keep work area clean and uncluttered, with chemicals and equipment being properly labeled and stored.  Clean and disinfect the work area on completion of task or at the end of each day.

- Glassware that is chipped or broken should be discarded in broken glass receptacles.

- Do not centrifuge uncovered specimens.

- All employees must be aware of the following when working with any chemicals:

- The chemical hazards as determined from the Safety Data Sheet (SDS), labels, and other appropriate references.

- SDS files are available to every employee and should be reviewed before using chemicals with which one is unfamiliar.

- The appropriate safeguards when using a chemical, including personal protective equipment.

- The location of the nearest safety shower, eyewash station, fire extinguisher, and fire alarm

- Knowledge of emergency phone numbers.

- The location and the proper use of the nearest emergency spill control station.

- How and where to properly store the chemical when not in use.

- The proper methods for transporting chemicals within the facility including the use of secondary containment procedures to avoid spills.

- The importance of purchasing minimum amounts of hazardous chemicals.

- Appropriate procedures for emergencies, including evacuation routes, spill cleanup procedures, and proper waste disposal.

- Sniffing of reagents can cause injury to an employee's mucous membranes and is prohibited.

- Great care shall be used when performing dilutions.

- Always add acids or alkali to water.

- Allow acid/alkali to run down the side of the container slowly and mix by gentle rotation.

- Safety equipment, such as safety showers, eyewashes and appropriate fire extinguishers are located throughout the laboratory and are tested periodically.

- Operating fume hoods shall be used for the handling of hazardous materials.

- Personnel shall notify their immediate Supervisor/Lead Technologist, or the Laboratory Safety Officer, in the event of a hood failure.

- When an AFB smear is reported as positive, the physician and Infection Control section are notified.  Employee notification of potential exposure to this patient is subsequently initiated by the Infection Control Section.

Protective Apparel and Equipment

Universal or standard precautions must be used when handling all blood and body fluid specimens.  The term "universal precautions" refers to the concept of bloodborne disease control requiring all human blood and other potentially infectious materials to be treated as if infectious for HIV, HBV, HCV or other bloodborne pathogens, regardless of the perceived "low risk" status of a patient or patient population.  Alternative concepts in infection control are called Body Substance Isolation (BSI) and Standard Precautions.  These latter terms define all body fluids and substances as infectious.  All health care workers must routinely use appropriate barrier precautions (PPE) to prevent skin and mucous membrane exposure when contact with blood or other body fluids is anticipated. Personal Protective Equipment (PPE) includes gloves, goggles, face shields, aprons, fluid impervious lab coats/gowns, and masks.

- Use of  Personal Protection Equipment

- Eye

- Chemical safety goggles and/or face shields, rather than safety glasses shall be used when pouring any hazardous chemicals or hazardous waste as they provide the best protection against splashes.

- Protective face shields (either personal or stationary) must be utilized in all areas where splashing of hazardous substances is reasonably anticipated.

- Chemistry

  - Processing station

  - Fusion workstation

- Histology

  - When any specimen requires opening

  - During grossing and grossing assistance of histology specimens

  - During grossing of frozen section specimens

- Blood Bank

  - While working at the bench with the specimen (uncapping specimen, decanting waste). Reading the tubes under the shield is not required.

  - When emptying the waste boxes in the sink.

- Cytology

  - Cytology processing station

- Hematology

- Osmolality bench

- Special Hematology

- Manual Hematology

- When pouring specimens into IRIS tubes in Urinalysis

- Sysmex bench if popping the sample cap is required

- Microbiology/Serology

  - Blood culture set up counter

  - Serology processing counter

  - PCR set up counter

- Phlebotomy

  - Specimen processing stations

- Special Reference Laboratory

  - IFA stations

  - Aptus specimen processing area in A139

  - Aptus benches in B102

  - Liaison Bench

  - Complement Fixation bench

  - Western Blot bench

  - HIV bench

  - PCR bench in B111

  - Mail sorting bench

- Protective eyewear should be easy to clean and disinfect.

- For those employees who wear glasses, face protection must fit over the glasses and cover mucous membranes (eyes and mouth).

- Aprons and Lab Coats

- Lab coats must be impermeable to fluids so that no blood or body fluid will soak through to the clothing or skin of the employee.

- Lab coats are required when handling all patient samples

- The proper technique for putting on (donning) a lab coat is as follows:

  - Insert one arm into a sleeve

  - Insert the other arm into the other sleeve

  - Adjust the shoulders and sleeves until comfortable

  - Completely button front of coat

- The proper technique for taking off (doffing) a lab coat is as follows:

  - Unbutton (unsnap)coat

  - Put hands behind your back

  - Slide fingers from one hand under the cuff of the opposite arm’s sleeve and pull sleeve off.

  - Repeat for the other sleeve.

- Lab coats should be fully buttoned with sleeves rolled down.

- Don't wear lab coats in public places, such as offices, lunch rooms, lounge areas, or elsewhere outside the laboratory, as they can transfer hazardous materials and contaminate these areas.

- If a safety apron is required for chemical handling, it must be rubberized and acid resistant.

- Shoes

- Shoes must be closed-toe.

- Shoes should have non-slip soles

- Shoes must provide adequate protection against blood, body fluids or chemicals passing through to the skin.

- Gloves

- Nitrile gloves are provided for all employees.  Latex gloves are no longer in use.

- Use properly fitting gloves.

- Employees are required to wear gloves when handling all patient specimens.

- Employees are required to wear gloves when performing phlebotomy.

- All gloves should be inspected before and after each use to ensure they are not torn, punctured or show signs of degradation.  If the integrity of the gloves is in question, replace immediately.

- Gloves are not to be torn or cut for any reason.  This damages the integrity of the protection.

- Disposable gloves are to be changed as soon as possible after contamination and between each patient.

- To remove gloves without skin exposure, remove the first glove by grasping the cuff and peeling the glove off the hand so that the glove is inside out. Repeat this process with the second hand, touching the inside of the glove cuff, rather than the outside.

- If your hands are visibly soiled or contaminated with blood or body fluids after glove removal,  cleanse your hands by washing them with soap and water for at least 15 seconds.

- If your hands are not visibly soiled, sanitize them with an alcohol-based waterless agent after glove removal.

- Do not wear gloves outside of the lab. When transporting hazardous materials between labs, use secondary containers that can be carried without gloves.

- Special acid resistant gloves will be provided in all areas utilizing acids.

- Disposable gloves should never be washed or disinfected and reused.

- Information, mandatory training and latex free (Nitrile) gloves will be provided to each employee annually.

Biohazardous Waste Disposal

- Non-hazardous waste should not be placed in the biohazardous waste receptacles.  Patient labels are the only non-hazardous waste that must be placed in the biohazardous waste. The incinerating process destroys the patient’s protected health information printed on the thermal labels.

- Urine waste may be discarded into the sanitary sewerage system.

- Biohazardous waste (including tissue and sharps disposal containers) will be placed in appropriately designated/labeled containers.  The double bagged waste and/or sharps disposal containers are to be placed in the large red biohazardous waste containers in room B113.  The large red biohazardous waste containers must be  labeled as “Biohazardous”, must be leak proof, and must have solid, tight-fitting covers that are applied before transport from the laboratory.  They are furnished by Environmental Management Service (EMS) and/or a contractor.  EMS personnel pick up the containers for incineration.

- Used glass slides, tissues and tissue embedded paraffin blocks are discarded in specific biohazardous lidded containers that have been provided by Environmental Management Service.  EMS personnel pick up the containers for incineration.

- Sterile syringes, needles, lancets, or other blood-letting devices ("sharps") that are capable of transmitting infection are to be used once only.  All waste sharps must be discarded in puncture-resistant containers that are easily accessible, located in areas where needles are commonly used, and properly labeled to warn handlers of the potential hazard.

Bloodborne Pathogen Exposure Control Plan

(see Attachment A:   BLOODBORNE PATHOGEN EXPOSURE PREVENTION PLAN   Memorandum No 111-01)

- Universal or standard precautions must be used when handling all blood and body fluid specimens.  The term "universal precautions" refers to a concept of bloodborne disease control requiring all human blood and other potentially infectious materials to be treated as if infectious for HIV, HBV, HCV or other bloodborne pathogens, regardless of the perceived "low risk" status of a patient or patient population.    All health care workers must routinely use appropriate barrier precautions to prevent skin and mucous membrane exposure when contact with blood or other body fluids is anticipated.

- The Lexington VA Bloodborne Pathogen Exposure Prevention Plan can be found on the Lexington Intranet http://vaww.lexington.med.va.gov/docs/memo.aspx or Attachement A at the end of this document.

- Exposures occur through:

- Needle sticks or cuts from sharp instruments contaminated with an infected patient’s blood

- Through contact of the eye, nose, mouth, skin, or mucous membranes with a patient’s blood or body fluids

- Preventive practices include:

  - Standard Precautions/Personal Protective Equipment (PPE)

- Gloves – protect hands

- Lab Coats/Aprons – protect skin and clothing

- Masks/Respirators – protect mouth/nose

- Goggles – protect eyes

- Face shields – protect face, mouth, nose and eyes

  - Hand Hygiene

See Attachment J for Hand Hygiene Memorandum No. 111-09

- Perform hand hygiene immediately after removing PPE.

- Use an alcohol-based hand rub or antimicrobial soap and water to routinely decontaminate their hands before and after having direct contact with a patient. A single act of using an alcohol-based hand rub or an antimicrobial soap and water after one patient and before the next patient suffices to decontaminate a HCW’s hands, if the hands are not subject to recontamination in-between patients (as in talking on the telephone, handling objects, pushing elevator buttons, turning doorknobs, etc).

- Procedures for hand hygiene are as follows:

Note: follow manufacturer’s instructions for all products.

| Soap and Water | Alcohol based hand rub |
| --- | --- |
| 1. Wet hands with water. (NOTE: Hot water increases dermatitis risk) | 1. Apply recommended amount of product to the palm. |
| 2. Apply soap. | 2. Rub hands together to cover all surfaces of hands and fingers. |
| 3. Rub all surfaces of hand, cuticles and between fingers for at least 15 seconds. | 3. Rub until hands are dry. |
| 4. Wash around each wrist and wash under rings. | 4. Do not touch electrical sockets or switches or engage in activities that may generate static electricity until hands are dry. |
| 5. Rinse hands and under rings. Keep hands and clothing away from sink. | 5. Do not use alcohol rubs with patients that have C. difficile. |
| 6. Dry hands by patting dry with a disposable towel and use towel to turn the water off. |  |

  - Engineering Controls

- using medical devices with safety features, i.e. self-blunting needles

  - the recapping, purposeful bending, breaking, removing from disposable syringes, or other manual manipulations of needles is prohibited

  - Resheathing instruments or self-sheathing needles may be used to prevent recapping of needles by hand.

- access to and use of sharps disposal containers

- use of hazard communication – warning labels and signs

- proper waste disposal

- Post-Exposure Procedures include:

- Cleaning

- wash needle sticks and cuts with soap and water

- flush splashes to the nose, mouth, or skin with water

- irrigate eyes with clean water using an eyewash station

- Reporting

- report accidents immediately to supervisor, lab safety officer and employee health (ext. 5270)

- seek appropriate medical care

- supervisor completes ASISTS

- possible post exposure treatment

Specimen Transport

- VAMC Leestown PLMS samples are transported to VAMC Cooper Drive via the shuttle.

- Blood and body fluid samples are to be transported as “Biological Substance, Category B”

- The shuttle drivers will complete the “Safety Guidelines for VAMC Bus Drivers Pertaining to Transporting Laboratory Samples” PowerPoint presentation annually.

- All samples will be collected in a primary container that is watertight and leak proof.  A couple of examples of primary containers are vacutainer tubes or a urine cups.

- The primary containers will be placed into a secondary container (biohazard zip-lock bag) that is watertight.  The secondary container must be securely sealed.  Up to 100 ml may be in the secondary container.  Absorbent material capable of absorbing the entire contents must be placed in the secondary container along with the samples.  Our absorbent material will be the Super Absorbent Wipes (the orange squares with fibrous backing) or paper towels.  Place the patient labels in the outside pocket of the biohazard bags.

- All secondary containers will be put into an outer container or packaging during transport.  The outer container should be made of strong material that can be cleansed and disinfected.  Contact information shall be on the outer container in case of a problem with the container or its contents.  Our outer containers are latched, leak proof coolers.

- Adequate cushioning absorbent materials shall be placed inside the cooler to absorb shocks during transport and absorb any spillage should it occur.

- After delivery and emptying of the coolers to Cooper Drive PLMS, the cooler shall be wiped down with a Sani-Cloth wipe before the shuttle driver picks up the cooler for transport back to Leestown.

Biological Spills and Accidents

- Large biological spills with significant aerosol formation

- Evacuate area immediately.

- Notify Safety Office immediately.

- Small biological spills with aerosol formation

- Place paper towels over spill.

- Flood with the hospital approved disinfectant, leave for 10 minutes.

- Wear gloves to pick up contaminated materials.  Dispose in biohazard bags.

- Small biological spills with broken glass

- Use UNI-SAFE (located on walls in multiple areas throughout the laboratory and phlebotomy areas) to pour around and over the spill to absorb it.  The UNI-SAFE will cause the spill to clump.

- Use provided scoops to scoop up spill with glass.

- Dispose of glass laden spill in a biohazard sharps container.

- Flood spill area with the hospital approved disinfectant, leave for 10 minutes.

- Wipe area with paper towels and dispose of in biohazard bags.

4.   Spill Evaluation

- All large spills and near spills will be carefully analyzed by Laboratory Service Safety Officer, Safety Subcommittee and Medical Center Safety Committee.

- Following analysis, results will be given to all personnel to prevent further incidents.

Viral Exposure

- The Bloodborne Pathogen Exposure Prevention Plan Memorandum No 111-01 (Attachment A) and The Bloodborne Pathogen Exposure Management Plan Memorandum No. 111-02 (Attachment B) are our plans for follow-up procedures after possible and known percutaneous, mucous membrane or abraded skin exposure to HIV, HBV or HCV.  Electronic copies of the prevention plan, management plan and the OSHA standard are accessible to all employees on the Lexington Intranet and on the Public Drive/+Infection Control folder.

- Individuals with percutaneous, conjunctival, non-intact skin or mucus membrane contact with patient blood/body fluids are encouraged to report the exposure and go immediately or as soon as possible to Occupational Health (OH) or the Emergency Department (ED) for post exposure evaluation and follow-up. Further management will depend on employee vs. non-employee status. Abnormal testing results of the source or exposed person are communicated to that individual by the ordering practitioner/surrogate.

Latex Allergy Protection

Note:  At this facility, we do not use latex gloves.  The College of American Pathologists (CAP) requires a documented program to protect personnel and patients from allergic reactions from exposures to natural rubber latex in gloves and other products.

Under OSHA’s Bloodborne Pathogens standard (29 CFR 1910.1030), your employer must ensure that gloves and other appropriate personal protective equipment (PPE)  are accessible at your worksite or issued to you. However, alternatives are available and must be readily accessible if you are allergic to latex-containing PPE. Glove liners or powder- free gloves can be used to reduce exposure or latex-free gloves can be used to eliminate exposure.Latex gloves have proved effective in preventing transmission of many infectious diseases to health care workers. But for some workers, exposures to latex may result in skin rashes; hives; flushing; itching; nasal, eye, or sinus symptoms; asthma; and (rarely) shock.

The National Institute for Occupational Safety and Health (NIOSH) has made several recommendations to minimize latex-related health problems in workers while protecting them from infectious materials.  These recommendations include reducing exposures, using appropriate work practices, training and educating workers, monitoring symptoms, and substituting nonlatex products when appropriate.

- In this document, the term "latex" refers to natural rubber latex and includes products made from dry natural rubber. Natural rubber latex is the product manufactured from a milky fluid derived mainly from the rubber tree, Hevea brasiliensis.

- Workers exposed to latex gloves and other products containing natural rubber latex may develop allergic reactions such as skin rashes; hives; nasal, eye, or sinus symptoms; asthma; and (rarely) shock.

- If you develop symptoms of latex allergy, avoid direct contact with latex gloves and products until you can see a physician experienced in treating latex allergy.

- Products Containing Latex

The following are examples of products that may contain latex:

- Emergency Equipment

- Blood pressure cuffs

- Stethoscopes

- Disposable gloves

- Oral and nasal airways

- Endotracheal tubes

- Tourniquets

- Intravenous tubing

- Syringes

- Electrode pads

- Personal Protective Equipment

- Gloves

- Surgical masks

- Goggles

- Respirators

- Rubber aprons

- Office Supplies

- Rubber bands

- Erasers

- Hospital Supplies

- Anesthesia masks

- Catheters

- Wound drains

- Injection ports

- Rubber tops of multidose vials

- Dental dams

- Household Objects

- Automobile tires

- Motorcycle and bicycle handgrips

- Carpeting

- Swimming goggles

- Racquet handles

- Shoe soles

- Expandable fabric (waistbands)

- Dishwashing gloves

- Hot water bottles

- Condoms

- Diaphragms

- Balloons

- Pacifiers

- Baby bottle nipples

- Latex in the Workplace

Workers in the health care industry (physicians, nurses, dentists, Medical technologists, etc.) are at risk for developing latex allergy because they use latex gloves frequently.

- Types of Reactions to Latex

Three types of reactions can occur in persons using latex products:

- Irritant Contact Dermatitis

The most common reaction to latex products is irritant contact dermatitis—the development of dry, itchy, irritated areas on the skin, usually the hands. This reaction is caused by skin irritation from using gloves and possibly by exposure to other workplace products and chemicals. The reaction can also result from repeated hand washing and drying, incomplete hand drying, use of cleaners and sanitizers, and exposure to powders added to the gloves. Irritant contact dermatitis is not a true allergy.

- Chemical Sensitivity Dermatitis

Allergic contact dermatitis (delayed hypersensitivity, also sometimes called chemical sensitivity dermatitis) results from exposure to chemicals added to latex during harvesting, processing, or manufacturing. These chemicals can cause skin reactions similar to those caused by poison ivy. As with poison ivy, the rash usually begins 24 to 48 hours after contact and may progress to oozing skin blisters or spread away from the area of skin touched by the latex.

- Latex Allergy

Latex allergy (immediate hypersensitivity) can be a more serious reaction to latex than irritant contact dermatitis or allergic contact dermatitis. Certain proteins in latex may cause sensitization (positive blood or skin test, with or without symptoms). Although the amount of exposure needed to cause sensitization or symptoms is not known, exposures at even very low levels can trigger allergic reactions in some sensitized individuals.

Reactions usually begin within minutes of exposure to latex, but they can occur hours later and can produce various symptoms. Mild reactions to latex involve skin redness, hives, or itching. More severe reactions may involve respiratory symptoms such as runny nose, sneezing, itchy eyes, scratchy throat, and asthma (difficult breathing, coughing spells, and wheezing). Rarely, shock may occur; but a life-threatening reaction is seldom the first sign of latex allergy. Such reactions are similar to those seen in some allergic persons after a bee sting.

- Levels and Routes of Exposure

- The proteins responsible for latex allergies have been shown to fasten to powder that are used on some latex gloves. When powdered gloves are worn, more latex protein reaches the skin. Also, when gloves are changed, latex protein/powder particles get into the air, where they can be inhaled and contact body membranes. In contrast, work areas where only powder-free gloves are used show low levels or undetectable amounts of the allergy-causing proteins.

- Wearing latex gloves during episodes of hand dermatitis may increase skin exposure and the risk of developing latex allergy. The risk of progression from skin rash to more serious reactions is unknown. However, a skin rash may be the first sign that a worker has become allergic to latex and that more serious reactions could occur with continuing exposure.

- Who Is at Risk?

- Workers with ongoing latex exposure are at risk for developing latex allergy. Such workers include health care workers (physicians, nurses, aides, dentists, dental hygienists, operating room employees, laboratory technicians, and hospital housekeeping personnel) who frequently use latex gloves and other latex-containing medical supplies. Workers who use latex gloves less frequently (law enforcement personnel, ambulance attendants, funeral-home workers, fire fighters, painters, gardeners, food service workers, and housekeeping personnel) may also develop latex allergy. Workers in factories where latex products are manufactured or used can also be affected.

- Atopic individuals (persons with a tendency to have multiple allergic conditions) are at increased risk for developing latex allergy. Latex allergy is also associated with allergies to certain foods especially avocado, potato, banana, tomato, chestnuts, kiwi fruit, and papaya

- People with spina bifida are also at increased risk for latex allergy.

- Diagnosing Latex Allergy

- Latex allergy should be suspected in anyone who develops certain symptoms after latex exposure, including nasal, eye, or sinus irritation; hives; shortness of breath; coughing; wheezing; or unexplained shock. Any exposed worker who experiences these symptoms should be evaluated by a physician, since further exposure could result in a serious allergic reaction. A diagnosis is made by using the results of a medical history, physical examination, and tests.

- Taking a complete medical history is the first step in diagnosing latex allergy. In addition, blood tests approved by the Food and Drug Administration (FDA) are available to detect latex antibodies. Other diagnostic tools include a standardized glove-use test or skin tests that involve scratching or pricking the skin through a drop of liquid containing latex proteins. A positive reaction is shown by itching, swelling or redness at the test site. However, no FDA-approved materials are yet available to use in skin testing for latex allergy. Skin testing and glove-use tests should be performed only at medical centers with staff that are experienced and equipped to handle severe reactions.

- Testing is also available to diagnose allergic contact dermatitis. In this FDA-approved test, a special patch containing latex additives is applied to the skin and checked over several days. A positive reaction is shown by itching, redness, swelling, or blistering where the patch covered the skin.

- Occasionally, tests may fail to confirm a worker who has a true allergy to latex, or tests may suggest latex allergy in a worker with no clinical symptoms. Therefore, test results must be evaluated by a knowledgeable physician.

- Treating Latex Allergy

Once a worker becomes allergic to latex, special precautions are needed to prevent exposures during work as well as during medical or dental care. Certain medications may reduce the allergy symptoms, but complete latex avoidance (though quite difficult) is the most effective approach.

- Conclusions

Latex allergy in the workplace can result in potentially serious health problems for workers, who are often unaware of the risk of latex exposure. Such health problems can be minimized or prevented by following the recommendations outlined in this segment.

- Responsibility of Employees

Employees should take the following steps to protect themselves from latex exposure and allergy in the workplace:

- Use nonlatex gloves.

- Appropriate barrier protection is necessary when handling infectious. If you choose latex gloves, use powder-free gloves with reduced protein content:

  - Such gloves reduce exposures to latex protein and thus reduce the risk of latex allergy (though symptoms may still occur in some workers).

  - So-called hypoallergenic latex gloves do not reduce the risk of latex allergy. However, they may reduce reactions to chemical additives in the latex (allergic contact dermatitis).

- Use appropriate work practices to reduce the chance of reactions to latex:

- When wearing latex gloves, do not use oil-based hand creams or lotions (which can cause glove deterioration) unless they have been shown to reduce latex-related problems and maintain glove barrier protection.

- After removing latex gloves, wash hands with a mild soap and dry thoroughly.

- Use good housekeeping practices to remove latex-containing dust from the workplace:

- Frequently clean areas contaminated with latex dust (upholstery, carpets, and ventilation ducts).

- Frequently change ventilation filters and vacuum bags used in latex-contaminated areas.

- If you develop symptoms of latex allergy, avoid direct contact with latex gloves and other latex-containing products until you can see a physician experienced in treating latex allergy.

- If you have latex allergy, consult your physician regarding the following precautions:

- Avoid contact with latex gloves and other latex-containing products.

- Avoid areas where you might inhale the powder from latex gloves worn by other workers.

- Tell your employer and your health care providers (physicians, nurses, dentists, etc.) that you have latex allergy.

- Wear a medical alert bracelet.

- Carefully follow your physician's instructions for dealing with allergic reactions to latex.

- Action

- When latex allergies occur, Employee Health will be notified.

- Evaluation of current prevention and control strategies whenever an employee is diagnosed with latex allergy will be reported in the annual safety report and the quarterly SAFETY 1 QI monitor.

Ergonomics

(see Attachment C for Ergonomics VAMC Memo 138-01)

Ergonomics is the science of designing tools, furniture, machines, or work processes to make them safer and more efficient for human use.  Ergonomic design attempts to prevent musculoskeletal disorders (MSD), such as repetitive motion injuries and back strain.  For more information:  https://www.osha.gov/Publications/laboratory/OSHAfactsheet-laboratory-safety-ergonomics.pdf

- MSD signs and symptoms

It is important to report signs and symptoms as early as possible to prevent serious injury or permanent damage.  Users at risk for MSDs associated with computer use may experience some of the following signs or symptoms:

- Numbness or a burning sensation in the hand

- Reduced grip strength in the hand

- Swelling or stiffness in the joints

- Pain in wrists, forearms, elbows, neck, or back

- Reduced range of motion in the shoulder, neck, or back

- Dry, itchy, or sore eyes

- Blurred or double vision

- Aching or tingling

- Cramping

- Weakness

- Good Working Practice

- Computer Workstation

- Hands, wrists, and forearms are straight, in-line and roughly parallel to the floor.

- Head is level or bent slightly forward, forward facing, and balanced.  Generally, it is in-line with the torso.

- Shoulders are relaxed and upper arms hang normally at the side of the body.

- Elbows stay in close to the body and are bent between 90 and 120 degrees.

- Feet are fully supported by the floor or a footrest may be used if the desk height is not adjustable.

- Back is fully supported with appropriate lumbar support when sitting vertical or leaning back slightly.

- Thighs and hips are supported by a well-padded seat and generally parallel to the floor.

- Knees are about the same height as the hips with the feet slightly forward.

- Use adjustable keyboard platforms under lab benches that accommodate use of the mouse

beside the keyboard.

- Where possible, position computer workstationsin corners or other areas away from doors,entrances and passageways.

- Place monitor so that viewing distance is between 18 and 30 inches.

- Place monitor so the top of the screen is approximately at eye level. This allows the eyes to gravitate naturally toward the center of the screen.

- Use a document holder placed adjacent to and in the same plane as the computer screen.

- Use footrests, where possible, in order to allow them to change leg positions throughout the day.

- Use an appropriate keyboard, mouse or other input devices if they have existing musculoskeletal problems.

- Take mini-breaks of 3 to 5 minutes for every 20-30 minutes of keyboarding or mouse work.These breaks can be spent doing mild hand exercises or stretches.

- Repetitive Pipetting

- Use pipettes with newer trigger mechanisms requiring less force to activate, and use the pointer finger to aspirate, and the thumb to dispense

- Use pipettes that fit comfortably in the user’s hand.

- For tasks such as mixing or aliquotting, use an electronic pipettor with mixing functions.

- Use a multichannel pipettor for large aliquotting tasks.

- Use shorter pipettes.  This decreases hand elevation and consequent awkward postures.

- Elevate chair rather than reaching up to pipette.

- Do not twist or rotate their wrist while pipetting.

- Use low profile waste receptacles for used tips.

- Take micro-breaks of 1-3 minutes for every 20-30 minutes of pipetting.  Mild hand exercises and stretches are beneficial.

- Clean pipettors on a regularly scheduled basis - this reduces "sticking" and improves quality of work.

- Adjust the workstation so the individual doesn’t have to work with their arms in an elevated position.  Work with arms close to the body.

- Rotate pipetting activities between laboratory tasks, hands, and people.

- Use thin-wall pipette tips that fit correctly and are easy to eject.

- Use minimal force when applying pipette tips.

- Keep samples and instruments within easy reach.

- Use an adjustable stool or chair when sitting at a lab bench.

- If is necessary to stand for long periods of time during pipetting, provide anti-fatigue matting.

- Microscopy

  - Use a cut-out work table.  This puts you close to the scope and gives an area for supporting forearms.

  - Sit close to the work surface.

  - Avoid leaning on hard edges.

- Keep elbows close to their sides. Maintain neutral spine.

- Adjust chair, workbench, or microscope asneeded to maintain an upright head position.

- Elevate, tilt or move the microscope close to the edge of the counter to avoid bending their neck.

- Use adjustable eyepieces or mount your microscope on a 30° angle stand for easier viewing.

- Keep scopes repaired and clean.

- Take short breaks. Every 15 minutes, close the eyes or focus on something in the distance.

- Every 30-60 minutes, get up to stretch and move.

- Don’t use a microscope for more than 5 hours per day.  Spread the use out over the entire work day, avoiding long uninterrupted periods of microscope work.

- Provide arm rests to support the operator’s forearms while using adjustment knobs.

- Use an ergonomically designed chair that provides adequate back support, adjustable height, and adjustable seat angle.

- Make sure there is adequate room under the work surface so the operator can pull the chair up to the ocular(s).

- Provide footrests and discourage the use of foot rings on stools.

- Provide sit-stand seats for areas where there is restricted leg room.

- Biosafety Cabinets and Laboratory Workbenches

- Remove unnecessary supplies from the work area.

- Perform all work 6 inches inside the hood.

- Position work supplies in their order of use, with those most frequently used near the front of the hood, but not closer than 6 inches from the face of the hood.

- Place equipment on approved elevated turntables for easy retrieval.  This reduces excessive reaching and twisting, which places an increased load on the low back.

- Use diffused lighting to limit glare.

- Take short breaks to stretch muscles and relieve forearm and wrist pressure.

- Adjust chair/stool to a height that allows the shoulders to relax.Use an ergonomically designed chair that provides adequate back support, adjustable seat angle, and height adjustability between 28 inches to 33 inches.

- Use footrests for individuals whose feet do not rest comfortably on the floor.

- Remove drawers, supplies, refrigerators, and etc. from under the workbenches and cabinet doors from under biosafety cabinets in order to provide leg room.

- Use anti-fatigue matting for laboratory personnel who must stand for extended periods of time.

- Micro-Manipulation & Fine Motor Skills

- If feasible for your work, use plastic vials with fewer threads.  This will reduce twisting motions during capping and uncapping lids.

- Practice using the forceps between the 1st and 2nd digits instead of using the thumb and 1st digit.  Then try alternating between the two positions to reduce the use of the thumb.  The thumb is used repetitively with almost every job task performed in the laboratory.

- Tilt storage bins toward the worker to reduce wrist flexion while reaching for supplies.

- Encourage mini-breaks and hand exercises.

- Microtome & Cryostat Work

- Lower the workstation to keep arms closer to body.

- Apply padding to the front edge of work surface to eliminate sharp edges and increase the amount of blood flow to the hands.

- Retrofit the existing handle with an adapter that will allow the operator to use the hand wheel in a pistol grip position.  This will alleviate repetitive wrist flexion and extension.

- Consider the use of an automatic foot operated cryostat when frequent cryosectioning is performed.

- Avoid placing utensils such as forceps inside the cryostat.

- Use an ergonomically designed chair.

- Take frequent mini-breaks.  These breaks should be used to perform stretching exercises, especially the hands

- Overhead Lifting

- Store heavy objects on shelves below shoulder height whenever possible.

- Use a stable footstool or stepladder to reach objects that are stored on shelves.

- Avoid asymmetric lifting (twisting while lifting).  The object to be lifted should be directly in front of the worker.

- Store materials that are frequently used on shelving units which are located higher than shoulder height.

- Use rotating carrousels to store material close to the worker.  This reduces excessive reaching for objects.

- General Lifting

- Use your thighs rather than the back for lifting.  Keep the head up and trunk tall and lift steadily with the legs.

- Squat down instead of bending at the waist.

- Push loads instead of pulling them.

- Do not twist the body when carrying a heavy load.

- Lift loads close to the body.

- Assessment (see Attachment K for Ergonomic Symptoms Survey Checklist)

- A work-site ergonomics self -assessment will be conducted each year.

- Employees are encouraged to report ergonomic related hazards so that early intervention can be implemented.

- Action

- When ergonomic concerns occur, Employee Health will be notified.

- Employee Health will evaluate employees’ concerns to determine if they are ergonomically related.  If the injury/illness is determined to be ergonomically related, the Employee Health nurse will notify the Safety Office.

- The Safety Office will evaluate medical center work areas for ergonomic hazards based on Employee Health recommendation, upon request from a supervisor, or based on incident reports.

- If indicated, The Safety Office will recommend the engineering controls (equipment, furniture, supplies) or administrative controls (work station setup, use, and work habits) that will assist in mitigating ergonomic problems.

- All ergonomic actions will be included in the annual safety report and the quarterly SAFETY 1 QI monitor.

Noise

(See Attachment F for the Noise Control and Hearing Conservation Program SOP-ENG-138-02)

The Occupational Safety and Health Administration’s (OSHA’s) Noise standard (29 CFR 1910.95) requires employers to have a hearing conservation program in place if workers are exposed to a time-weighted average (TWA) noise level of 85 decibels (dBA) or higher over an 8-hour work shift. OSHA’s permissible exposure limit (PEL) for noise exposure is 90 dBA for an 8-hour TWA and the standard uses a 5 dBA exchange rate. This means that when the noise level is increased by 5 dBA, the amount of time a person can be exposed is cut in half. For example, a person who is exposed to noise levels of 95 dBA (5 dBA above the OSHA PEL of 90 dBA) can be exposed for only 4 hours in order to be within the daily OSHA PEL.

- The laboratory will provide protection against the effects of noise exposure when sound levels equal or exceed an 8-hour time-weighted average sound level of 85 decibels (when people have to shout to be heard)

- Should noise levels become too high, the Lab Safety Officer and/or Hospital Safety Officer should be notified.  An appropriate resolution to reduce decibel levels will be sought or appropriate protection will be provided.

- Noise exposure levels are monitored by the Medical Center Safety Office upon request.  Baseline data is available on the P&LMS SharePoint site at: https://vaww.v09.r03.portal.va.gov/sites/lexvamc/pathology/DocumentReferences/Pathology%20Noise%20Survey%20ver%2012-1-%202010.doc and https://vaww.v09.r03.portal.va.gov/sites/lexvamc/pathology/DocumentReferences/Pathology%20noise%20survey%20ver%2011-7-06%20REF.doc

Housekeeping

- Cleaning/Disinfection

- Environmental Management Service is responsible for routine cleaning of all floors within the Laboratory Service.

- Laboratory personnel are responsible for daily cleaning of all bench tops and other work areas such as fume hoods and laminar flow hoods.

- Laboratory personnel are responsible for daily disinfection of all contaminated bench tops.  A viracidal solution such as Virex or bleach diluted 1:10 should be used according to manufacturer’s instructions to achieve best results.

- Telephones within the laboratory should be disinfected daily.

Maintenance

- All biological safety cabinets will be certified annually.  All fume hoods will be evaluated semi-annually by the Safety Office to ensure proper air flow is being maintained.

- Eye wash fountains

- Pathology & Laboratory Medicine Service will inspect all eyewash fountains on a weekly basis.  Each section of the laboratory is responsible for checking the stations in their section.

- Emergency drench type showers

- Engineering Service will inspect all emergency showers at periodic intervals.

- Fire extinguishers

- Medical Center Safety Office will inspect all fire extinguishers within P&LMS every month.

Inspections

- The Laboratory Safety Officer or designee will perform monthly inspections of P&LMS using the Service Safety Team Inspection Program Checklist (see Attachment I).

- The Medical Center Management Inspection Team will conduct semi-annual inspections (EOC Rounds) to carry out the basic principles of environmental sanitation, safety, security and orderliness for the protection of patients, staff, and visitors.  These inspections are in addition to those safety and sanitation inspections which are part of the normal responsibilities of Engineering and Environmental Management Services.

Passageways

- Stairwells and hallways should not be used as storage areas.

- Access to emergency exits, emergency equipment and utilities control should never be blocked.

- Emergency exit routes are posted in the front of the laboratory in the foyer.

Recordkeeping

- Documentation of all laboratory accidents resulting in property damage or involving spillage of hazardous substances

- All laboratory accidents resulting in property damage or involving spillage of hazardous substances must be reported to the employee’s supervisor and the laboratory safety officer immediately.

- Corrective actions shall be initiated quickly to prevent repeating the incident and to remove the hazardous condition.

- The laboratory accidents resulting in property damage or involving spillage of hazardous substances will be reviewed at the P&LMS safety meeting.

- The accidents and corrective actions will be included in the annual safety report and the quarterly SAFETY 1 QI monitor.

- Documentation of all laboratory accidents resulting in work related injury or illness using ASISTS (see Attachment L for complete ASISTS instructions)

- Employees who incur a work related injury or illness must notify their supervisor immediately.

- The supervisor must initiate the ASISTS accident report (VA Form 2162) immediately followed by an accident investigation within 72 hours of the notification to allow the organization to meet the OSHA recordkeeping requirements

- The electronic ASISTS form may be found on the facility’s Application Server

- Select the  in the bottom left of your screen

- Select “All Programs”

- Select “Application Server”

- Select “ASISTS”

- Corrective actions shall be initiated quickly to prevent repeating the incident and to remove the hazardous condition.

- If the employee needs to be seen by a physician send them to the ED or they may opt to see their private physician.

- If the employee is filing for Worker’s Compensation, the CA1 or CA2 on ASISTS will need to be filled out.

- Printed copies of all reports will be submitted to the Laboratory Safety Officer and kept on file.

- The accidents and corrective actions will be included in the annual safety report and the quarterly SAFETY 1 QI monitor.

Medical Recordkeeping

- Medical records will be retained in accordance with the requirements of State and Federal regulations.  Employees have access to their exposure and medical records upon request.  Findings of periodic monitors for chemical vapors and/or radiation are provided to employees.

Signs

- Prominent signs and labels of the following types should be posted or affixed

- Location signs for safety showers, eyewash stations, and other safety and first aid equipment.

- Location of emergency exits.

- Location of spill kits.

Employee information and training

- The Laboratory Safety Manual, The Chemical Hygiene/Hazardous Communication Plan, and The Infection Control Plans are readily available to all employees and students.  A hard copy of the plan is kept on the bookshelf outside of Room B112, outpatient phlebotomy areas and A606.  Electronic copies of each are stored on the P&LMS SharePoint site at:  https://vaww.v09.r03.portal.va.gov/sites/lexvamc/pathology/DocumentReferences/Chemical%20Hygiene%20-%20Hazardous%20Communication%20Plan.doc and https://vaww.v09.r03.portal.va.gov/sites/lexvamc/pathology/DocumentReferences/Lab%20Safety%20Manual.docx  and https://vaww.v09.r03.portal.va.gov/sites/lexvamc/pathology/DocumentReferences/Infection%20Control%20Guidelines.doc

- Each new employee will be oriented at hospital/medical center VA New Employee Orientation with regard to hazard communication standards, fire, and disaster plans.

- Each new employee will be oriented to the P&LMS Safety Program by the Laboratory Safety Officer and/or their immediate supervisor.  This orientation will include initial safety training and walk through of the lab identifying all lab safety equipment.

- Laboratory personnel will receive annual safety training on the following topics:

- Bloodborne Pathogen Exposure Control Plan

- Infection Control Plan

- Laboratory Safety

- Chemical Hygiene/Hazardous Communication  Plan

- Environment of Care Management Plan (which includes the P&LMS Disaster and Contingency Plans)

- Fire Extinguisher Use

- Fire Drill

- Emergency Eyewash Training

- Ergonomics

- Latex allergies

Electrical safety

(see attachment G for Biomedical Engineering Annual checklist and attachment H for Biomedical Engineering initial checklist for all laboratory instruments)

- Grounding

- All instruments must be grounded and checked by a Biomedical Engineer before use, after repair or modification, and when a problem is suspected.

- Shocks

- Never operate electrical equipment with wet hands.

- All shocks of any magnitude must be reported immediately.  Small shocks often precede major shocks.

c.   Corrective action

- Shut off the current and/or unplug the instrument(s).  Do not attempt to use an instrument that is causing shock.  Report the incident to Biomedical Engineering.

- Repairs

- Repairs on the electrical system of the building are prohibited.

- Any work performed on switches, outlets, circuit boxes, or equipment must be referred to Engineering Service.

- Extension cord usage

- All extension cords and adapters are distributed by the Engineering Service.

- Three-to-two wire adapters (cheater plugs, multiple outlet plugs), and unapproved extension cords are prohibited from use on all equipment.  Modifications to electrical distribution systems shall be accomplished in areas where their use seems to be indicated.  Excessively long line cords are to be avoided.

- Adapters are to be avoided whenever possible.  However, if adapters are needed for equipment fitted with distinctive plugs, they shall be inspected by Engineering Service to assure continuity of all conductors.

- Personal Electric Appliances

- Only corded electrical appliances that receive approval for use in the medical center shall be placed in service.  The approval process for registering a piece of personally owned equipment is outlined in the Lexington, KY VA Medical CenterUse of Personally-Owned Electrical Devices Memorandum No. 138-13.

- No personally owned corded electrical appliances shall be utilized until inspected by Engineering Service.

- The only items that may be permitted, after they pass inspection, are radios, coffee makers, desk fans, desk lamps, microwave ovens or staff purchased refrigerators.  All other corded electrical devices are prohibited.

- Crock pots and similar warming appliances may be permitted by exception for one day use during service pot lucks, farewells, and related events.  They must be removed at the end of the shift.  The Service Chief shall provide approval for these appliances in advance of their use.

- Portable Patient Care Electrical Equipment

- Portable patient care electrical equipment will be inspected before initial use, after repair or modification, and when a problem is suspected.

Fire Safety

- Reporting a fire

- Rescue all persons in immediate danger of life

- Activate Fire Alarm and notify all other personnel in the immediate area

- Call 4500 to inform operator

- The operator will make an announcement over the medical center public address system “Code Red at (give location)”.

- Confine the fire by closing the doors and windows.

- Evacuate the fire area according to Pathology and Laboratory Medicine’s Fire Evacuation Flow Chart (posted in each department).

- Prevention

- Be aware of ignition sources i.e. open flames, heating elements and spark gaps.

- Do not use flammable liquids in the presence of ignition sources.

- Flammable liquids give off vapors which may also ignite or explode.  Be sure flammable liquids are properly stored in approved flame cabinets.  Bulk storage of 5 gallons or more should be stored in approved storage cabinets or approved storage rooms.  The bulk storage in a safety cabinet should be less than 60 gallons per 5,000 square feet of laboratory area.  Do not store any flammable liquid in areas exposed to direct sunlight or heat.

- Fire Drills

  - Hospital fire drills are the responsibility of the Medical Center's Fire and Safety Section.  Records of these fire drills are maintained by the Fire and Safety Section.

  - Pathology and Laboratory Medicine Service fire drills are conducted at least annually

    - All employees will participate in at least one annual fire drill which

includes actually exiting the area via appropriate escape routes

    - Records of fire drill participation will be kept:

      - Manually by the Laboratory Safety Officer in the “Signed Certificates of Completion XXXX (year)

      - Electronically in the Pathology and Laboratory Medicine Education Management Program

- Control of fires

  - Evaluation

- Evaluate the type and extent of the fire.  If it appears to be a large fire, Rescue, Alarm, Confine and Evacuate.  Control measures should only be undertaken for small fires.

- Evaluate the type of material burning (wood, flammable liquids, electrical or gases).

- Our service uses ABC dry chemical fire extinguishers suitable for combustibles, flammables and electrical fires

  - Solid combustibles

- Small objects may be handled with heat resistant gloves and extinguished with water.

- Dry chemical extinguishers may be needed for larger fires.

  - Flammable liquids

- Dry chemical extinguishers are usually needed for safe and effective control of burning liquids.  Do not use water; it enhances the fire's spread.

  - Electrical equipment

- Do not use water unless the circuit has been shut down.

- Shut down circuit if possible

- CO2 is most suitable to prevent further damage to computer equipment, but dry chemical extinguishers are also safe and effective.

  - Gas

- Shut off source if possible.

- Extinguish flame after gas has been shut off.

- Keep flames away from gas cylinders.

- Fire Safety Equipment

- Sand or absorbent material is to be used to contain spread of liquids that have not ignited.

- Our fire extinguishers are ABC dry chemical type.

- Heat resistant gloves may be used to move or handle small burning objects, to handle hot vessels, to turn off hot valves or handles, etc.

Ultraviolet Light Safety

- Definition

- UV radiation is used in chemical synthesis and analysis, medical diagnosis and treatment, sterilization, electrostatic processes, as well as in fluorescent lamps, instrument panel lights, tissue culture rooms, and biosafety cabinets.

- Exposure

- UV light may cause corneal or skin burns from direct or deflected light sources.  Wherever UV light sources are used, suitable and adequate personal protective equipment must be provided, and appropriate approved signage displayed. Continued exposure to UV radiation accelerates aging and may cause skin cancer, cataracts, conjunctivitis, and other conditions.

- Protection

- Protective clothing, gloves and face shields or glasses (rated for UV wavelengths) should be worn when there is danger of exposure to UV radiation.

- All UV lights in use in P&LMS are used as a disinfection aid in the PCR area.  In order to minimize potential harm, they should not be on turned on until the end of the day or when personnel are gone from the work area.

- All UV lights used in P&LMS shall be appropriately labeled with warning signage (i.e. Warning:  This device produces potentially harmful ultraviolet (UV) light.  Protect eyes and skin from exposure.)

Guidelines for Handling Patient Emergencies in Outpatient Phlebotomy Areas

*Note: Anytime a patient becomes critical, call Ext. 4500 (Code 500 team will respond).

- Cooper Drive Division  (General Assistance)

- Whenever a patient becomes ill or faint while in the Outpatient Phlebotomy Area, the aid given will depend on the situation.  However, the following general guidelines should be followed:

- Stay with patient and call on other laboratory personnel for help.

- Call Emergency Room, Ext. 4660/4966 for a physician.

- Alert the Chief of P&LMS and/or the Chief Medical Technologist & Laboratory Manager of the emergency event

- Minimum number of people should remain in the area.

- Patient fainting during phlebotomy.

- Hold his/her head between their knees.

- Use ammonia capsule.

- Stay with patient and call on other laboratory personnel for help.

- Alert the Chief of P&LMS and/or the Chief Medical Technologist & Laboratory Manager of the emergency event

- Transfer patient to ER area when appropriate if he/she feels the need to lie down for a period.

- Leestown Division - Same general guidelines as above with these exceptions:

- Anytime a patient becomes critical, call Ext. 4500 (Code 500 team will respond).

- For general assistance, Primary Care may be called at Ext. 3947 between the hours of 8am-3:30pm.

Laboratory Safety Committee

- The laboratory safety committee will meet semi-annually or as needed.

- The laboratory safety committee will consist of the laboratory safety officer and a representative from:

  - Phlebotomy

  - Clinical laboratory

  - Anatomic laboratory

  - Community Based Outreach Clinic (CBOC’s)

- The laboratory safety committee will:

  - periodically inspect the workplace and prompt elimination of unsafe conditions

  - review accident reports

  - ensure compliance with the safety program

  - report activities to the facility safety committee

# References

| Bloodborne Pathogen Exposure Management Plan | VAMC Memo 111-02 |
| --- | --- |
| Bloodborne Pathogen Exposure Prevention Plan | VAMC Memo 111-01 |
| Creutzfeldt-Jakob and Other Prion Induced Disease | VAMC Memo 111-10 |
| Emergency Eyewash/Shower Guide | VAMC 138-10 |
| Environment of Care Management Program | VAMC Memo 001-07 |
| Ergonomics Plan | http://www.cdc.gov/niosh/topics/ergonomics/#guide |
| Ergonomics Program | VAMC Memo 138-01 |
| Formaldehyde | SOP-ENG-138-12 |
| Hand Hygiene | VAMC Memo 111-09 |
| Hazard Communication Program | VAMC 138-07 |
| Hazardous Product Removal | VAMC Memo 90-07 |
| Hazardous Waste Management | VAMC Memo 138-16 |
| Infection Control Guidelines for Staff | VAMC Memo 111-06 |
| Laboratory Notification Flowcharts for Hazardous Spills/Fire Evacuation |  |
| #### Maintaining A Safe Electrical Environment | SOP-ENG-138-21 |
| Management of Hazardous Chemicals | VHA Directive 2003-030 |
| Management of Health Care Workers Infected with HIV, HepB, or Hep C | VAMC 111-07 |
| Management of Mercury in VHA Facilities | VHA Directive 2002-018 |
| #### NIOSH Publication No. 97-135: Preventing Allergic Reactions to Natural Rubber Latex in the Workplace | http://www.cdc.gov/niosh/docs/97-135/ |
| Noise Control And Hearing Conservation Program | SOP-ENG-138-02 |
| #### OSHA Fact Sheet, Laboratory Safety Ergonomics for thePrevention of Musculoskeletal Disorders | https://www.osha.gov/Publications/laboratory/OSHAfactsheet-laboratory-safety-ergonomics.pdf |
| #### OSHA Fact Sheet, Laboratory Safety Noise | https://www.osha.gov/Publications/laboratory/OSHAfactsheet-laboratory-safety-noise.pdf |
| PLMS Medical Device Related Adverse Patient Event |  |
| Respiratory Protection Program | VAMC Memo 138-14 |
| Service Memo on Designated Clean Areas |  |
| Standard and Transmission Precautions | VAMC Memo 111-11 |
| TB Exposure and Prevention Control Plan | VAMC 111-03 |

### Attachment A:   BLOODBORNE PATHOGEN EXPOSURE PREVENTION PLAN   Memorandum No 111-01

VA MEDICAL CENTER						MEMORANDUM NO. 111-01

JULY 5, 2013								LEXINGTON, KENTUCKY

BLOODBORNE PATHOGEN EXPOSURE PREVENTION PLAN

- PURPOSE:  To eliminate or minimize occupational exposure to bloodborne pathogens, or reduce it to the lowest feasible extent among medical center staff and for complying with Occupational Safety  and Health Administration’s (OSHA) Bloodborne Pathogen Standard, 29 CFR 1910.1030 “Occupational Exposure to Bloodborne Pathogens.” (Also see VA memorandum; Bloodborne Pathogen Exposure Management Plan)

- POLICY:  To provide a safe and healthful work environment for our staff. The Bloodborne Pathogen Exposure Prevention Plan (BBP EPP) is a key document to assist us in implementing and ensuring compliance with the OSHA standard, thereby protecting our employees. All employees have an opportunity to review this BBP EPP at any time. A copy of the prevention plan and the OSHA standard is accessible to all employees on the Lexington Intranet and on the Public Drive/+Infection Control folder.

- RESPONSIBILITIES

- Infection Control Committee.  Review and or update this BBP EPP at least annually or whenever necessary to include new or modified tasks and procedures that affect occupational exposure, and to reflect new or revised employee positions with occupational exposure.

- Employee health

- Maintain employee program and ensure that all medical actions required are performed with the employee’s cooperation.

- Implementation of the BBP EPP and making the written BBP EPP available to OSHA and National Institute for Occupational Safety and Health (NIOSH) representatives.

- Educate employees about Hepatitis B and other bloodborne pathogens, offering and giving the Hepatitis B vaccine. Records of employee participation in the Hepatitis B vaccine program will be maintained in employee health.

- Manage blood/body fluid exposures and maintaining appropriate employee health records.

- Service chiefs/designees

- Ensure that their employees attend mandatory bloodborne pathogen education during orientation and annually and that all training for bloodborne pathogens is provided during working hours (see ATTACHMENT A) and ensuring that Category I and II (see ATTACHMENT B) employees receive annual bloodborne pathogen training.  Learning Resources Center (LRC) is to be notified of the employee names and T&L number so they may be added to the Talent Management System (TMS) for initial and annual bloodborne pathogen training.

- Ensure that education is provided whenever there is a change in an employee's responsibilities, tasks, procedures, or work situation where an employee's occupational exposure is affected.

- Maintain and provide all necessary personal protective equipment (PPE), engineering controls (e.g., sharps containers), labels, and red bags as required and ensure that procedures are followed.  They are also responsible to ensure that adequate supplies of equipment are available in the appropriate sizes; ensuring that employees have and use the appropriate PPE and wear appropriate clothing/shoes to protect from exposures and injuries.

- Ensure use of safe sharp devices and working to identify additional needs to protect their employees. This includes performing an exposure determination of employees/employee tasks where exposure may occur.

- Learning Resources Center.  Coordinates with services to provide initial and annual education on bloodborne pathogens via TMS (see ATTACHMENT A).

- Safety Office. Maintains the OSHA 300 log for the recording of percutaneous injuries from contaminated sharps.  Sharps injuries, involving contaminated objects, are to be recorded on the OSHA 300 Log of Work-Related Injuries and Illnesses and the OSHA 301 Injury and Illness Incident Report. Percutaneous exposure incidents that occur must be recorded on a sharps injury log. The OSHA 300 and 301 forms (initiated through ASISTS) may be used to meet the sharps injury log requirements, provided two conditions are met.  First, the supervisor must enter the type and brand of the device on either form 300 or 301.  Second, the supervisor must maintain the records in a way that segregates sharps injuries from other types of work-related injuries and illnesses, or allows sharps injuries to be easily separated.  The ASISTS reporting system meets both requirements. The sharps injury log will serve as a tool for identifying high risk areas and evaluating devices.  The information in the sharps injury log shall be recorded and maintained in such manner as to protect the confidentiality of the injured employee. The sharps injury log shall contain, at a minimum:

- The type and brand of device involved in the incident. The level of detail presented should be sufficient to allow ready identification of the device, location, and circumstances surrounding an exposure incident, (e.g., the procedure being performed, the body part affected, objects or substances involved, and how they were involved) so that the intended evaluation of risk and device effectiveness can be accomplished.

- The service or work area where the exposure incident occurred, an explanation of how the incident occurred and corrective actions implemented to prevent a repeat of the same or similar incidents.

- Infection Control staff Assists the Infection Control Committee in updating this BBP EPP, provides bloodborne pathogen educational materials and assists with initial or annual education of employees as requested.

- Engineering service.  Ensures that sharps disposal unit holders are installed where needed.

- Environmental Management Service (EMS).  Ensures that there is proper disposal of sharps containers and other infectious waste.

- Employees and volunteers

- Comply with the education requirements upon hire and annually, and procedures and work practices outlined in this BBP EPP (for those who are determined to have occupational exposure to blood or other potentially infectious materials (OPIM) see ATTACHMENT B).

- Inform employee health if he/she is immunocompromised and to follow this memorandum.

- Perform hand hygiene per policy

- Know where the BBP EPP is located and the contents

- Utilize appropriate PPE.

- Report needle sticks and body fluid exposures and following up with employee health. (7). Comply with this policy when staff administers medication with sharps.

- PROCEDURES

- The BBP EPP shall be reviewed and updated at least annually, and whenever necessary, to reflect new or modified tasks and procedures which affect occupational exposure, and to reflect new or revised employee positions with occupational exposure.  Employees that are covered by this BBP EPP receive an explanation of it during orientation.  It will also be reviewed in their annual refresher training (see ATTACHMENT A). The review and update of the BBP EPP shall:

- Reflect changes in procedures and technology that eliminate or reduce exposure to bloodborne pathogens.  This would include, but would not be limited to, newly available medical devices designed to reduce the risk of percutaneous exposure to bloodborne pathogens. Safer medical devices that are appropriate, commercially available, and effective will be used (see ATTACHMENT C).  If a safer device is not available in the marketplace, it is not required. Effective safer medical device is a device that, based on reasonable judgment, will make an exposure incident involving a contaminated sharp less likely to occur in the application in which it is used.

- Document annual consideration and implementation of appropriate commercially

available and effective safer medical devices designed to eliminate or minimize occupational exposure (see ATTACHMENT C).

- Document input solicited from non-managerial employees responsible for direct patient care who are potentially exposed to injuries from contaminated sharps in the identification, evaluation, and selection of effective engineering and work practice controls. Employee input will be solicited on the identification, evaluation, and selection of controls.

- Evaluate for resistance when new devices require staff members to adopt new

techniques, or when staff members perceive that use of the device might have an adverse effect on the patient.  Participation of front-line workers can help to overcome the following barriers:

  - Safer medical devices often require adjustments in technique and staff members are often reluctant to revise practices to which they have become accustomed. With the broad array of devices being used in healthcare settings, it is critical to ensure that devices will work together when necessary so that there are not equipment compatibility problems.

  - The need for continued evaluation of devices and the allotment of sufficient time for adequate device evaluation and after initial use by employees, it may be necessary to replace the device originally selected with a more suitable device.

- Engineering controls and proper work practices

- Engineering controls will be used. These are controls (e.g., needleless systems, blunt suture needles, plastic or mylar-wrapped capillary tubes, biosafety cabinets, and sharps disposal containers) that isolate or remove the bloodborne pathogens hazard from the workplace encompassing not only sharps with engineered sharps injury protections and needleless systems, but also other medical devices designed to reduce the risk of percutaneous exposure to bloodborne pathogens.  Where engineering controls will reduce employee exposure either by removing, eliminating or isolating the hazard, they must be used.

- Work practice controls will be used. These reduce the likelihood of exposure by the manner in which a task is performed.  Work practices include, for example, no-hands procedures in handling contaminated sharps, eliminating hand-to-hand instrument passing in the operating room, employees washing their hands immediately or as soon as possible after removal of gloves or other PPE, contaminated needles and other contaminated sharps are not bent, recapped, removed or broken. Contaminated reusable sharps are placed in puncture-resistant, leak-proof, color-coded containers until properly processed.  If engineering and work practice controls do not eliminate exposure, the use of PPE cabinets is required.

- Identification of the need for changes in engineering control and work practices

- VAMC identifies the need for changes in engineering control and work practices

through for example, evaluation of OSHA records, employee interviews or committee activities. The evaluation of products or procedures is accomplished by the following process: A suggestion about a product or procedure is made for example, by an employee, director, suggestion form, committee, or review of new law or regulation. Committees that can review or make suggestions include but are not limited to the following committees; Infection Control, Commodity Standardization, Environment of Care, Executive Leadership Board, and Accident Review Board.

- Evaluation of the need for new procedures or products to protect against bloodborne pathogens is accomplished by taking a suggestion from the source and referring it to the Logistics Chief (chair of the

commodity standardization committee) per their policy. Part of this evaluation will include whether other devices (safety scalpel handles, blunt suture needles, safer syringes, and/or safer phlebotomy devices) that are commercially available would reduce exposures. Such evidence might include Centers for Disease Control (CDC) studies of efficacy, pilot tests, or data available in published studies.  It should be taken into consideration that the availability or use of an engineering control is not enough to guarantee that an employee cannot be injured. Employee acceptance and employee training are required for the engineering control to be effective. Engineering controls are to be examined and maintained or replaced on a regular schedule to ensure their effectiveness.

- Implementation and Method

- Standard precautions will be practiced at this facility. All blood/body fluids/non-intact skin/mucus membranes will be considered infectious, regardless of the perceived status of the source individual (see Waste Procedures in EMS policies).

- Engineering and work practice controls will be utilized to eliminate or minimize exposure to employees at this facility.  Where occupational exposure remains after institution of these controls, PPE shall be utilized.  The following engineering controls will be used in addition to sharps containers when using the following:

  - Surgical power tools, lasers, and electro-cautery devices

    - May generate aerosols and be a source for splashing and spattering.  Some of these devices include labeling recommendations such as local exhaust ventilation which must be followed by the employee. Anticipated spattering or generation of droplets would necessitate use of eye protection and mask or a face shield, to prevent contamination of the mucous membranes of the eyes, nose, and mouth.

    - Cleaning procedures should not cause unnecessary splashing, spraying, spattering, or generation of droplets of blood or OPIM.

  - Special attention should be given when disposing of needles; drawing blood; handling trash and dirty linens, starting IV's or administering injections to all patients. Gloves will be worn for these procedures and the hands washed when the gloves are removed. The safety device is to be activated immediately after use and then immediately disposed of in a sharp’s container.

  - Needles will not be sheared, recapped, bent, broken, or removed from the syringe. Certain circumstances may exist when no alternative is feasible, and the action is required by the medical procedure, in which recapping, bending, or removing needles is necessary (e.g., administering incremental doses of a medication, such as an anesthetic to the same patient).  In these procedures recapping must be performed by means of forceps or a one-handed technique. The use of the properly performed one-hand scoop method (in which the hand holding the sharp is used to scoop up the cap from a flat surface) for recapping is a recognized and acceptable method; however, the scoop method must be performed in a safe manner, and must also be limited to situations in which recapping is necessary (i.e., with the exception of dental service and the endocrinology clinic, where special devices are available for this purpose).

  - Performing procedures: All procedures involving blood or OPIM shall be performed in such a manner as to minimize splashing, spraying, spattering, and generation of droplets of these substances.

  - Specimens of blood or OPIM

    - Will be placed in a closed and labeled container, which prevents leakage during collection, handling, processing, storage, transport, or shipping, and the specimen container shall be placed within a second container, which prevents leakage.

    - Staff at this facility utilize standard precautions in the handling of all specimens (all specimens are considered potentially infectious material).

    - If the specimens leave the facility, (e.g., during transport to outside lab, shipment, or disposal) a label or red color-coding is required.

    - Extracted teeth, that are being discarded or used as specimens, are subject to the containerization and labeling provisions.  Extracted teeth, gallstones, and kidney stones may be given to the patients.  In these situations, the teeth and stones are not subject to the containerization and labeling provisions of the standard.

  - Bandages that are dripping with blood, or body fluids, or bandages that the blood or body fluids could be squeezed out of, will be placed in a red biohazard container or bag. Bandages that contain the drainage, or have dried blood, will be disposed of in the regular trash.

  - In work areas, where there is a reasonable likelihood of exposure to blood or other potentially infectious materials, employees are not to eat, drink, apply cosmetics or lip balm, or insert contact lenses.  Food and beverages are not to be kept in refrigerators, freezers, cabinets, or on shelves, counters, or bench tops where blood or OPIM are routinely present. Mouth pipetting of blood, or other potentially infectious materials, is not done under any circumstances regardless of the danger or potential infectious nature of the substance.  All procedures will be conducted in a manner which will minimize splashing, spraying, splattering, and generation of droplets of blood, or other potentially infectious materials.

  - Hand washing facilities are available in all areas to employees who have a possibility to incur exposure to blood and/or body fluids and are close to points of possible exposure.  Hand washing for at least 15 seconds, with soap and water, must be performed as soon as feasible, in cases of gross contamination, to adequately flush contaminated material from the skin.  An alternative hand washing method as an interim measure when soap and water are not a feasible means of washing the hands, or other parts of the body, is the waterless hand cleanser. When this type of alternative is used, employees must wash their hands (or other affected area) with soap and running water as soon as feasible thereafter. Employees shall wash hands and any other skin with soap and water, or flush mucous membranes with water immediately, or as soon as feasible following contact of such body areas with blood or other potentially infectious materials.

- Safer medical devices will be used (See ATTACHMENT C).

- Labels and signs (See ATTACHMENT D).

- Equipment which has become contaminated with blood, or other potentially infectious materials, shall be decontaminated per the manufacturer’s instructions before any maintenance (in house or contract) is performed on it. The service using the equipment is responsible for contacting Supply, Processing and Distribution (SPD) for making arrangements to have equipment cleaned prior to any service work being performed. Decontamination will be by the use of the hospital approved disinfectant.

- PPE (see ATTACHMENT E)

- Blood or Body Fluid Spills

  - Personnel must wear gloves when cleaning blood or body fluid spills. Other PPE/barriers (gowns or apron) will be used, if indicated. EMS may be notified and assistance requested for large spills (greater than 50 cc’s).

  - An EPA-approved germicide/disinfectant or a 1:10 dilution of sodium hydrochloride (bleach) must be used to disinfect the area after the spill is wiped. Recommended contact time for the EPA- approved germicide/disinfectant is 10 minutes, or until the solution is dry.  For carpeted areas, remove gross material, and call EMS immediately for cleaning and disinfecting.

  - Disposable cloths used to wipe the spills are discarded in regulated infectious wastes containers if they contain 50cc or more of blood or a body substance.

  - Contaminated clothing of any kind is to be removed as soon as possible, and laundered according to healthcare requirements.

- Housekeeping. VAMC is maintained in a clean and sanitary condition by units and EMS, (See EMS Policies) by following a written schedule for cleaning and method of decontamination based upon the location within the facility, type of surface to be cleaned, type of soil present, and tasks or procedures being performed in the area. All work surfaces, patient rooms, floors, protective coverings, etc. are cleaned with a broad-spectrum disinfectant.  If an area becomes overtly contaminated, it is decontaminated immediately.  While extraordinary attempts to disinfect or sterilize environmental surfaces, such as walls or floors, are rarely indicated, routine cleaning and removal of soil are required. The particular disinfectant used, as well as the frequency with which it is used, will depend upon the circumstances in which the housekeeping task occurs. Since environmental contamination is an effective method of disease transmission for HBV (the CDC states that HBV can survive for at least one week in dried blood on environmental surfaces or contaminated needles and instruments), there are minimum requirements for the cleaning and decontamination of equipment and environmental and working surfaces that come into contact with blood or OPIM.

  - The work surface decontamination is to be performed at the end of the work shift, if the work surface may have become contaminated since the last cleaning by, for example, setting down contaminated instruments or specimens on the work surface.

  - The use of protective coverings (i.e. plastic bags) is an acceptable alternative for protecting clean items and surfaces against contamination.  If this option is chosen, the covering must be removed and replaced as soon as feasible following overt contamination or at the end of a work shift, if it may have become contaminated during the shift.

  - Inspection and decontamination is required, on a regularly scheduled basis, of cans, bins, pails, wash or emesis basins and so forth, which are intended for reuse.

  - Since contaminated broken glass (e.g., phlebotomy tubes) is capable of inflicting injury, broken glassware which may be contaminated must not be picked up directly with the hands.  It shall be cleaned up using mechanical means, such as a brush and dust pan, tongs, or forceps and placed in the red sharps containers or the blue trash cans (on patient care units) for glass. The tools used in cleanup (e.g., brush and dustpan or forceps) must be properly decontaminated or discarded after use. Vacuum cleaners are not appropriate for cleanup of contaminated broken glass.

  - Employees are not to place their hands into containers whose contents include reusable sharps contaminated with blood or OPIM because the contents cannot be seen and safely handled.  (i.e., employees must not reach into sinks filled with soapy water into which sharp instruments have been placed, instead use strainer type baskets to hold the instruments and forceps to remove the items).

  - Proper decontamination of reusable medical equipment (RME), such as glassware or hand instruments cannot be achieved in the presence of organic debris (e.g., blood) because it interferes with the disinfecting/sterilizing process.  In other words, the dried blood or secretions must be removed before the item can be disinfected or sterilized. RME is to be reprocessed per the applicable standard operating procedure or cleaning schedule.

- Regulated waste (see ATTACHMENT F)

- Laundry procedures (see ATTACHMENT G)

- Hepatitis B Vaccine (see ATTACHMENT H)

- Record keeping (see ATTACHMENT I)

- REFERENCES: Under Secretary for Health’s Information Letter Blood-Borne Pathogens (BBP) Hazards and Management IL 10-2012-010, OSHA Bloodborne Pathogen Standard; 29 CFR 1910.1030. CDC Updated U.S. Public Health Service Guidelines for the Management of Occupational Exposures to HBV, HCV, and HIV and Recommendations for Post-exposure Prophylaxis. MMWR June 29, 2001; 50 (RR-11), pp 1-42.

http://www.osha.gov/pls/oshaweb/owadisp.show_document?p_table=STANDARDS&p_id=10051

- FOLLOW-UP RESPONSIBILITY: Infection Control (111), Occupational Health, and Safety Office.

- RECERTIFICATION: On or before July 5, 2014.

Pamala Thompson, RN, MSA, MSN, NEA-BC, FACHE

Interim Director Attachments (9)

ATTACHMENT A

ANNUAL BLOODBORNE PATHOGEN TRAINING

- PURPOSE: Bloodborne pathogen training upon hire or during orientation and follow-up annual re- training is given to promote patient and employee safety and knowledge.  This memorandum defines what is to be included in annual bloodborne pathogen training and assists in complying with Occupational Safety and Health Administration (OSHA) regulations.

- POLICY: All employees will receive initial bloodborne pathogen training upon hire or beginning work during the orientation process.  Employees are to receive additional training when changes such as modification of tasks or procedures or institution of new tasks or procedures affect the employee's occupational exposure.  The additional training may be limited to addressing the new exposures created. Annual re-training of employees with the potential of exposure to blood or body fluids is required.

- RESPONSIBILITIES

  - Medical center director is responsible for ensuring that necessary resources are provided to educate employees and for requiring full compliance with all aspects of this policy.

  - Section/service chiefs or designee are responsible to provide or assure staff have completed;

- Bloodborne pathogen training upon hire or during orientation. Employees will be instructed in the specifics of the current guidelines and in any guideline, memorandum or procedure related to their specific duties.

- Annual Bloodborne Pathogen Training as outlined in this policy.

- Documentation of education will be the responsibility of the Section/Service Chiefs and is to include; the dates of the training session, names and qualifications of persons conducting the training, the names and titles of all persons attending the training sessions and records are to be maintained for 3 years from the date on which training occurred.  Training or certification of training may be done by the employee on TMS.

  - The Infection Control Practitioner and/or the Hospital Epidemiologist will provide ongoing inservice educational activities, consultation, and or assistance on an as needed basis.

  - All employees with the potential of exposure to blood or body fluids as listed in Attachment B (Exposure Determination as required by Federal Bloodborne Pathogen Law 1910.1030) are required to complete annual bloodborne pathogen training. Employees are responsible for knowing infection control policies, memorandum and guidelines.

- PROCEDURES

- Training may be done as follows.

- Use this policy (4.b and 4.c) as a guideline for annual training by the service/section chief or designee.  Following education have staff self certify education in TMS.

- Having the employee complete a Bloodborne Pathogen Training module on the VA Talent Management System (TMS).

- Managers, supervisors, service chiefs, infection control staff or the ID on call may be contacted for questions.

- Annual blood borne pathogen training must cover the following.

- A copy of the regulatory text of the OSHA bloodborne pathogen standard is available to each employee on the Public Drive in the +Infection Control folder or at www.OSHA.gov or www.osha.gov/pls/oshaweb/owadisp.show_document?p_table=STANDARDS&p_id=10051).  This standard outlines federal requirements to protect employees.

- Review the definition of standard precautions and that all body fluids are considered infectious.

- Give a general explanation of the epidemiology and symptoms of bloodborne diseases including discussion of HIV/AIDS, Hepatitis B, Hepatitis C and others depending on group-possible symptoms, and meaning of tests done on source patient and employee.

- Review the modes of transmission of bloodborne pathogens and ways most likely to be transmitted in the community and in the health care setting.

- All employees are responsible for memorandum 111-02 Bloodborne Pathogen Exposure Management Plan and memorandum 111-01 Bloodborne Pathogen Exposure Prevention Plan. Explain what is in and where the bloodborne pathogen exposure control plans can be found.

- Give an explanation of the appropriate methods for recognizing tasks and other activities that may involve exposure to blood and other potentially infectious materials.

- Explain the use and limitations of methods that will prevent or reduce exposure including appropriate engineering (sharps containers and needles and scalpels with safety devices) and work practices (hand hygiene and use of PPE) controls. Train on specific devices used at this facility, including newly purchased devices.

- Give an explanation of biohazard signs/labels and/or color coding (red bags and red sharps containers).

- Review information on the types, proper use, location, removal, handling, decontamination and disposal of PPE.  Identify where to get PPE and who to tell if it doesn't work/fit, etc.

- Give an explanation of the basis for selection of PPE.  If splash or splatter is expected or anticipated, wear PPE.  Discuss use of gloves, gowns, masks, protective eyewear, ambu-bags, sharps containers, not using one-handed recapping of needles and how to transport specimens.

- Give the following information on the hepatitis B vaccine; Three shots in the arm (deltoid) over 6 months usually provides immunity. The shot is safe, can provide immunity to Hepatitis B and the vaccination is offered free of charge through employee health.

- Review appropriate actions to take and persons to contact in an emergency involving blood or other potentially infectious materials; 1) wash the area with soap and water or flush mucous membranes with water, 2) notify supervisor and 3) report to Employee Health during administrative hours or to the Emergency Room after hours.

- During follow-up immediately after an exposure, baseline tests will be ordered for the employee and source patient.  Depending on risk, medicines are available for employees to prevent specific diseases.  The employee is to get a copy of the assessment done by employee health on the post- exposure evaluation and follow-up.

- Employees receiving training shall have an opportunity for questions to and answers from the person doing the education.  Managers, supervisors, service chiefs, infection control staff or the ID on call may also be contacted for questions.

ATTACHMENT B

EXPOSURE DETERMINATION

OSHA requires employers to perform an exposure determination concerning which employees may incur exposure percutaneously, to mucus membranes, eyes, or to intact or non-intact skin by occupational exposure to blood, (includes plasma, platelets) and OPIM. OPIM includes the following human body fluids: semen, vaginal secretions, cerebrospinal, synovial, pleural, pericardial, peritoneal, or amniotic fluids, saliva in dental procedures, any body fluid that is visibly contaminated with blood, and all body fluids in situations where it is difficult or impossible to differentiate between body fluids; any unfixed tissue or organ (other than intact skin) from a human (living or dead). Also included are medications derived from blood.

    - The exposure determination is made without regard to the use of personal	protective equipment (i.e., employees are considered to be exposed even if they wear PPE).

    - This exposure determination is required to list all job classifications in which all employees may be expected to incur such occupational exposure, regardless of frequency.

    - Exposures to intact skin are usually not of concern except in those instances where there was an exposure for a considerable length of time (i.e. covered in blood for hours). Non-intact skin exposures include skin with dermatitis, hangnails, cuts, abrasions, chafing, acne, etc.

    - Pathogens include, but are not limited to: HBV (hepatitis B), Human Immunodeficiency virus (HIV), hepatitis C virus, malaria, syphilis, babesiosis, brucellosis, leptospirosis, arboviral infections, relapsing fever, viral hemorrhagic fever, malaria, syphilis, human T-lymphotrophic virus Type 1, adult T-cell leukemia/lymphoma (caused by HTLV-I), HTLV-I associated myelopathy, and diseases associated with HTLV-II.

JOB CLASSIFICATIONS

- The following job classifications include those classifications where employees are likely to incur occupational exposure to blood or other potentially infectious materials: The following job classifications have been identified in which all (Category I) employees in these job classifications have occupational exposure.  The list below identifies the individuals who will receive the training, protective equipment, vaccination, and other protections of the BBP EPP. These guidelines extend only to employees in the workplace and do not extend to students if they are not considered employees, to state, county, or municipal employees, to health care professionals who are sole practitioners or partners, and to the self-employed.  Independent Contractors are companies that provide supervisory personnel and/or rank-and-file workers, to carry out a service such as Construction Companies and are responsible for complying with all provisions of the standard and in accordance with the guidelines of VAMC. VAMC expects contract workers to follow hospital policies.  PPE and inservices will be provided to VAMC employees.

| Category I Employees Likely To Incur Occupational Exposure |  |  |
| --- | --- | --- |
| Anesthesia Technicians | Health Technicians (Vascular, Ortho, Urology, Endoscopy, Ophthalmology, Cardiac Cath) | Physical Therapist Assistants |
| Certified Registered Nurse Anesthetists | Histo-technicians (Series 0646) | Physician Assistants |
| Chemists (Series 1320) | Infection Control Practitioners | Physicians |
| Chief of Staff | Licensed Practical Nurses | Registered Respiratory Therapists/Anesthesia Technicians |
| Clinical Staff Social Workers | Medical Technologists | Registered Nurses |
| Clinical Clerks (Medical Students) | Medical Instrument Technicians (Endoscopy, Pulmonary, EKG, Endocardiography, EEG, PSG, Echocardiography, Biomedical) | Respiratory Therapists |
| Cytotechnologists (Series 0601) | Medical Supply Technicians | Secretary/Stenographers - Anesthesiology |
| Dental Lab Technicians | Medical Laboratory Aides (Series 0645) | Speech Pathologists |
| Dental Assistants | Microbiologists (Series 0403) | Students/trainees in any of the above |
| Dental Hygienists | Nursing Assistants | Supervisory Staff Social Workers |
| Dentists | Occupational Therapists | Team Administrators |
| Diagnostic Radiologic Technologists | Patient Escort, Specimen Transport, Pathology, Nursing and Rehab Volunteers | Veterans Benefits Counselor |
| EMG Technicians | Physical Therapists | Vocational Rehabilitation Therapists |

- Some of the employees in the following class (Category II) may incur exposure to blood or OPIM during specific tasks or procedures.

The job classifications and associated tasks for these categories are as follows:

| Category II Employees May Incur Occupational Exposure |  |
| --- | --- |
| Job Classification | Tasks/Procedures |
| A/C Mechanics | Working in patient rooms (air handlers, thermostats, etc.) |
| Admission and Clinical Clerks – | Interviewing patients |
| Audiovisual Production Specialists Medical Photographers | Photographing or videotaping patients during invasive procedures (i.e., surgery) |
| Bio Lab Tech - Research | Fixation of human cranial bone samples, Isolation of monocytes from human blood |
| Carpenters | Furniture repair, changing floor tile |
| Chaplains | Bedside patient visits |

| Category II Employees May Incur Occupational Exposure |  |
| --- | --- |
| Chemists | Research Working with human blood and tissue, performing cytokine assays on human serum samples from healthy volunteers, using various clinical samples from human sources containing viral pathogens |
| Clinical Dietitians | At patient bedside |
| Clothing Clerks | Handling patient clothing and personal effects |
| Computer Specialists and Electronic Technicians | Working on computers that have body fluids on them at times |
| Electricians | Working on electrical items/systems in patient rooms |
| Electronics Mechanics | Working on pillow speakers and communication equipment |
| Food Service Workers - Dietetic | Handling food trays from patients' bedside and during dishwashing process, potential for Needle stick from needles left on tray |
| Health Science Specialists - Research | Processing blood samples and transferring plasma by disposable pipettes into vials |
| Health Unit Coordinators | Patients in area |
| Housekeeping Aids/Foreman | Cleaning patient rooms, OR, ER, ICUs, rest rooms, trash cans and handling soiled linen |
| Insurance Clerks | Patient visits |
| Laborers | Handling garbage |
| Laundry Machine Operators | Handling soiled patient clothes |
| Laundry/Linen Section Staff | Handling soiled linen and uniforms |
| Machinists | Working on patient beds |
| Mail Clerks | Mail contaminated with blood/body fluids |
| Medical Technicians - Research | Positioning the human foot for microcirculation studies |
| Medical Record Technicians | Records contaminated with blood/body fluids |
| Office Automation Assistants, Office Automation Clerks, Secretary, Clerks - Pathology | When opening mail containing biological specimens, surveys or other such material |
| Painters | Working in patient rooms |
| Pest Control Staff | Pest treatment throughout |
| Pharmacists | Working at bedside |
| Plasterers | Working in patient rooms |
| Plumbers | Working on plumbing drains and pipes |
| Police Officers | Exposure to body fluids e.g. patients/visitors |
| Prosthetic Clerks | Patient visits |
| Recreation Therapists/Assistants | Working with patients |
| Release of Information Clerks | Patient visits |
| Secretary - EMS, CDD | Handling patient clothing and personal effects |
| SPD Medical Supply Technicians | Decontaminate all the grossly soiled instruments/equipment used throughout the medical canter |

| Category II Employees May Incur Occupational Exposure |  |
| --- | --- |
| Transportation Staff - Engineering | Transporting patients |
| Utility System Operators | Changing filters in air handlers in patient rooms |

- Category III Job Classifications: Any employees and volunteers with job classification not listed in Categories I or II above, have no occupational exposure. These individuals are offered the Hepatitis B vaccine, however, are not required to receive the training, protective equipment, vaccination, and other protections of the BBP EPP.  Any of the employees in Category III, having a change in duties that would cause an actual or potential for exposure, must notify the Employee Health Nurse immediately.

ATTACHMENT C

SAFER MEDICAL DEVICES

- Safer medical devices will be used and are generally of two types

  - Needleless systems (e.g., intravenous medication delivery systems that administer medication or fluids through a catheter port or connector site using a blunt cannula or other non-needle connection) are devices that do not use needles for:

    - The collection of bodily fluids, or withdrawal of body fluids after initial venous, or arterial access is established;

    - The administration of medication or fluids; or

    - Any other procedure involving the potential for occupational exposure to bloodborne pathogens due to percutaneous injuries from contaminated sharps.

  - Sharps with engineered sharps injury protection (e.g., syringes with a sliding

sheath that shields the attached needle after use or needles that retract into a syringe after use self-sheathing needles on syringes) means a non-needle sharp or a needle device used for withdrawing body fluids, accessing a vein or artery, or administering medications or other fluids, with a built-in safety feature or mechanism, that effectively reduces the risk of an exposure incident (includes, but is not limited to, shielded or retracting catheters used to access the bloodstream for intravenous administration of medication or fluids, and intravenous medication delivery systems that administer medication or fluids through a catheter port or connector site using a needle that is housed in a protective covering.) There are specific design features for recessed needle systems that are important in preventing percutaneous injury and which have the following characteristics: A fixed safety feature provides a barrier between the hands and the needle after use.  The safety feature should allow or require the worker's hands to remain behind the needle at all times:

    - The safety feature is an integral part of the device and not an accessory.

    - The safety feature is in effect before disassembly and remains in effect after disposal to protect users and trash handlers, and for environmental safety.

    - The safety feature is as simple as possible, and requiring little or no training to use effectively.

- Evaluation of replacement of new products for safe needle devices will be ongoing. Trending of needle stick and sharp injuries will be accomplished at least quarterly by the Infection Control Committee who will make recommendations for the evaluation of new products if indicated. Other groups which might make recommendations for the evaluation of new products include user groups, Commodity Standard Committee, Environment of Care Committee or Nursing

Clinical Practice Council.  Input from user groups responsible for direct patient care who are potentially exposed to injuries from contaminated sharps will be solicited in the identification, evaluation, and selection of safer medical devices.  This could take place as involvement in informal problem-solving groups, participation in analysis of exposure incident data or participation in the evaluation of devices through pilot testing.

- Education by the service of their staff regarding the appropriate and correct use of safe needle devices will be ongoing.  Educational needs of staff will in part be determined by trending of needle stick and sharps injuries. Training for new equipment will involve company representatives.  Other training, provided by knowledgeable staff, will include interactive classes during orientation, annual reviews, and any other time deemed necessary.

- A sharps injury log will be maintained by Employee Health.  This information will be presented to the Infection Control Committee on a quarterly basis.

- Annual evaluation of safety devices where safer medical devices are available to employees, their use is mandatory and exceptions will be cleared by the Infection Control Committee prior to the excepted use.

- Per memorandum 90-31 Product and Technology Standardization Team/Commodity Standardization Committee, companies are contacted for information or trial products and evaluation forms for products are devised or selected.

- Employees will be involved in the evaluation of new products or procedures. New products will be obtained to be sampled on each applicable unit. During the time, a product is sampled and evaluation forms will be given to employees sampling the product.  The evaluation forms will be analyzed by the department/unit doing the trial, then the forms will be routed to the infection control practitioner for the evaluation in the annual BBP EPP revision. The summary of the evaluations from each department/unit will be taken into consideration along with costs and availability in the selection of the product(s). New procedures will follow the same process.

- Once a new product or procedure is selected and put into place, staff will be encouraged to continue to give feedback by, for example, sending comments using evaluation forms, e-mail, or calling a specified contact person for the new product or procedure.

- When the Product & Technology Standardization Team/Commodity Standardization Program denies a product, they are to document why the item cannot be purchased.  This establishes a procedure for the denial of a product especially when a user requests the device and it meets the safe design needs of sharps.

2013 Review

| TYPE OF DEVICE | BRAND OF SAFETY DEVICE AVAILABLE | EVALUATION OF SAFETY DEVICES |
| --- | --- | --- |
| Blood Collection Devices |  |  |
| TYPE OF DEVICE | BRAND OF SAFETY DEVICE AVAILABLE | EVALUATION OF SAFETY DEVICES |
| Butterfly blood collection device | Safety-Lok blood collection set (Vacutainer) | No change anticipated |
| Blood collection | BD Vacutainer Eclipse (Lab) | No change anticipated |
| Finger stick lancets | 1-step Safety Lancet (Surgilance) | No change anticipated |
| Lab capillary tubes | Plastic (BD) | No change anticipated |
| Lab vacuum tubes | Plastic (BD) | No change anticipated |
| IV delivery system |  |  |
| PIV Catheter Insertion Device | Introcan Safety (B. Braun) | No change anticipated |
| IV delivery system | MicroClave Clear needleless end cap | No change anticipated |
| Port-a-cath needles | B. Braun WHIN safety Huber needle set | No change anticipated |
| Radiopharmaceuticals injections | Saf-T-Intima (BD) access system (Nuc Med) | No change anticipated |
| Syringes & Needles |  |  |
| IM/SQ, PPD and Insulin | Monoject Safety Syringe (Kendal) | PCS reviewing need for insulin syringe change |
| Carpuject | Leur loc | No change anticipated |
| Needle | Mini-spike Dispensing pin (Nuc Med) | No change anticipated |
| Butterfly | Safety-Lok blood collection set (Nuc Med) | No change anticipated |
| Blood Gas Syringe | Arterial Blood Sampling Kit (Sims Portex) | No change anticipated |
| Venipuncture Needle Protection Device | Punctur-Guard ( Bio-plexus) | No change anticipated |
| Biopsy needle | Bard biopsy needle device (Urology) | No change anticipated |
| Spinal needle | Havel’s 22gaugex 8” spinal needle (Urology) | No change anticipated |
| Suture needle | None | Blunt and tapered needles used for muscle and fascia tissue |
| Huber needles | Surecan Safety Huber (B.Braun) | No change anticipated |
| Hemodialysis needles | AV Fistula with Wing-eater (JMS) | No change anticipated |
| PICC introducer needles | None | Continue to search for new technology |
| Syringe and needle for Nuclear Medicine | None | Explore new technology because current technology unacceptable to users |
| Central Venous Catheterization Kit | Tray with safety features; locking disposal cup staple anchoring device, safety scalpel (Arrow) | No change anticipated |
| Needles which have no | None | Trial of Point-Lok Safety Device |

| TYPE OF DEVICE | BRAND OF SAFETY DEVICE AVAILABLE | EVALUATION OF SAFETY DEVICES |
| --- | --- | --- |
| current safety technology |  | (Portex) – Already in use in Research |
| Medication Access |  |  |
| Medication vial access | Needleless system (B. Braun) Mini- Spike dispensing pin with ultrasite valve, monoject smartip needle, monoject blunt needles | No change anticipated. |
| Scalpels |  |  |
| Scalpel, disposable safety | Sandel | No change anticipated. |
| Scalpel handles, reusable safety | DeRoyal's Canica | No change anticipated |
| Other Safety Devices |  |  |
| Urine sampling devices | Foley tray with needleless Safe Sampling Port (Bard) | No change anticipated |
| Dental Devices | Dental Dams (Rubber) Needle sheath prop one handed recapper (ProTector) | Continue to search for new technology |
| Marker | Civco fiducial marker kit (Urology) | No change anticipated |

ATTACHMENT D

Labels and Signs

- Biohazard warning labels (fluorescent orange with lettering and symbols in contrasting color) shall be used to warn employees about the presence of bloodborne pathogens. These labels shall be on the following: containers of regulated waste, refrigerators and freezers containing blood or other potentially infectious materials, containers used to store, transport or ship blood or other infectious materials and on fully or partially contaminated equipment. Red bags and containers shall be used for infectious waste and disposable, used sharps.   Employees are to notify their supervisor if they discover regulated waste containers, refrigerators containing blood or OPIM, contaminated equipment, etc. without proper labels.

  - Warning labels shall be affixed to containers of regulated waste, refrigerators, and freezers containing blood, or other potentially infectious material, and other containers used to store, transport, or ship blood, or other potentially infectious materials.

  - These labels shall be fluorescent orange with lettering and symbols in a contrasting color.  Red bags or red containers may be substituted for labels.

  - Labels shall be affixed as close as feasible to the container by string, wire, adhesive, or other method that prevents their loss or unintentional removal.

  - Containers of blood, blood components, or blood products that are labeled as to their contents and have been released for transfusion, or other clinical use, are exempted from the labeling requirements.

  - Individual containers of blood, or other potentially infectious materials, that are placed in a labeled container during storage, transport, shipment, or disposal are exempted from the labeling requirement.

  - Labels required for contaminated equipment shall be in accordance with this paragraph and shall also state which portions of the equipment remain contaminated.

  - Signs shall be posted at the entrance of work areas specified above.

ATTACHMENT E

Personal Protective Equipment (PPE)

- All PPE used will be provided without cost to employees. PPE are chosen based on the anticipated exposure to blood or other potentially infectious materials.  The PPE will be considered appropriate only if it does not permit blood, or other potentially infectious materials, to pass through or reach the employees' clothing, skin, eyes, mouth, or other mucous membranes under normal conditions of use, and for the duration of the time the protective equipment will be used.

  - Due to safety issues, sandals, CROCs (with holes in the top) or other open toed shoes are NOT

to be worn in direct patient care areas.

  - Gloves are available in a variety of types and sizes.  Gloves are required for any phlebotomy or vascular access procedure. They are not required for an injection as long as hand contact with blood or other potentially infectious materials is not reasonably anticipated

  - Utility gloves may be decontaminated for reuse if their integrity is not compromised. Discard utility gloves if they show signs of cracking, peeling, tearing, puncturing, or deterioration.  Never wash or decontaminate disposable gloves for reuse.

  - Wear appropriate gloves when it can be reasonably anticipated that there may be hand contact with blood or OPIM, and when handling or touching contaminated items, or surfaces. Replace gloves if torn, punctured, contaminated, or their ability to function as a barrier is compromised.

  - Masks in combination with eye protection devices, such as goggles or glasses with solid side/top shields, or chin length face shields, are required to be worn wherever splashes, spray, splatter or droplets of blood/body fluids may be generated and eye, nose or mouth contamination can reasonably be anticipated.

  - Fluid repellant gowns are available to prevent minor amounts of blood/body fluids from getting on clothing.  In areas such as the operating room (OR) and endoscopy suite, the gowns are impervious to blood/body fluids.

  - To minimize migration of contamination beyond the work area, employees must:

    - Remove immediately, or as soon as feasible, any garment contaminated by blood or OPIM in such a way as to avoid contact with the outer surface PPE after it becomes contaminated, and before leaving the work area.

    - Dispose PPE contaminated by blood or OPIM in red biohazard bags.

    - Dispose of PPE that is not contaminated by blood or OPIM in the black leak proof plastic trash bags.

    - Perform hand hygiene immediately or as soon as feasible after removal of gloves or other PPE.

    - PPE is to be worn in the area of contamination and is not to be worn in the halls. This includes gloves, masks and shoe covers.

    - If uniform/clothing becomes contaminated, change into scrubs before leaving a

work area.  The employee may sign out a pair of scrubs from the supervisor. Wash any areas of skin contaminated by blood or body fluids.  Soiled clothing is to be placed in black plastic leak proof bags and taken home and cleaned in the in the home washer and dryer with detergent and bleach according to the labels.

ATTACHMENT F

Regulated Waste and Disposal

The term "regulated waste" refers to categories of waste, which require special handling such as liquid or semi-liquid blood or OPIM and include: contaminated items that would release blood or OPIM in a liquid or semi-liquid state if compressed (i.e. bloody dressings), items that are caked with dried blood or OPIM and are capable of releasing these materials during handling (i.e. specimen containers, test tubes, tissue specimen, liquid blood specimens, used blood tubes), and contaminated sharps; cultures, pathological and microbiological wastes containing blood or OPIM.

- Sharps containers are for the disposal of contaminated needles and other sharp	items.

These will be as close to the point of use as possible. Needle products and contaminated disposable sharps shall be discarded immediately into sharps containers.  It is suggested that guide wires be coiled and enclosed in a glove before disposal into a sharps container.

- Glass that is broken or unbroken is to be placed in a box, or in the container in the soiled utility room that is labeled “glass”.

- All contaminated sharps shall be discarded as soon as feasible in sharps containers, which may be found in the following locations; medication carts, patient rooms (except on 4S), exam/treatment rooms, dietetic dining room, hemodialysis, employee health, intensive care units, emergency room, or suite.

- Regulated waste (blood or body fluids that could spill or leak) other than sharps shall be placed in plastic bags that are marked as infectious waste or biohazard.  The bags are to be placed in red plastic barrels marked as biohazard. These barrels may be found in the following locations: Soiled SPD rooms; Emergency Room; Dental Clinic; Hemodialysis; Endoscopy; Wards; Soiled Utility Rooms.

- The containers are inspected, maintained, and replaced by staff on a schedule assigned by each service chief or nurse manager.  Environmental Management Staff (EMS) may also check the sharps containers.

- Check sharps container and do not allow overfilling.  If it is 2/3 to 3/4 full, carefully

close the lid, remove the box, and replace it with a new one.  Seal the filled container tightly; take to the area designated by the unit, and place in infectious waste container for proper disposal.

- Containers are to be maintained upright throughout use.

- New sharps containers are to have the lids securely fastened when placing them in the holder.

Containers may also be locked onto a mobile cart, if one is used by healthcare workers in these units, or they may be brought to the site and removed by the employee upon leaving.

- When moving containers of contaminated sharps from the area of use, the containers shall be closed immediately prior to removal, or replacement to prevent spillage, or protrusion of contents during handling, storage, transport, or shipping, and placed in a secondary container, if leakage is possible. The second container shall be: closable and constructed to contain all contents, and prevent leakage during handling, storage, transport, or shipping.

- Immediately, or as soon as possible after use, contaminated reusable sharps shall be placed in appropriate containers until properly reprocessed.  Reusable sharps, such as large bore needles, scalpels, and saws, pose the same percutaneous exposure hazard as disposable sharps, and must be contained in a manner that eliminates or minimizes the hazard until they are reprocessed. Therefore, the containers for reusable sharps must meet the same requirements as containers for disposable sharps, with the exception that they are not required to be closable, since containers used for collecting and holding reusable sharps will be reused. Contaminated sharps, that are reusable, are to be placed back on the procedure tray, as soon as possible, after use.  All contaminated reusable sharps are considered as potentially infectious material and will be sent to SPD for sterilization.

- Sharps containers and red bag wastes are picked up by environmental management service from designated areas in each department.

ATTACHMENT G

Laundry Procedures

All soiled linen will be handled and bagged as potentially infectious.  All soiled linen will be placed in self-closing linen bags that are designed to contain moisture and airborne contaminants.  Linen will be double bagged if the outside of the first bag is contaminated with body substances. All employees who handle soiled linen will utilize PPE to prevent contact with blood/body fluids.  Laundry from this facility will be cleaned at the Louisville VA Medical Center whose personnel also handle all laundry as being potentially contaminated.

- Standard Precautions will be used in the handling of all laundry.

- All soiled linen is considered contaminated. Contaminated laundry shall be handled as little as possible with a minimum of agitation and shall be bagged at the location where it was used and shall not be sorted or rinsed in the location of use.

- Whenever contaminated laundry is wet and presents a reasonable likelihood of soak through of or leakage from the bag or container, the laundry shall be placed and transported in bags or containers which prevent soak-through and/or leakage of fluids to the exterior.

- Not all contaminated laundry must be placed in such bags or containers; only laundry that is wet enough to leak or soak through and expose employees handling the bags/containers to blood or OPIM, or contaminate other surfaces.

- Employees who have contact with contaminated laundry wear protective gloves and other appropriate PPE in order to prevent or reduce contact exposure to blood or OPIM.

ATTACHMENT H

HEPATITIS B VACCINE AND FORM

The Hepatitis B Vaccine series is available at no cost to all employees and volunteers who have occupational exposure.  The form on the following two pages is used to document the consent or declination by the employee and the administering of the vaccine and follow-up testing.

  - The employee health nurse will provide information to employees prior to hepatitis B vaccinations addressing the safety, benefits, efficacy, methods of administration, availability and then administer the vaccine.

  - The vaccine is offered after the employee has received training and within ten (10) working days of the employee’s initial assignment to a job where exposure to bloodborne pathogens is possible.

  - Employees who do not respond to the primary vaccination series will be offered revaccination with a second three-dose vaccine series and be retested. Non-responders will be informed that they are considered non-immune.  Non-vaccinated employees, who have rendered assistance in any situation involving the presence of blood or OPIM, will have the full hepatitis B vaccination series made available to them within 24 hours.

  - If the CDC recommends a routine booster dose(s) of hepatitis B vaccine at a future date, such booster dose(s) shall be made available.

EMPLOYEE HEALTH RECORD: HEPATITIS B VACCINE COMPLIANCE FORM page 1

(Print Name)	_ (Birth date)	(SS#)  	                  has been determined to be at risk of contracting Hepatitis B as a result of job related duties. Hepatitis B vaccine must be offered within 10 days of employment. Begin by answering each of the three following questions. Do you have?

| Yes | No | Hypersensitivity to yeast or any vaccine component (aluminum hydroxide or |
| --- | --- | --- |
|  |  | thiomersal)? |
| Yes | No | Currently with moderate to severe acute febrile illness? |
| Yes | No | Current or history of Hepatitis B infection? |

If all the above questions are “No”, medical authorization per VAMC Memorandum No. 111-02 is given for the series of three (3) doses with recombinant vaccine. Most people have no side effects from the Hepatitis B vaccine. The injection is given in the muscle of the upper arm. This may cause soreness for a day or two at the injection site. As is the case of most drugs or vaccines, there is a possibility that allergic or more serious reactions will occur in a very small percentage of cases.

Effectiveness: Seroconversion, (immunity to Hepatitis B) following the series of three injections has occurred in more than 90% of those who received the vaccine. How long protection lasts after vaccination and the need for booster doses are not known. If you have any questions about the vaccine, please ask now or sign the declination below.

CONSENT: I have read the information about Hepatitis B, the Hepatitis B vaccine and have received the Hepatitis B Vaccine Information Statement. I have had an opportunity to ask questions and understand the purpose, benefits, and risks of Hepatitis B vaccination. I understand that I must have at least 3 doses of vaccine to confer immunity. However, as with all medical treatment, there is no guarantee that I will become immune or that I will not experience an adverse side effect from the vaccine. I request that it be given to me.

Consent of Recipient	Witness	Date

Vaccination is encouraged unless:

- documentation exists that the employee has previously received the complete series,

Yes	No	Date:

- antibody testing shows the employee is immune, and/or

Yes	No	Titer:

- Medical evaluation shows that vaccination is contraindicated.

Yes	No

Those employees meeting 1, 2 or 3 above are exempt and shall have this information documented in their medical record. If the employee initially declines hepatitis B vaccination but at a later date decides to accept the vaccination, the employer shall make available at no cost the hepatitis B vaccination at that time.

Employees who decline to accept hepatitis B vaccination offered by the employer sign a statement of the Hepatitis B Vaccine Declination (Mandatory) which is as follows;

DECLINATION: I understand that due to my occupational exposure to blood or other potentially infectious materials I may be at risk of acquiring hepatitis B virus (HBV) infection. I have been given the opportunity to be vaccinated with hepatitis B vaccine, at no charge to myself. However, I decline hepatitis B vaccination at this time. I understand that by declining this vaccine, I continue to be at risk of acquiring hepatitis B, a serious disease. If in the future I continue to have occupational exposure to blood or other potentially infectious materials and I want to be vaccinated with hepatitis B vaccine, I can receive the vaccination series at no charge to me.

_

Employee Signature	Date

_

Witness Signature	Date

VAMC EMPLOYEE HEALTH HEPATITIS B VACCINE RECORD page 2

Hepatitis B Vaccination shall be made available after the employee has received the training required and within

10 working days of initial assignment to all employees who have occupational exposure unless the employee has previously received the complete hepatitis B vaccination series, antibody testing has revealed that the employee is immune, or the vaccine is contraindicated for medical reasons.

Hepatitis B - The Disease

Hepatitis B is a viral infection caused by Hepatitis B virus (HBV) which causes death in 1-2% of patients. Most people with Hepatitis B recover completely, but approximately 5-10% become chronic carriers of the virus. Most of these people have no symptoms, but can continue to transmit the disease to others. Some may develop chronic active hepatitis and cirrhosis. HBV also appears to be a causative factor in the development of liver cancer. Thus, immunization against Hepatitis B can prevent acute hepatitis and also reduce sickness and death from chronic active hepatitis, cirrhosis, and liver cancer. The Hepatitis B vaccine is also known as the first anti-cancer drug by helping to prevent liver cancer.

Information about Hepatitis B Vaccine (Recombinant)

Hepatitis B vaccine is made using recombinant DNA technology with no serum from human donors. A high percentage of healthy people who receive three doses of vaccine achieve high levels of Hepatitis B Antibodies (anti- HBs) and protection against Hepatitis B. Full immunization requires three doses of vaccine over a six month period. There is no evidence that the vaccine has ever caused Hepatitis B or AIDS. However, persons who have been infected with HBV prior to receiving the vaccine may go on to develop clinical hepatitis in spite of immunization. The duration of immunity is unknown at this time but is probably long term.

Possible Side Effects

The incidence of side effects is very low. No serious side effects have been reported with the vaccine. A few persons experience tenderness and redness at the site of injection. Low grade fever may occur. Rash, nausea, joint pain, and mild fatigue have also been reported. The possibility exists that more serious side effects may be identified in the future.

First Series

Give 20 micro-gram dose IM (1” to 1.5”, 20 to 25 gauge needle) in the deltoid muscle at 0, 1 & 6 months. Have titer drawn one to two months after the 3rd dose.

If 6 months or less have passed since the 1st dose, continue the series, 2nd dose now, & 3rd dose in 4 to 6 months

| Date/Time Given VIS Date | Manufacturer Lot # /Expiration date | Site | Dose | Given By Whom | Date & Titer Results |
| --- | --- | --- | --- | --- | --- |
| 1st dose |  |  |  |  |  |
| 2nd dose |  |  |  |  |  |
| 3rd dose |  |  |  |  | * |

*After the third dose, if the Hepatitis B Surface Antibody (anti-HBs) test is POSITIVE, the person is immune, stop series.

Second Series for Non-Responder

If the Hepatitis B Surface Antibody (anti-HB) test is NEGATIVE after the third dose, the individual is not immune. Start the fourth dose and give 20 micro-gram dose IM (as above). Check the titer one month after the fourth dose. If individual does not convert, complete the second series. Recheck titer one to two months after the sixth dose. Approximately 20% to 50% of non-responding individuals have been shown to produce antibody after the fourth dose, and 50 to 75% will respond after six doses.

| Date/Time Given VIS** Date | Manufacturer Lot # /Expiration date | Deltoid Site | Dose | Given By Whom | Date & Titer Results |
| --- | --- | --- | --- | --- | --- |
| 4th dose |  |  |  |  | * |
| 5th dose |  |  |  |  |  |
| 6th dose |  |  |  |  | *** |

*If the Hepatitis B Surface Antibody (anti-HBs) test is POSITIVE, the person is immune, stop series.

** Vaccine Information Statements

*** If the Hepatitis B Surface Antibody (anti-HBs) test is Negative, do not repeat the series

ATTACHMENT I

RECORD KEEPING

  - All training records for employees will be maintained by the service of the employee. If requested, services shall furnish a copy of attendance for the annual refresher training to the infection control practitioner.

  - All exposure incident records shall be maintained by employee health. The employee records are not disclosed or reported without the employee's express written consent to any person within or outside the workplace, except as required by law.

  - All required medical records will be made available to OSHA upon request. VAMC shall maintain the records for at least the duration of the employee's employment, plus 30 years. Employee medical records shall be provided upon request for examination and copying to the subject employee or to anyone having written consent of the subject employee.  If VAMC ceases to do business and there is no successor employer to receive and retain the records for the prescribed period, VAMC shall notify OSHA at least three months prior to their disposal and transmit them to OSHA, if required by OSHA to do so, within that three month period.

  - The record shall include:

    - The name and social security number of the employee.

    - Routes and circumstances of all exposure incidents.

    - Results of source individual’s blood testing, if available.

    - Documentation that the employee was informed of the evaluation of the results of post-exposure medical evaluation and the need for follow-up shall be included.

    - A copy of the employee's hepatitis B vaccination status, including the dates of all the hepatitis B vaccinations, any medical records relative to the employee's ability to receive vaccination or signed declination statements, and a copy of all results of examinations, medical testing, and follow-up procedures, including the post-exposure prophylaxis provided if indicated:

    - An evaluation of any reported illness related to the exposure incidents.

    - The counseling information provided.

    - The employer's copy of the healthcare professional's written opinion.

    - A copy of the information provided to the healthcare professional.

### Attachment B:  BLOODBORNE PATHOGEN ECPOSURE MANAGEMENT PLAN Memorandum No. 111-02

###### BLOODBORNE PATHOGEN EXPOSURE MANAGEMENT PLAN

- PURPOSE:  To manage individuals (for example employees, trainees, residents, volunteers, contractors, visitors) exposed to bloodborne pathogens. This includes evaluation, testing, and post- exposure prophylaxis (PEP), if indicated for individuals exposed to bloodborne pathogens.

- POLICY:  Individuals with percutaneous, conjunctival, non-intact skin or mucus membrane contact with patient blood/body fluids is encouraged to report the exposure and go immediately or as soon as possible to Occupational Health (OH) or the Emergency Department (ED) for post exposure evaluation and follow-up.  Further management will depend on employee vs. non-employee status (see attachment A).  Abnormal testing results of the source or exposed person are communicated to that individual by the ordering practitioner/surrogate.

- RESPONSIBILITIES

  - OH physician/designee.

- Conduct initial evaluation, testing and prophylaxis during normal work hours (07:30 to 16:00).

- Follow-up evaluations, monitoring of laboratory tests during the period of chemoprophylaxis and post-treatment follow-up for a period of up to one year.

- OH Nurse.

- Maintain the OH Program and ensuring that all medical actions required are performed

- Manage blood/body fluid exposures.

- Maintain appropriate employee health records.

- Follow-up documentation of the route of exposure, the circumstances related to the incident, and the type and brand of device involved if applicable.

- If at a Community Base Outpatient Clinic (CBOC). Order the labs and inform the employee of the protocol follow-up.

  - Laboratory Information System.  Sends a CPRS view alert to the ordering practitioner for positive Hepatitis C results and calls the ordering practitioner/surrogate for positive HIV results so that the practitioner/surrogate may notify the patient.

  - ED physician.  To evaluate, counsel, and PEP of individuals during the evening hours (16:00 to 07:30), weekends, and holidays (see ATTACHMENTS A-B).

  - House Supervisor or designee. Ensure VA policy for Human Immunodeficiency Virus (HIV) Testing and Patient Follow-up, is followed when there is an employee exposure, the OH Nurse is not available, and verbal consent for testing from the { XE "HIV" } source person is required.

- Practitioners.

- Place orders for source patient testing (see Attachment A) and follows up the results of any orders they have placed.

- Provide HIV testing per VA policy titled HIV Testing and Patient Follow-up. { XE "HIV" }

- The ordering practitioner/surrogate is to notify the patient of all HIV, Hepatitis B or Hepatitis C testing results.

- On call Infectious Disease attending physician.  Is available for consultation and assistance.

- Pharmacy Service.  Ensuring an adequate supply of the medication for employees who will need PEP and individuals who will need PEP for HIV exposure.{ XE "HIV" }

- Exposed Individuals.  Report needle sticks/body fluid exposures and obtain follow-up of exposures. Evaluation and management of non-employee will vary depending on any written contracts.

- (CBOC) employees.  Notify the OH nurse of all exposures.  The OH nurse will order employee’s lab tests per policy, unless the OH nurse is not available, then the provider orders the employee’s lab tests (see Attachment A).   The appropriate CBOC staff is to obtain the verbal consent for HIV testing from the source patient and document that consent was obtained in CPRS. The provider of the source patient is to order necessary lab tests on the source patient (see Attachment A). All blood work drawn is to be sent to the Lexington VAMC lab, but may be drawn at the CBOC.

- Infection Control Committee.  Reviews the Bloodborne Pathogen Exposure Management Plan annually.

- PROCEDURES

- See ATTACHMENT A for detail of procedure to follow for immediate bloodborne pathogen exposure.

- Source individual tested positive.  For previously positive test results for Hepatitis B, Hepatitis C, or HIV{ XE "HIV" }, then those tests will not need to be repeated on that person.

- Employee tested positive.  For any previously positive results of those tests in 4.b., then those tests will not need to be repeated on that person.  If the employee refuses HIV{ XE "HIV" } testing, the blood sample obtained for Hepatitis B and Hepatitis C testing will be preserved for at least 90 days to allow the employee to decide if the blood should be tested for HIV serological status.  If testing refused, then after the 90 days, the blood sample may be discarded.The employee will be offered PEP{ XE "HIV" }.  There is no exclusion for PEP unless the HCW refuses.  In general, administration of the prophylactic regimen should be started as soon as possible (preferably immediately) after exposure because effectiveness may diminish over time.

- HBV: see attachment B (in accordance with current recommendations of http://www.cdc.gov/mmwr/PDF/rr/rr5011.pdf ).

- HIV: PEP for any HIV{ XE "HIV" } Exposure or from a source patient for whom there is reasonable suspicion of HIV infection (see Table  below)

- REFERENCES: US Public Health Service Guideline - Updated US Public Health Service Guidelines for the Management of Occupational Exposures to Human Immunodeficiency Virus and Recommendations for Postexposure Prophylaxis ICHE, September 2013; OSHA Bloodborne Pathogen Standard; 29 CFR 1910.1030.

http://www.osha.gov/pls/oshaweb/owadisp.show_document?p_table=STANDARDS&p_id=10051; CDC, Revised Recommendations for HIV Testing of Adults, Adolescents, and Pregnant Women in Health-Care Settings{ XE "HIV" }{ XE "AIDS" }, MMWR 2006, 55:RR:14, pgs 1-17; CDC, Updated

U.S. Public Health Service Guidelines for the Management of Occupational Exposures to HIV{ XE "HIV" } and Recommendations for Post-exposure Prophylaxis. MMWR September 30, 2005; 54(RR- 9), pp 1-24; Ordering and Reporting Test Results, VHA Directive 2009-019, March 24, 2009

- FOLLOW-UP RESPONSIBILITY:  Infection Control, OH, and Safety Office.

- RECERTIFICATION: On or before September 24, 2014.

Emma Metcalf, MSN, RN Director

Attachments

### Attachment C:  ERGONOMICS PROGRAM MEMORANDUM NO. 138-01

VA MEDICAL CENTER					MEMORANDUM NO. 138-01

LEXINGTON, KENTUCKY					NOVEMBER 27, 2012

ERGONOMICS PROGRAM

- PURPOSE:  To provide guidelines, procedures and assign responsibilities for the establishment of an ergonomics program for this medical center.

- POLICY: To identify and eliminate ergonomic job-related hazards within the medical center to the extent feasible. When hazards cannot feasibly be eliminated, the most practical method of control will be implemented.  The goal of an ergonomics program is to make the workplace fit the employee.

- DEFINITION:  Ergonomics is the science that seeks to adapt the job and workplace to the worker by designing tasks and tools that are within the worker's capabilities and limitations.

- RESPONSIBILITIES

  - Service Chiefs.

- Provide the resources to ensure staff is aware of the basic ergonomic principles and utilize equipment properly. Attachments 1 and 2 provide algorithms to assist in selecting the appropriate equipment to move patients.

- Ensure that employees are trained in proper body mechanics, the use of mechanical lifting devices and job-specific procedures to reduce excessive body stress.

  - Safety Office. Coordinates the ergonomics program, to include:

- Develop and provide awareness training

- Evaluation of work areas for ergonomic hazards.

- Review of injury/illness records for trends indicating potential ergonomic hazards.

- Recommendation of controls that will mitigate identified ergonomic problems.

- Provide technical recommendations on workstation adjustability requirements prior to selection and purchase.

  - Engineering Service

- Ensure new workstations and supporting furniture are adjustable to support a variety of employees.

- Ensure keyboard trays are included in new work stations and added to existing stations as a basic setup.

  - Information Management Services.

- Provides an appropriate, soft wrist rest for keyboard trays and point device (mouse, track ball, touchpad, etc) at every computer workstation.

- Maintains a supply of alternate ergonomic keyboards to support keyboard modifications when recommended.

  - Supervisors.

- Ensure proper setup of work stations and use of ergonomic equipment

- Submit a work order to request ergonomic evaluation when employee(s) indicates symptoms of soft tissue discomfort or requests an evaluation.

  - Employee Health.  Notifies the Safety Office when employee displays symptoms related to soft tissue discomfort.

  - Employees.

- Notify their Supervisor and/or Employee Health of soft tissue discomfort to ensure early intervention and preventive actions through an ergonomic evaluation.

- Implement the work station modifications recommendated during the  ergonomic evaluation.

- Use safe patient handling equipment provided to move and position patients as a first choice.

- PROCEDURES

  - Training and Awareness:

- Employees will complete ergonomic awareness training during new employee orientation and through TMS, Office Ergonomics SAH0457 or Ergonomics in the Workplace, periodically.

- Employees will complete training on patient lift devices to ensure proper use in patient handling.

  - Evaluation of Injuries/Illnesses:

- Employee Health will evaluate employee injuries/illnesses to determine if they are ergonomically related.  If the injury/illness is determined to be ergonomically related, the Employee Health nurse will notify the Safety Office.

- The Safety Office will conduct an ergonomic evaluation for the employee identified by Employee Health.

- The Safety Office will review incident reports for trends indicating ergonomic hazards.

  - Evaluation of Workplace:  The Safety Office will conduct ergonomic evaluations as requested. or based on incident reports. Employees are encouraged to report ergonomically related hazards to their supervisor so that early intervention can be implemented. The Safety Office will recommend the engineering controls (equipment, furniture, supplies) or administrative controls (work station setup, use, and work habits) that will assist in mitigating ergonomic problems.

  - Furniture: New furniture or relocations must be planned through the Interior Designer who will design the work station and  be responsible for initiating all furniture requests (for example, chairs, workstations, keyboard trays, etc.). The Interior Designer will work with the Safety Office to ensure purchases will provide the flexibility to adjust for the majority of employees.

- REFERENCE:  OSHA Ergonomic Guidelines, 1990 and VISN 8 Patient Center of Inquiry, James

A. Haley Veteran Center, Tampa Fl.

- FOLLOW-UP RESPONSIBILITY: Chief Engineering Service (138).

- RECERTIFICATION:  On or before November 27, 2017.

DeWayne Hamlin

Director Attachments (2)

ATTACHMENT B

ERGONOMIC PRINCIPLES

- Back Injury Prevention

  - Use proper mechanical devices correctly to lift/transfer heavy loads.

  - Seek adequate assistance when lifting or moving heavy or awkward loads.

  - Use proper body mechanics, mechanical lifting devices and job-specific procedures to reduce excessive body stress.

  - When working in/at a computer workstation or sedentary job, use a chair which supports the lower back; ensure the chair is adjusted so that feet are flat on the floor, lower back is supported, seat pan leaves two to three finger width between the back of the calf and front edge of the seat and the arms are adjusted to the user.  Look at the instructions for the chair to be familiar with the adjustments.

  - When standing for prolonged periods, change positions frequently by putting body weight on one foot and then the other.  Consider use of appropriate mats to reduce stress from hard floor.

  - Plan lifts by testing the weight of the load before attempting the lift.  Lift with the legs and keep back straight.  If container is awkward in size, configuration or weight distribution, get assistance to move the object.

  - Maintain good posture throughout the day. Maintain personal physical fitness. Both are paramount to proper ergonomic practices.

- Computer workstations:

  - Monitors should be positioned to avoid glare from overhead lighting or light from windows and eliminate contrast of the screen with background lighting and reflections.

  - Monitors should be adjustable and must be adjusted by individual users.

  - Monitor screens should be arms length from the user face with the top line of print positioned at eye level.  Check this by having a coworker check to see if the jaw line is parallel to the floor.

  - Position the keyboard so that the user's hands and forearms are straight and parallel to the floor and the upper arm is at least a right angle to the forearm. Keyboard and pointing device wrist rests should be utilized to lessen the stress to the wrist.  Rests may be obtained from IM.

  - Chair should be adjustable and allows the feet to rest on the floor. The users back must be firmly against the chair for maximum support with shoulders straight and feet on the floor. Keep wrists straight, stop occasionally to stretch and shake out the wrists and fingers.

  - Document holders, when used, should be the same height as the screen.

  - Foot rests, when recommended, should be used periodically throughout the day to provide adjustments to user’s posture.  The chair must be positioned with the feet flat on the floor. This may require adjustments to the work surface.

  - Keep frequently used material close at hand and infrequently used materials on shelves or distant from the user.

  - Alternate work with other tasks so not to remain in one position too long. Every 55 minutes, users should stand up and stretch. Every 30 minutes, users should look away from the monitor for 30 seconds to give their eyes a break.

  - Report any pain, swelling or numbness in the wrists or hands to your supervisor immediately.

### Attachment D:  MEDICAL EQUIPMENT MANAGEMENT PROGRAM Memorandum No. 138-09

VA MEDICAL CENTER						MEMORANDUM NO. 138-09 LEXINGTON, KY 40502						NOVEMBER 27, 2012

MEDICAL EQUIPMENT MANAGEMENT PROGRAM

- PURPOSE:  To continue the Medical Equipment Management Program which shall assess and control the clinical and physical risks of patient care equipment used in the medical center.

- POLICY:  The Medical Equipment Management Program shall be used as a method to evaluate, implement, monitor and improve upon the clinical and physical risks of fixed and portable equipment for the diagnosis, treatment, monitoring and care of patients, and of other fixed and portable electrically powered equipment.  Building Service Equipment, which is part of the Utilities Management Program, as well as administrative/ADP and non-patient care equipment is excluded from the Medical Equipment Management Program.

- DEFINITIONS

  - Component of a Major System.  Items which have been assigned a preventive maintenance (PM) number, but are used only in conjunction with other equipment items which have been evaluated for inclusion in the program.

  - Patient Care Equipment. Any equipment used in monitoring, diagnosis or treatment of patients.

  - Electronic Protected Health Information (EPHI). Protected Health information under HIPAA.

- RESPONSIBILITIES

  - The Environment of Care Committee will:

- Review and analyze summaries of maintenance, training and problem resolution while recommending corrective actions.

- Track equipment problems until resolution.

- Review and analyze established performance standards and indicators to ensure the effectiveness of the program.

  - Service Chiefs will:

- Document the orientation and, at least annually thereafter, continue education for users and maintainers of equipment in the program.  Service Chiefs should decide which types of equipment, in the program or not, require annual training.  For example, equipment used on a daily basis and in the program may not require annual training.  The Environment of Care Committee may decide that, due to information provided through performance standards, a particular device may require annual training. The appropriate service Chief(s) will be notified of this recommendation and be required to provide such training.

- Document and send to the Biomedical Engineer (138-CDD), evidence of equipment failures or user errors that have or may have an adverse effect on patient or employee safety/quality of care. Failures, which are suspected to have caused a patient’s death or serious injury or illness, must be reported to the Biomedical Engineer within 24 hours. The Biomedical Engineer will report these incidents to the Medical Device Event Review Team for investigation (see VAMC Memorandum No. 90-04, Medical Device Incident Reporting). Such reports must be reviewed and sent to the manufacturer/FDA, per the Safe Medical Devices Act of 1990, within 10 working days.

- Inform Engineering Service of the arrival of all equipment at this facility, if it has by-passed the Warehouse or Personal Property Management (e.g., leased, rented, borrowed, loaned or staff/patient- owned equipment).

  - Acquisition and Logistics Service (A&LS) will:

- Hold all incoming electrical equipment for Engineering Service inspection.  If there is a rush order to have a piece of equipment checked, A&LS will contact Engineering Service.  Items, which cannot be inspected due to installation limitations or other factors, will be inspected by Engineering Service after

installation.

- Use the Automated Equipment Management System/Medical Equipment Reporting System (AEMS/MERS) to document equipment deletions and new equipment acquisitions to the equipment inventory file. Certain devices considered as components of a major system, and requiring inclusion into the program, will be included in the equipment inventory by Engineering Service.

  - Supervisory Biomedical Engineer will:

- Ensure incoming equipment inspections are performed and new equipment is evaluated for inclusion in the Medical Equipment Management Program.

- Ensure periodic PM inspections (not to exceed two years) are performed for all equipment in the Medical Equipment Management Program. All work orders against the equipment shall be documented in the AEMS/MERS computer system.

- Develop and maintain a unique inventory of all equipment that is current within 6 months.

- Monitor, report and recommend process improvements of the Medical Equipment Management Program performance standards and indicators to the Environment of Care Committee.

- Ensure locally developed procedures are in place for each type of device in the Medical Equipment Management Program.

- Ensure equipment maintenance technicians receive appropriate training to maintain equipment in the Medical Equipment Management Program.

- Ensure emergency procedures are in place for certain critical items in the Medical Equipment Management Program.

of.

- Ensure EPHI data is removed from storage media of all medical equipment before it is disposed

- PROCEDURES

  - Criteria for Inclusion of Equipment into the Medical Equipment Management Program: All equipment, regardless of ownership, will be evaluated against the attached criteria for inclusion into the Medical Equipment Management Program.  The review will be documented and will encompass equipment function, risk, maintenance requirements and work history factors attendant with its normal expected use (Attachment A).  Documentation of the evaluation process shall be maintained in the Engineering Service/Biomedical Engineering office.

  - Biomedical Engineering will evaluate all new medical equipment. Equipment with a cumulative score of nine or above will be included in the Medical Equipment Management Program unless its exclusion is warranted as a component of a major system. Scoring results will be indicated in the AEMS/MERS equipment inventory field labeled, "Criticality," and equipment meeting inclusion criteria will have the "JCAHO" field noted as "YES".  This equipment will have certain performance characteristics verified during their PM inspections. Equipment with scores of eight or below may still require at least annual electrical safety inspections (per NFPA). Their JCAHO fields in the equipment inventory will be designated as "NO", but will be scheduled for this required inspection.

  - Incoming Inspections:  All equipment will be tested prior to initial use, whether or not it is included into the Medical Equipment Management Program.  All equipment evaluated and placed into the Medical Equipment Management Program will receive at least annual inspections, unless specific documentation warrants a more frequent or extended maintenance schedule. This incoming inspection will also include, if appropriate, an electrical safety inspection and the performance procedure to be used in subsequent inspections.  Documentation of the incoming inspection's performance will be a work order referenced to an equipment inventory entry.  An exception will be rental equipment necessary for patient treatment that is brought into the medical center during non-administrative hours and shall be inspected on the first working day after its arrival (see VAMC Memorandum No. 138-25, Technical/Safety Testing of Incoming Equipment).

  - Equipment Inventory:  The Biomedical Engineering Section, in conjunction with Asset Management Section, A&MMS, will maintain an inventory of all equipment in the Medical Equipment Management Program and other equipment not included in the program. This database will be maintained using AEMS/MERS software, which integrates the inventory, work order requests and purchase requests for equipment. The inventory will be continuously updated, but a formal review of accuracy will be made on a 6-month basis.

  - Frequency:  Equipment in the equipment management program will receive a PM frequency based on the manufacturer and industry standards. Any significant changes to a PM frequency shall be reported to the Environment of Care committee for approval.

  - PM Completion:  The PM completion goal for life support equipment medical will be 100% (Attachment B).  The PM completion goal for non-life support medical equipment will be greater than 95%.  Biomedical Engineering will provide off-schedule PMs on deferred equipment as identified by staff.

  - Re-inspection labels:   Biomedical engineering labels all medical equipment upon incoming inspection. There are three different labels that are used to indicate the preventive maintenance strategy followed. A green label indicates that the equipment is included in the Medical Equipment Management Program and a PM will be performed accordingly. A blue label indicates that the equipment is non-VA owned and specifies whether it is rental, loaned, or demo. Re-inspection and inclusion into the PM program is required if the equipment remains in the facility for longer than 90 days. A white label indicates that no PM is necessary for the device. All medical equipment will have one of these labels affixed prominently to the device and will indicate to staff that the equipment is safe to use. Staff should notify biomed if they locate a piece of medical equipment that has an overdue PM sticker, as the equipment needs inspection by biomed to ensure safety.

  - Removal of EHPI data before disposal: Biomedical engineering will evaluate all medical devices for EHPI data before disposal.  Follow Biomedical “Disposing of storage media containing sensitive information” SOP for those medical devices that contain EHPI data. (Attachment C).

  - Testing Procedures: Testing procedures will be locally developed for all equipment in the Medical Equipment Management Program.

  - Training:  It is encouraged that equipment training be focused on what items the employee and supervisor assess require such training, to make these training events meaningful. Records will be maintained by the service of the employee using or maintaining the equipment. Engineering Service is available to assist in user training upon request. Certain training events may be required to be held for users and maintainers of equipment by the Environment of Care Committee. This will be done when information is presented (trends in user errors, repair frequency, critical equipment failures, etc.) which indicates a need for such training.

  - Performance Standards:  Performance Standards and applicable Performance Indicators are established and monitored to benchmark the state of the Medical Equipment Management Program.  The results are reported to the Environment of Care committee.

  - Rented, Leased, Loaned, Cost per Test, Trial and Personal Property: Certain items are rented from local vendors. These vendors are responsible for all equipment maintenance. Biomedical Engineering may at any time request maintenance records for any medical equipment rented from any vendor.  Leased or cost per test equipment will be initially inspected, evaluated and, if appropriate, incorporated into the program.  This equipment often has maintenance as part of the agreement. Such maintenance, including scheduled PM, will be tracked by the program. Medical equipment that is being trialed will be initially inspected and evaluated. This trial will comply with Lexington VA memorandum 90-06. Personally owned medical equipment is also initially inspected, evaluated and incorporated.  Such equipment should be used on an infrequent basis; new equipment requests should be made to obtain the needed equipment by the medical center.  Loaned equipment will be initially inspected and

documented with a work order. The definition of loaned assumes a short time period.  If the duration of the loan is over 90 days, the equipment is considered leased at no cost and incorporated, if applicable.

- REFERENCES: The Comprehensive Accreditation Manual for Hospitals; NFPA 99; VA Joint Commission Environment of Care Guidebook,; VA Equipment Management Guidebook; Lexington VA memorandum 90-06 (Testing and Demonstration of Supplies and Equipment); Engineering Service Memorandum 31( Preventative Maintenance of Medical Equipment); VA Directive 6550.

- FOLLOW-UP RESPONSIBILITY: Chief, Engineering Service (138).

- RECERTIFICATON: On or before November 27, 2017.

DeWayne Hamlin Director

Attachments “3” DISTRIBUTION:  "B"

ATTACHMENT A

MEDICAL EQUIPMENT MANAGEMENT PROGRAM INCLUSION/EXCLUSION SCORING SHEET

The following evaluation (scoring) of patient/nonpatient care equipment will determine inclusion in the Medical Equipment Management Program.  Equipment classification will be based on the following criteria:

- Equipment Function:  Determine if the equipment is of a critical nature, necessary for life support, diagnosis, or monitoring.

  - Life Support or Surgical/Intensive Care	5

  - Physical Therapy/Treatment	4

  - Monitoring-Intensive Care/Other	3

  - Laboratory-Analytical	2

  - All Others	1

- Equipment Risk: Determine the severity of risk during use of equipment, based on documented histories of the equipment device type.

  - Potential for death to patient/employee	5

  - Potential for injury to patient/employee	4

  - Potential for misdiagnosis or inappropriate therapy	3

  - Potential for equipment/fire damage	2

  - No known potential of risk	1

- Equipment Maintenance Requirements: Determine frequency of use and manufacturer's, VA or NFPA recommendations.

  - Monthly PM required	5

  - Quarterly PM required	4

  - Semi-Annual PM required	3

  - Annual/Bi-annual PM required	2

  - No PM required	1

- Location of Equipment Usage: Determine history of risk based on location.

  - Anesthetizing location	5

  - Critical Care location	4

  - Wet location	3

  - General Patient Care location	2

  - Non-Patient location	1

TOTAL:

IF TOTAL SCORE IS LESS THAN 9,

Does specific device history warrant inclusion?   Yes No Should equipment be included for required

electrical safety inspections?	Yes  No Should equipment be excluded because it is

inspected as part of another system?	Yes No

Equipment category:

(9 and above - Include)	INCLUDE/EXCLUDE

Reviewed by:

Date:

ATTACHMENT B

LIFE SUPPORT EQUIPMENT

Addressing JCAHO EC.2.04.03, a life support list has been established.  Life Support Equipment will be given priority during PM inspections.

The definition of Life Support equipment is equipment that provides all or most of the items for maintaining life.

- Anesthesia Machine

- Aortic Balloon Pump

- Defibrillator-AED

- Defibrillator-Biphasic

- Heart Lung Machine

- Pacemaker-Noninvasive

- Ventilator

ATTACHMENT C

DISPOSING OF STORAGE MEDIA CONTAINING SENSITIVE INFORMATION

Ref. VISN 9 Directive 10N9-39-02

- Medical device with hard drive containing sensitive information that can run Data Eraser

- Run Data Eraser program from floppy drive.

    - Input ID# of medical device.

    - Input serial# of hard drive.

    - Run Triple pass (DOD Specs.).

  - Complete forms:

    - Print VA Form 0571 (IT Equipment Sanitization Certificate).

    - Print DE_CERT form from Data Eraser floppy

    - Sign forms.

    - Supervisory Biomedical Technician signs in “supervisor of staff official”.

- Contact ISO via email to schedule verification.

- ISO verifies that no sensitive data is left on drive.

- ISO signs VA Form 0571 (IT Equipment Sanitization Certificate)

- Make 2 copies of all signed forms and distribute to:

- Original kept by supervisory biomedical technician.

- Copy to ISO.

- Copy to AMMSL (keep with PC or hard drive).

- Ensure no media is left in Tape/CD/DVD/Floppy drives.

- Take medical device to warehouse along with copy of documents.

- All other media from Medical equipment:

- Put in box marked “Storage media for disposal” box located in Biomed.

- These items will be taken to IM when they are sending out their media for destruction.

- Examples are:

    - Hard drives that can’t be wiped.

    - Removable media (backup tapes, optical disks, CD’s, DVD’s, etc.)

### Attachment E:  FORMALDEHYDE SOP-ENG-138-12

VA MEDICAL CENTER	SOP-ENG-138-12

LEXINGTON, KENTUCKY	OCTOBER 1, 2012

FORMALDEHYDE

- PURPOSE:  To establish a program for controlling occupational exposures to formaldehyde. This program involves the use of engineering and administration controls, personal protective equipment, medical surveillance and employee training.

- POLICY:  Appropriate measures will be taken to evaluate and control the use of formaldehyde at this medical center to limit occupational exposures and reduce the risk to personnel and the environment.

- PROCEDURES:

  - An evaluation will be conducted of all operations that involve the use of formaldehyde to determine the potential for exposure to personnel through inhalation and contact with skin and eyes.

  - The potential for exposure through skin contact will be assumed whenever personnel handle liquids containing one percent or more formaldehyde.  It will also be assumed that the potential exists for eye contact during operations involving the use of these materials outside of a containment device (e.g., fume hood) in quantities large enough to cause a splash or vapor.

  - Evaluation of exposure through inhalation is based on the likelihood that airborne formaldehyde vapor or mist is being released to the work environment in concentrations greater than the action level of 0.5 parts per million (ppm) averaged over an 8-hour workday or greater than the short-term exposure limit (STEL) of 2 ppm during any 15 minute period.

  - Where the potential for skin contact exists, personnel will wear rubber gloves and, if necessary, aprons.  Butyl and nitrile rubber gloves provide the greatest protection from formaldehyde.  Aprons are necessary when handling larger volumes of liquid outside of a containment device (e.g., when dumping formalin solution in Pathology, filling large containers in the veterinary medical unit and transporting waste containers).

  - Chemical splash goggles must be worn as needed to protect eyes from splash, spills and vapors.  Face shields should be worn in addition to, rather than in place of, chemical goggles.

  - Where it is determined that the potential exists for inhalation exposure, air sampling will be performed to determine exposure levels.  Sampling will include full shift monitoring to ensure compliance with the action level and 15 minute sampling to determine compliance with the  STEL (Short Term Exposure Limit).  STEL sampling is important where tasks are performed  that involve brief but intense formaldehyde exposures.

  - Examples of operations where periodic air sampling may be necessary include the following:

    - Dumping formalin solution in Histology.

    - Grossing procedures in Pathology.

    - During necropsy in Research.

  - Results of air sampling will be provided in writing to the employee(s), service, employee representatives and Occupational Health.  Also, all documentation of employee monitoring will be maintained in the Safety Office.

  - Where air sampling indicates that exposures may reach or exceed the action level or STEL, an evaluation will be performed to validate the findings and determine measures necessary to reduce formaldehyde concentrations.

  - In general, tasks involving formaldehyde should be performed using local exhaust ventilation where feasible.  These ventilation devices will be checked periodically to ensure adequate airflow is provided to control and remove vapors from the worker’s breathing zone. Examples of ventilation devices include fume hoods, gross lab and necropsy table exhaust.

  - In areas and operations where the use of local exhaust ventilation is not feasible, personnel should ensure adequate general ventilation is provided.  The use of formaldehyde in confined areas where ventilation is limited should be avoided.

  - In the event of an accident involving skin and eye contact, personnel will use quick drench showers and eyewash stations and seek emergency medical care via Occupational Health or the Emergency Department. All Research labs are equipped with eyewash stations at the sinks and have showers installed at numerous locations in the hallways.  Eyewash stations and a quick drench shower are also installed in Pathology service.  If clothes become contaminated with formaldehyde, they are to be removed and laundered on site.  Contaminated clothes will not be worn home or outside the work area.

  - Spill, Leak, and Disposal Procedures

For small containers, place the leaking container in a well ventilated area (i.e. fume hood). Utilizing proper PPE, take up small spills with absorbent material (spill kits such as Spill-X-FP or Biotex neutralizing pads) and place the waste into properly labeled containers for later disposal.

For large spills, contact the Safety Office (ext 4176 or 3365) or the Police Service during non administrative hours. Leave the area immediately and prohibit/extinguish all flames in the hazard area. Designated/trained persons may stop the leak and shut off ignition sources if these procedures can be done without risk. The hazard area will be isolated and entry denied except for designated response personnel protected by suitable protective clothing/respirators.

Dispose of waste containing formaldehyde in accordance with applicable local, state, and Federal law and in a manner that minimizes exposure of employees at the site and of the clean-

up crew.

  - Medical Surveillance Program:

- Employees to be included in a medical surveillance program include:

  - Employees who are or may be exposed to formaldehyde at or above 0.5 ppm calculated as an 8-hour time-weighted average ten or more days per year.

  - Employees exposed to formaldehyde at or above 2 ppm calculated on a 15-minute basis (STEL).

  - Employees who develop signs and symptoms of overexposure to formaldehyde in emergencies.

- Medical surveillance must follow the protocols established in 29 CFR 1910.1048. Additional guidance is provided in the Clinical Occupational Health Guidebook (http://vaww.ceosh.med.va.gov/ceosh/Guidebooks/COH2009/2009_COH.htm).

  - Formaldehyde training will be provided to the employees by their service as part of the Hazard Communication Program.  Substance-specific formaldehyde training will be customized to each work area, be provided to all affected employees and include all required elements. (HazCom Training Outline – Formaldehyde and OSHA 29 CFR 1910.1048 Appendix A.)

- RESPONSIBILITIES:

- The Safety Office is responsible for:

    - Conducting evaluations of procedures and tasks involving the use of formaldehyde, determining the need for air monitoring and personal protective equipment and conducting air monitoring.

    - Receiving and interpreting the results of air monitoring, determining compliance with established standards and providing written results to the employee(s), service, and Occupational Health and employee representatives.

    - Performing periodic surveys to check the effectiveness of controls.  These surveys include taking air velocity measurements in fume hoods or other mechanical ventilation systems, checking the availability of personal protective equipment and inspecting emergency showers and eyewashes.

    - Responding to spills/emergencies, performing clean-up or calling in outside spill response team.

    - Providing information, in addition to air monitoring results, to Occupational Health pertinent to medical surveillance.

    - Conducting special evaluations of procedures, tasks and situations that have resulted in overexposure, medical removal or employee restrictions to determine whether changes or additional controls will result in reduced exposure.

    - Evaluating compliance with the Hazard Communication Program requirements regarding formaldehyde.

- Services will:

- Ensure formaldehyde containing materials are properly labeled, material safety data sheets are available and substance-specific training on formaldehyde is provided to affected employees.

- Assist the Safety Office in identifying particular operations involving formaldehyde within their service that may require further evaluation.

- Supervisors will ensure that:

- All employees follow proper procedures when working with formaldehyde.

- Personal protective equipment is available and worn as required.

- Substance specific formaldehyde training, which includes health hazard data, proper safeguards, spill recognition and measures to limit exposure, is conducted for new employees, when new methods or materials are introduced, and annually.

- REFERENCES:  OSHA 29 CFR 1910.1048 Formaldehyde Standards, VAMC Memorandums 00-38 Environment of Care Management Program, 138-07 Hazard Communication and 138-20 Hazardous Waste Management.

- FOLLOW-UP RESPONSIBILITY:  The Chief, Engineering Service has follow-up responsibility.

- RESCISSION:  This memorandum will be due for recertification on October 1, 2012, in accordance with procedures outlined in VAMC Memorandum No. 001-1.

Michael M. Young, P.E., CHFM Chief, Engineering Service

Attachment

ATTACHMENT A

VA MEDICAL CENTER	SOP-ENG-138-12

LEXINGTON, KY	OCTOBER 1, 2012

HazCom Training Outline – Formaldehyde (Covered by OSHA, 29 CRF 1910.1048)

- Training must be provided:

  - To all employees who:

    - Use formaldehyde, including mixtures or solutions containing greater than 0.1 percent formaldehyde.

    - Work in areas where the potential exists for exposures to reach or exceed 0.1 parts formaldehyde per million parts of air (ppm).

  - At the time of initial assignment, annually thereafter, and whenever a new hazard is introduced.

- General information on formaldehyde:

- Review of MSDS.

- Appendix A of 29 CFR 1910.1048.

- Health hazards:

    - Exposure from inhalation of vapors, ingestion and contact with skin and eyes.

    - Suspected human carcinogen – believed to cause cancer of lung, nasopharynx, oropharynx and nasal passages.

    - Sensitizer – causes allergic reactions to skin and respiratory system, in previously exposed personnel, in low concentrations.

    - Irritant to skin, respiratory tract, eyes and throat.

    - Acutely toxic – immediate adverse health effects include difficulty in breathing, tearing eyes, numbness and discoloration of the skin, burning nose and throat.

    - As workers become adapted to the presence of formaldehyde, its warning properties become less reliable for detection.

- Physical hazards – moderate fire and explosion hazard.  Reacts with oxidizers, nitrogen dioxide, nitro methane, perchloric acid, aniline or peroxyformic acid to yield explosive compounds.

- Exposure Monitoring:

- Full shift personal air sampling is conducted to determine exposure concentrations relative to the action level of 0.5 ppm.

- Short-term (15 minute) sampling is conducted during high-risk operations (e.g., grossing, pouring formalin solution) for evaluating exposures relative to short-term exposure limit (STEL) of 2 ppm.

- Medical Surveillance:

- Provided to employees exposed above the action level or STEL.

- Provided to employees who develop signs and symptoms related to formaldehyde exposure or those exposed during emergencies.

- Personnel instructed to report any such signs/symptoms to supervisor.

- Operations which may result in exposures include:

- Tissue fixation (i.e., grossing operation) in Pathology.

- During necropsy in VMU-Research.

- Control Measures:

- Mechanical ventilation is the most common form of engineering control used to control vapors.  Examples include chemical fume hoods, lateral exhaust slotted hood, etc.  General ventilation is used to maintain labs under negative pressure relative to surrounding areas (i.e., hallways).

- Administrative controls include filling containers with only minimal quantities of formalin to perform task, keeping container lids closed and substituting less toxic materials in place of formaldehyde whenever possible.

- Personal Protective Equipment – impervious gloves (butyl or neoprene), face shield, aprons, respiratory protection (last resort) etc.

- Emergency Response:

- Take steps to stop flow and isolate if it can be performed safely.

- Only very minor spills should be cleaned up by employees.  Use absorbent material (e.g., Spill-X-FP or Biotex neutralizing pads) to contain and polymerize spill.

- Notify supervisor and Safety Office (ext 4176 or 3365).

- If the spill cannot be managed by the Safety Office, they or an assigned designee will call for spill response either from a contracted vendor or 911 (Lexington Fayette Urban County Hazmat Team).

### Attachment F:  NOISE CONTROL AND HEARING CONSERVATION PROGRAM  SOP-ENG-138-02

VA MEDICAL CENTER	SOP-ENG-138-02

LEXINGTON, KENTUCKY	OCTOBER 1, 2012

NOISE CONTROL AND HEARING CONSERVATION PROGRAM

- PURPOSE: To restate policy regarding the Hearing Conservation Program for personnel occupationally exposed to potentially hazardous noise levels.

- POLICY: Where it is determined that personnel are exposed to hazardous noise levels, as defined by OSHA Standard 29 CFR 1910.95, these employees will be included in the Hearing Conservation Program. It shall be the policy of the Medical Center to ensure employees are informed of work areas where hearing protection devices are required or recommended, devices are provided for use in those areas at no cost to the employees, audiometric testing is completed, and testing is completed on the hearing conservation program.

- RESPONSIBILITIES:

  - Safety Office will:

(1) Coordinate the Noise Control and Hearing Conservation Program, for monitoring, and evaluation of noise exposure and developing training tools on hearing protection.

2)  Perform evaluations to determine noise levels and personnel exposure levels to identify hazardous noise-producing equipment and personnel to be included in on the conservation program.

- Provide written notification of the hazardous noise survey to the service of the affected employees, the Employee Health nurse and Audiologist. The written notification will either designate individuals by name, work area, or job classification.

- Provide written notification to employees (via the service) of the results of the noise survey.

- Shall maintain the Federal agency OSHA 300 log on reportable hearing loss when an employee’s hearing test shows a standard threshold shift (STS) in overall hearing.

- Service Chiefs will:

- Ensure employees are informed of the written results of noise evaluations performed by the Safety Office.

- Provide employees or their affiliates with the opportunity to observe any noise measurements conducted.

- Take appropriate measures to reduce occupational noise exposure in the work area.

- Compile a list of affected employees (with help from the Safety Office if designation is by other than name) and contact Audiology to schedule baseline, annual and job termination audiograms for all affected employees within their service.

- Audiology Staff will:

- Maintain responsibility for audiometric test requirements for performing hearing screening on an annual basis. Determine if further evaluation and follow-up is necessary and provide an oral review of testing procedures and results to the employee at the time of testing.

- Forward results of all audiometric testing, including a threshold shift, to the Employee Health Nurse or designee and provide written notification to employees in the event of a threshold shift within 21 work days of the determination. Provide written notification to the Safety Officer in the event of an STS for inclusion in the OSHA 300 log.

- Conduct hearing conservation training and fit personnel with hearing protection. See training outline (Attachment B).

- Supervisors will:

- Oversee the daily, appropriate use of hearing protection in areas where protection is required.

- Ensure staff completes their annual testing and training to assure hearing conservation.

- Notify the employee of the need to avoid high levels of non-occupational noise exposure 14 hours before the audiometric examination.

- Employee Health will:

- Be Responsible for monitoring audiometric tests and record retention of baseline and annual audiograms.

  - Update employee’s demographical data and most recent noise exposure assessment, in accordance with [29 CFR 1910.95 (m) (2) (ii)], during each testing session on the electronic VA Form 10-2364 electronic audiogram.

  - Retain employee noise exposure measurement records for two years.

  - Retain Audiometric test records for the duration of the affected employee’s employment.

- Review and evaluate audiograms as well as confer with effected employees to determine work relatedness.

  - Consult with the Audiologist to determine if the standard threshold shift is work related.

  - Take appropriate measures to reduce occupational noise exposure in the work area.

- Provide All access to records, when applied under section 29 CFR 1910.1020 (a)—(e) and (g), upon request to employees, former employees, representatives, designated by the individual employee, and the Assistant Secretary.

- Transfer all records required to be maintained, to the successor employer to be retained for the duration of the period prescribed if the employer ceases to do business or if the employee relocates to a different employment.

- Engineering Services will:

- Utilize available technology to control/reduce noise exposure.

- Maintain the records for the last acoustic or exhaustive calibration of the audiometer on the equipment and file the Certificate with the Biomedical Engineering Department.

  - Employees working in hazard noise areas will:

- Wear authorized and properly fitted hearing protective devices.

- Report for all examinations and health education training.

- Report any hearing problems.

- Comply with recommendations and provisions of this program.

- PROCEDURES

- The Safety Office will perform noise measurements in areas where it is suspected that levels may exceed 85 decibels (dB), A-weighted (dBA). Where measurements indicate that personal exposures may exceed the 85 dBA limit as an 8 hour Time Weight Average (TWA), noise dosimetry is conducted to determine the equivalent continuous sound level (ECL). When circumstances make monitoring generally inappropriate, the employee will wear a personal noise monitoring dosimeter to evaluate noise level results. Evaluations may be initiated when:

- Environmental monitoring surveys indicate significant noise levels may be present.

- Employee complaints are received.

- Employees are having difficulty hearing normal conversations in the work area.

- Noise reevaluation will be conducted when there are operational changes or any of the above occurs in areas previously identified as no significant noise level present.

- All continuous, intermittent and impulse sound levels from 80 to 130 dB will be integrated into the noise measurements.

- Where noise dosimetry indicates that personnel are exposed to an ECL of 85 dBA or grater greater, the Safety Office will recommend that the affected employees be placed on the Hearing

Conservation Program. Consideration will be given to personnel who are subject to spending time in hazardous noise areas or whose exposure may exceed 85 dBA on an irregular basis.

- Audiology will administer the following audiometric testing:

- All employees exposed to an 8 hour TWA of 85 dBA or more shall be identified for inclusion within the audiometric testing program. This program shall result in the acquisition and assessment of baseline and annual audiograms. Full audiometric evaluation will also be performed prior to the employee termination date. The Service Administrator is responsible for scheduling the audiograms. The audiometric testing process is outlined in flowchart (Attachment C).

- In the event that pulse-tone audiometers are used, they shall have a tone on-time of at least 200 milliseconds. Rooms used for audiometric testing shall not have a background sound pressure level exceeding Table D-1 of appendix D.

- The audiometric test chamber shall be evaluated periodically in accordance with current American National Standard Institute (ANSI) standards.

- Rooms used for audiometric testing shall not have a background sound pressure level exceeding Table D-1 of 29 CFR 1910.95 Appendix D. and shall maintain accurate records of the measurements.

- The audiometer shall be checked before each day’s use and the audiometer’s output is listened-to to ensure the output is free from distorted or unwanted sounds.

  - The Audiometer calibration is checked acoustically biannually in accordance to VA standards for Veteran Compensation and Pension testing.

  - Baseline Audiogram:

  - A baseline audiogram shall be performed within 6 months of an employee’s first exposure to an 8 hour TWA of 85 dBA. A copy of the baseline results shall be forwarded to Employee Health Nurse or designee. Employees should not be exposed to workplace noise for more than 14 hours before the baseline test or wear hearing protectors during this time period.

  - A licensed or certified audiologist, an otolaryngologist, or a technician who is certified by the Council of Accreditation in Occupational Hearing Conservation shall perform audiometric testing.

- Annual Audiograms:

  - An annual audiogram shall be performed on all employees included in the hearing conservation program. Subsequent annual audiograms shall be performed within one year of the previous audiogram.

  - At risk employees shall be monitored annually unless more frequent testing is needed or if there is suspect of medical pathology of the ear is caused by the hearing protection devices available.

  - The annual audiogram shall be compared with the baseline audiogram in order to determine if the employee has developed a STS which is an average shift in either ear of 10 dB or more at 2,000; 3,000; and 4,000 hertz.

  - An annual audiogram may be substituted for the baseline audiogram when, in the judgment of the audiologist, otolaryngologist, or physician who is evaluating the audiogram notices the STS is persistent or if the annual audiogram indicates significant improvement over the baseline audiogram.

  - Follow-up audiograms are performed if the annual audiogram shows that a STS has occurred. The follow-up is performed within 30 days of the annual audiogram.

  - In determining whether a STS has occurred, allowances may be made for the contribution of presbycusis  to the change in hearing level by correcting the annual audiogram according to the procedure described in 29 CFR 1910.95 Appendix F.

  - The person performing the audiometric testing may confer with the Occupational Health Physician on threshold shifts, work relatedness, and potential need for referral for other medical assessments.

- Termination of Employment Audiograms:

  - Full audiometric evaluation will be performed prior to the employee leaving station date. (b)Evaluation will include tympanometry, pure tone air/bone conduction testing and speech

testing.

(c)  The Service Administrator is responsible for scheduling the exit exam.

- Personal Protection Equipment:

- The Medical Center shall provide personal hearing protective equipment to all employees exposed to 8 hour TWA of 85 dBA or greater at no cost to the employee. Hearing protectors shall be replaced as necessary.

- Employees who are exposed for durations in excess of the OSHA permissible exposures listed in Table 1 shall be required to wear hearing protective equipment.

  - All employees included in the program shall be trained by the Audiology Service, Safety Office and/or Supervisor prior to their assignment in high noise areas

  - The use, fitting, care and maintenance of hearing protective devices to be used by the employee.

- The VA Medical Center uses the hearing protector attenuation set by the National Institute of Occupational Safety and Health (NIOSH) method No. 3: “Using a Sound Level Meter set on the C-Scale” as described in 29 CFR 1910.95 Appendix B.

- Safety Office shall recommend hearing protective devices that reduce employee exposure to below the permissible limits in Table 1 (OSHA standard 29 CFR 1910.59).

TABLE 1 – PERMISSIBLE NOISE EXPOSURES

Sound level

| Exposure Duration per Day (hours) | dBA Slow Response |
| --- | --- |
| 8 | 90 |
| 6 | 92 |
| 4 | 95 |
| 3 | 97 |
| 2 | 100 |
| 1 ½ | 102 |
| 1 | 105 |
| ½ | 110 |
| ¼ or less | 115 |

- Designated hazardous noise areas will be posted with warning that informs personnel that the area is a hazardous noise and hearing protection is required.

- Equipment that produces hazardous noise (levels at or above 85 dBA) will be labeled with a warning informing personnel that hearing protection is required during its operation. See attached list of previously identified hazardous noise areas and equipment (Attachment A).

- If subsequent audiometric testing of an employee whose exposure to noise is less than an 8-hour TWA of 90 dBA indicates that a STS is not present, the employer:

  - Shall inform the employee of the new audiometric interpretation.

  - May discontinue the required use of hearing protectors for that employee.

- For employees who have experienced a STS, and/or if the noise exposure increases to the extent that the hearing protectors provided no longer provide adequate attenuation, the employer shall provide more effective hearing protectors where necessary.

  - Properly fitted earplugs or muffs reduce noise 15 to 30 dB. The better earplugs and muffs are approximately equal in sound reduction, although earplugs are better for low frequencies noise and earmuffs for high frequency noise.

  - Simultaneous use of earplugs and muffs usually add 10 to 15 dB more protection than either used alone. Combined use should be considered when noise exceeds 105 dB.

g.  Training:

- All employees included in the program shall be trained by the Audiology Service, Safety Office and/or Supervisor prior to their assignment in high noise areas in the following

  - The use, fitting, care, and maintenance of hearing protective devices to be used by the employee.

  - The affects of noise on hearing.

  - The purpose of audiometric testing and an explanation of the testing procedures by the Audiologist/Speech Pathology Department.

- A power point presentation and the Learning Management System (LMS)-item number: NFED 27236, located on the public drive / PM&R / Audiology / Hearing Conservation, is available for employee training.

- A copy of the OSHA Standard, 29 CFR 1910.95 “Occupational Noise Exposure” shall be posted in the workplace and made available to employees, as well as, any informational material pertaining to the standard that are supplied to the employer by the Assistant Secretary.

  - Upon request, all material related to the employer’s training and education program pertaining to this standard is available through the Audiology and Engineering Departments.

- REFERENCE: OSHA Standard, 29 CFR 1910.95 “Occupational Noise Exposure”

- FOLLOW-UP RESPONSIBILITIES: Safety Officer, Engineering Service.

- RECERTIFICATION: This memorandum rescinds VAMC Memorandum No. 138-05, dated December 31, 2009 and will be due for review December 31, 2012 in accordance with procedures outlined in VAMC Memorandum No. 001-01.

Michael M. Young, P.E., CHFM Chief, Engineering Service

ATTACHMENT A

VA MEDICAL CENTER	SOP-ENG-138-02

LEXINGTON, KENTUCKY	October 1, 2012

KNOWN HAZARDOUS NOISE EQUIPMENT/AREAS

Leestown Division:	Decibels (dBA)

Carpenter Shop

DeWalt Radial Arm Saw S#: 228869200817-49	99.0 (Regular Wood) DeWalt Radial Arm Saw S#: 228869200817-49	110 (Treated Wood) DeWalt Radial Arm Saw S#: 228869200817-49	104.0 (Metal)

Powermatic Table Saw EE#: 324	94.4 (Regular Wood)

Delta Shaper EE#: 12589	100.0 (Wood)

Band Saw EE#: 318	93.5 (Treated Wood)

Band Saw EE#: 318	101.2 (Metal)

Bench Grinder EE#: 6213	84.1 (Idol)

Bench Grinder EE#: 6213	107.7 (Metal)

Delta Belt Sander S#: 013727	86 Wood

Delta Belt Sander S#: 013727	86.0 (Metal) Porter Cable Portable Scroll Saw (.4HP) S#: 27811	94.6 (Idol) Porter Cable Portable Scroll Saw (.4HP) S#: 27811	95.7 (Wood) Milwaukee Rotary Hammer, 3/8” S#: 672D503300693	104.0 (Brick) Stihl Gas Powered Concrete Saw S#: 133309191	100.0 (Idol) Stihl Gas Powered Concrete Saw S#: 133309191	100.2 (Concrete) Air Hose psi = 100		91.5

Tile Cutter Saw, Model 660 S#: 510174	72.1 (Idol)

Tile Cutter Saw, Model 660 S#: 510174	100.0 (Tile)

Bosch Router Model #: 0601905439	92.0 (Idol)

Bosch Router Model #: 0601905439	98.4 (Wood)

Super SawZall S#: Unreadable	95.6 (Idol)

Super SawZall S#: Unreadable	97.9 (Wood)

Super SawZall S#: Unreadable	106.8 (Metal)

Hitachi Circular Saw 7 ¼” S#: 73154	96.3 (Idol)

Hitachi Circular Saw 7 ¼” S#: 73154	102.8 (Wood)

Bosch Rotozip (18v) S#: 88200	95.5 (Idol)

Bosch Rotozip (18v) S#: 88200	95.3 (Metal)

DeWalt Nail Gun (Cordless) S#: 023061	95.2

Air Hose psi=145	103.6

Bosch Jack Hammer Model #: 061130403A	97.0

Plumbing/Machine Shop

| Pipe Machine EE#209 | 94.4 |
| --- | --- |
| Air Hose psi = 160 | 112.6 |
| Kalamazoo Band Saw S#: 394144 Bench Grinder S#: 111985 | 94.3 (Metal) 102.2 |
| Sewer House (bldg. 38) EE: 2001 | 89.3 |

Spartan 2001 Sewer Machine (bldg. 67) S#: 44226000	78.7

Electric Shop

Bench Grinder (1/2HP) S#: Unknown	84.0 Metals

Bench Grinder (1/2 HP) S#: Unknown	74.8 (Idol)

Air Hose 9-45 Code #: 7120	104.1 (Steady stream of air)

DeWalt Cordless Drill ID#: 867569	80.6 (idol) Grounds & Miscellaneous Services

Troybuilt Push Mower 22” (6.5 HP) S#: 1125K11698	80.9

Edger (3.5HP) (Non Functional)	87

John Deere Backhoe EE#: 20213	90.4 (Full Throttle)

International Tractor S#: 20200010010847	90.7

Sthil Br550 Leaf Blower EE#: 45234	101.9

Sthil Weed eater 250 EE#: 58088	78.8 (Idol)

Sthil Weed eater 250 EE#: 58088	98.1 (Full Throttle)

Stihl Chain Saw EE#: 50807	74.3 (Idol)

Stihl Chain Saw EE#: 50807	106.8 (Full Throttle)

Poulan Pro Hedger S#: 07180D401285	75.5 (Idol)

Poulan Pro Hedger S#: 07180D401285	96.2 (Full Throttle)

John Deere Tractor S#: 26210A	101.4

Prosthetics

| Bench Grinder EE#: 34335 | 81.1 |
| --- | --- |
| Dust Collect Portable EE#: 34335 | 81.1 |
| S400 Grinder EE#: 43713 | 86.8 |
| Dayton Band Saw EE#:34365 | 86.2 |
| Air Hose psi = 84 | 82.1 |
| Dotco Air Saw Model #: 12L2065-90 | 85.5 |

Dental

| Ultra Sonic | 67.8 |
| --- | --- |
| High Speed Hand Piece | 67.0 |
| Slow Speed Hand Piece | 75.6 |
| High Volume Evacuator | 73.0 |
| High Speed Hand Piece & High Volume Evacuator | 74.3 |
| Slow Speed Hand Piece & High Volume Evacuator | 73.4 |
| High Volume Evacuator & Air Water Syringe | 76.6 |
| Air Water Syringe | 65.0 |
| Saliva Injector | 70.7 |
| Promix DENTSPLY Agitator | 69.5 |
| Air Hose psi = ~30 (no gage) | 93.0 |
| Red Wing Bench Lathe & Ventilation Unit EE#: 51691 | 79.3 |
| Whip Mix Grinder EE#: 51569 | 83.9 |
| Yames and Bird Vibrator | 76.8 |
| Lab Combo Fume Hood EE#: 835 | 80.2 |
| Pneumatic Drill EE#: 50666 | 84.7 (stone) |
| Pneumatic Drill EE#: 50666 | 79.1 (dry) |

Nutritional Services

| Hobart Dishwasher - Pots & Pans S#: 73206652 | 78.3 |
| --- | --- |
| Hobart Dishwasher – Dishes EE#: 36303 | 83.3 |
| Hobart Dishwasher – Dishes (Dish Exit) EE#: 36303 | 78.8 |
| Hobart Dishwasher – Dishes (food scraper) EE#: 36303 | 83.5 |
| Grinder/Pulper | 85.6 |
| Train Air Conditioner S#: S86E09332 | 70.8 |
| Hobart Stand Mixer EE#: 32484 | 65.4 |
| Commercial Blender EE#: 6246 | 92.0 |

Emergency Generator Rooms

Building 69	(For Bldg 1, 2, 3, 4)113

Building 71	(For Bldg 25)	104

Building 73	(For Bldg 16, 29) 105

Building 72	(For Bldg 27, 28) 105

Building 39	(For Chiller Plant & Boiler)

111

Building 00	(For Bldg 17)	107

Chiller and Boiler Plant

| Chiller Plant: |  |
| --- | --- |
| 4’ vicinity of Unit #2 while 40% capacity | 95 |
| On floor while Unit #2 40% capacity | 87 |

Bldg. 39, Boiler Plant Pump Room:

1 pump on, at valves	93

Fire Pump Bldg. 45 (Not in Use)	109

Cooper Division:	Decibels (dBA)

Carpenter Shop/Grounds & Miscellaneous Service

Delta Band Saw	85.3 (plastic)

Table Saw	98.0 (wood)

Portable Router – Porter Cable	85.0 (dry)

Porter Cable Portable Router	89.0 (vermica)

Skill Saw	105

Tile Cutter	106.0 (wet cut)

Tile Cutter	75.4 (dry)

Delta Sander/Grinder Combo EE#: 3182	91.7 (steel)

DeWalt Miter Saw	102.1 (wood)

DeWalt Portable Chop Saw EE#: 61805 [Generated sparks/odor)	106.9 (Metal) DeWalt Orbit Sander	77.5 (dry

DeWalt Orbit Sander	80.5 (wood)

Air Compressor	103.2

Hitachi Reciprocating Saw	96.7 (Metal)

Garage

Murray 22” Push Lawn Mower	99.3

Sthil F5 250R Weed eater	102.2

Sthil Farm Boss Chain Saw	80.8 (idol)

Sthil Farm Boss Chain Saw	110.3 (running)

Sthil Leaf Blower	104.2

Plumbing/Machine Shop

Powermatic Bench Grinder (1/2HP), RM B-12 EE#: 3180	83.3 (~15 min. runtime) Portable Grinder, Disk (Unknown)		104

Pipe Machine (Unknown)	90

Bridgeport Press EE#: 3207	74.2

Powermatic Model#: 315S269	65.2

Oxygen Delivery	77.0 (ambient near O2 tank)

Oxygen Delivery	95.7 (during deliver)

Room Numbers

Chiller Room	98.1B-39, Mechanical Room

84

Chiller Plant, N Door Entrance	94 1 unit on

Chiller Plant, Floor (Average)	94 1 unit on

Chiller Plant, 4’ Unit Vicinity	98 1 unit on

Chiller Plant, Glass Office	67 1 unit on Boiler Plant (B702)

Office	65.2

Boiler #1	85.8

Air Handling Units	73.0

Air Compressor for Comps EE#: 2877	82.6

Air Intakes	81.2 – 108.0

Boiler Plant, 4’ of unit vicinity	88

| Generator (only on running) | 88.6Emergency Generator, |
| --- | --- |
| (Room B038) |  |
| Generator Room (before start of generator) | 65.4 |
| Generator Room (Outer Room) | 65.0 |
| Generator Room (During Operation) | 112.1 |
| Generator Room (Outer Room during Operation) | 74.4 |

Ware House

Ambient Noise	79.0

Box Crusher	91.2 (~ 5 min a few times a

day)

Dental

Quantrex L&T Ultra Sonic EE: # 53016	84.8 (at source)

Quantrex L&T Ultra Sonic EE: # 53016	65.1 (at work station)

High Speed Hand Piece Pneumatic Drill	65.8

High Volume Evacuator (HVE) Suction Hose	71.3

HVE & High Speed Hand Piece Pneumatic Drill	74.8

Slow Speed Hand Piece	66.7

HVE & Slow Speed Hand Piece	77.8

Air Water Syringe	71.2

HEV & Air Water Syringe	74.0

Saliva                                                                                  Ejector 	66.8

ProMix DENTSPLY Algamator EE#: 33755                               77.1 (at Source)

ProMix DENTSPLY Algamator EE#: 33755                               72.2 (at work station)

Laboratory

Centrifuge EE#: 5891	80.4

Hermle Centrifuge EE#: 35724                                                     86.6 (Box open)

Hermle Centrifuge EE#: 35724                                         83.4 (Box Closed) Thomas Vertical Vibrator Plate Mixer EE#: 1463                87.0 (walks away) Germ Free Hood EE#: 49864	76.4

Research

Steris Basil 6700 Tunnel Cage Washer EE#: 61613	78.9

Bio Bubble EE#: 54911	78.9

Cage and Rack Washer Model #: RW4750	78.9

Nutrition and Food Services

| Dish Machine | 79.0 |
| --- | --- |
| Dish Machine | 77.7 |
| Pulp Machine & Dish Washer | 80.8 |
| Pulp Machine & Dish Washer | 80.0 |
| Pulp Machine | 84.5 |

VA MEDICAL CENTER	SOP-ENG-138-02

LEXINGTON, KENTUCKY		OCTOBER 1, 2012 ATTACHMENT  B

TRAINING OUTLINE

Training Outline: Hazardous Noise/Hearing Conservation Program Covered by: OSHA, 29 CFR 1910.95

- GENERAL PROGRAM ELEMENTS

  - Noise measurement.

- Sound level meter measures instantaneous noise levels in decibels (dB).

- Noise dosimetry measures the equivalent continuous sound level (ECL) over the entire work shift.

  - All personnel whose 8 hour time-weighted-average (TWA) exposure to noise exceeds 85 dBA must be placed on the Hearing Conservation Program.

  - Personnel must wear hearing protection while in areas where noise levels exceed 85 dBA (included lawn mowers, power saws, boiler plant, etc.).

- EFFECTS OF NOISE ON HEARING:

  - High noise levels may cause either temporary or permanent hearing loss.

  - Damage is caused by sound pressure on the hair cells in the inner ear (cochlea). Hearing loss normally begins at higher frequencies and spreads to lower frequencies.

  - Continued unprotected exposure makes hair cells less able to recover, resulting in permanent hearing loss.

  - Examples of various environmental noises and their equivalent decibel levels (see Attachment A).

- HEARING PROTECTION:

  - Ear plugs.

fit.

- Sized ear plugs (flange type) – come in various sizes, protection depends on quality of

  - Attenuation ranges from 15 – 35 dB.

- Malleable ear plugs (EAR type) – plug is compressed by hand and inserted into ear canal where it expands to provide fit for various sizes.

  - Attenuation ranges from 29 – 41 dB

  - Earmuffs.

    - Attenuation ranges from 20 – 50 dB

  - Combination of ear plugs and earmuffs.

    - Normally required for high noise level (above 115 dB) for extended periods.

  - Demonstration of appropriate insertion if earplugs and position of earmuffs.

- AUDIOMETRIC TESTING:

  - Summary of hearing/screening procedures (e.g., baseline, annual, termination and follow- up audiometric testing).

ATTACHMENT C

### Attachment G:  Biomedical Engineering Annual Checklist

Monthly PM List: 35-BIOMEDICAL CDD Shop for 2/11  Printed: 02/16/2011  Page 1

Order: LOCATION  DIV(All) BLDG(All) WING(All) ROOM(All)

Includes OUT OF SERVICE Equip.  Responsible Tech: HARRIS, RALPH

Entry #    Equipment Category             Model            Serial Number

[ROOM-BLDG-DIV (Wing)]   Manufacturer Equipment Name            Local ID

Status   PM #       Manufacturer                   Service

Work Order Number

-------------------------------------------------------------------------------

16890      BALANCE-ELECTRICAL             CT600S           16780

B102-01-CDD (1B)       BALANCE, ANALYTICAL                      35R-16890

IN USE   6670-5813  OHAUS SCALE                    PATHOLOGY-CDD

Proc:                               Crit: 6   Freq: A  Level: 1,2

PM-B1102M-192    Initials:_____ Date:_______ Hours:_____(.5) Cost:______

PM Status (circle): P C D0 D1 D2 D3  Condition (GOOD): LN G P  Y2K: FC

27723      BALANCE-ELECTRICAL             GT4000           5979

B102-01-CDD (1B)       BALANCE, ANALYTICAL                      35-27723

IN USE              OHAUS SCALE                    PATHOLOGY-CDD

Proc:                               Crit: 8   Freq: A  Level: 1,2   JCAHO: YES

PM-B1102M-193    Initials:_____ Date:_______ Hours:_____(.5) Cost:______

PM Status (circle): P C D0 D1 D2 D3  Condition (GOOD): LN G P  Y2K: FC

36302      BALANCE-ELECTRICAL             OHAUS BALANCE    1125140169

B102-01-CDD (1B)       BALANCE ANALYTICAL OHAUS                 35-36302

IN USE   6670-36302 OHAUS SCALE                    PATHOLOGY-CDD

Proc:                               Crit: 8   Freq: A  Level: 1,2   JCAHO: YES

PM-B1102M-194    Initials:_____ Date:_______ Hours:_____(.5) Cost:______

PM Status (circle): P C D0 D1 D2 D3  Condition (GOOD): LN G P  Y2K:

55883      COAGULATION PROFILER           STAGO B4600      BU29017583

B102-01-CDD (1B)       TIMER, BLOOD AND PLASMA COAGULATION'     35-55883

IN USE              DIAGNOSTICA STAGO, INC         PATHOLOGY-CDD

Proc:                               Crit: 6   Freq: A  Level: 1,2,3 JCAHO: YES

PM-B1102M-195    Initials:_____ Date:_______ Hours:_____(1) Cost:______

PM Status (circle): P C D0 D1 D2 D3  Condition: LN G P  Y2K:

Monthly PM List: 35-BIOMEDICAL CDD Shop for 2/11  Printed: 02/16/2011  Page 22

Order: LOCATION  DIV(All) BLDG(All) WING(All) ROOM(All)

Includes OUT OF SERVICE Equip.  Responsible Tech: HARRIS, RALPH

Entry #    Equipment Category             Model            Serial Number

[ROOM-BLDG-DIV (Wing)]   Manufacturer Equipment Name            Local ID

Status   PM #       Manufacturer                   Service

Work Order Number

-------------------------------------------------------------------------------

55884      COAGULATION PROFILER           58612            BU29017584

B102-01-CDD (1B)       TIMER, BLOOD AND PLASMA COAGULATION      35-55884

IN USE              DIAGNOSTICA STAGO, INC         PATHOLOGY-CDD

Proc:                               Crit: 6   Freq: A  Level: 1,2,3 JCAHO: YES

PM-B1102M-196    Initials:_____ Date:_______ Hours:_____(1) Cost:______

PM Status (circle): P C D0 D1 D2 D3  Condition: LN G P  Y2K:

### Attachment H:  BIOMEDICAL NEW EQUIPMENT CHECK IN SHEET

### Attachment I:  SERVICE SAFETY TEAM INSPECTION PROGRAM CHECKLIST

Service Safety Team Inspection Program Checklist

| - This checklist supports Other areas and activities that do not fit in PCS, Clinical, Administrative or Industrial. It mirrors most, but not all, items on the EOC checklist. Please identify your Service and the area(s) you have inspected. Also, list all the persons involved in the inspection. |  |
| --- | --- |
| - Service: Pathology & Laboratory Medicine - ALL Areas Inspected: Laboratory______________________________ | Team Member(s) Present: Vicki Richardson Date: Supervisor Present:________________________ |

| Items to Review/Inspect | Std. | Meets | Not Meet | N/A | Finding and Corrective Action(s): | Correction Date |
| --- | --- | --- | --- | --- | --- | --- |
| Violence of Prevention |  |  |  |  |  |  |
| Are there any easily accessible objects which might be used as a weapon? |  |  |  |  |  |  |
| Does staff know the early recognition of escalating behavior or recognition of warning signs? |  |  |  |  |  |  |
| Are there emergency numbers posted by phones? |  |  |  |  | Contact the ADPAC and put in a ticket for faded or if replacement stickers are needed. |  |
| Did you know the VA has zero tolerance for work place violence? |  |  |  |  |  |  |

| Infection Control |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Staff use proper hand washing procedures | SOARS |  |  |  |  |  |
| Standard precautions used for all patients | IC.1.10 |  |  |  |  |  |
| Eating, drinking, smoking, applying cosmetics or lip balm, and handling contact lenses are prohibited in work areas where there is a reasonable likelihood of occupational exposure. | IC.1.10 |  |  |  |  |  |
| PPE available and used correctly | IC.1.10 |  |  |  |  |  |
| Sharps containers working properly, not full, secured | IC.1.10 |  |  |  |  |  |
| Management of the Environment of Care |  |  |  |  |  |  |
| No chemicals or housekeeping carts are left unattended and unsecured in work areas |  |  |  |  |  |  |
| The environment is maintained in a clean, clutter free manner | EC.8.10 |  |  |  |  |  |

| Under sink areas have only household type cleaning supplies (i.e. dish soap) or cleaning materials specifically approved for Service activities. | EC.8.10 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Signage correct, framed and legible for the area | IC 4.10 |  |  |  |  |  |
| Floors clean, including edges, and all walking surface |  |  |  |  |  |  |
| Staff refrigerators clean, including gaskets? Defrosted? |  |  |  |  |  |  |
| Hallways and corridors clear of obstacles? | SOARS |  |  |  |  |  |
| Appropriate doors are locked, (Housekeeping closet, linen closet, electrical closets, electrical panels, etc.) |  |  |  |  |  |  |
| Staff are able to articulate their role in a security incident | EC.1.10 |  |  |  |  |  |
| Ceiling tiles in good repair, no sign of leakage | EC.2.10 |  |  |  | Must have corrective action: Room and WO #. |  |
| No evidence of mold or moisture (ceiling tiles, bathroom tiles, wallpaper, windows, etc) | EC.8.10 |  |  |  |  |  |

| All lights are operable, with no burned out bulbs | SOARS |  |  |  | Must have corrective action: Room and WO #. |  |
| --- | --- | --- | --- | --- | --- | --- |
| Utility rooms / closets clean and in good repair |  |  |  |  | Must have corrective action: Room and WO #. |  |
| Use of barriers in construction areas to prevent contamination of clean areas |  |  |  |  |  |  |
| Signs clear and easy to read? Exits clearly marked and lit? User friendly? | NFPA |  |  |  | Must have corrective action: Room and WO #. |  |
| Electrical Safety |  |  |  |  |  |  |
| All personal electrical appliances have been safety checked? | VAMC |  |  |  | Must have corrective action: Room and WO #. |  |
| No open element heating appliances? | VHA |  |  |  |  |  |
| Power strip only small appliances or equipment. No microwave or refrigerator on power strips. | NEC |  |  |  |  |  |
| Other Electrical: No extension cords No expansion outlet plugs. Acceptable cords not damaged; strained. | OSHA |  |  |  |  |  |
| Microwave has user responsibility label (VAMC) | NFPA |  |  |  |  |  |
| Safety |  |  |  |  |  |  |
| Staff can locate Service Emergency Contingency Plan and discuss use. | OSHA |  |  |  |  |  |
| Staff are current for respiratory fit testing | OSHA |  |  |  |  |  |
| Staff know where to find the Fire & Life Safety Mgt Plan | JC |  |  |  |  |  |
| Staff know where to find the Safety Mgt Plan | JC |  |  |  |  |  |
| Director’s Safety Policy is posted or had documented review with staff |  |  |  |  |  |  |
| No sprinkler heads are observed to be obstructed-18" clearance | NFPA |  |  |  |  |  |
| Staff are able to articulate their role in a fire drill- including RACE and PASS | OSHA EC 4.10 |  |  |  |  |  |
| Staff can locate nearest exit, fire extinguisher and pull station | OSHA EC.5.10 |  |  |  |  |  |
| All exits are free and unobstructed | OSHA EC.5.30 |  |  |  |  |  |

| Staff can verbalize where to locate MSDS information and Hazardous Communication Program is current in MSDS binder? | OSHA EC.5.10 |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Paper and other combustibles are not stored under desks, work stations. | EC.5.50 |  |  |  |  |  |
| Chairs, work stations, and supporting computer equipment are adjusted for user/multiple users. | SOARS |  |  |  |  |  |
| No items stored on tops of cabinets, lockers and ledges. | OSHA |  |  |  |  |  |
| Are stepladders and step stools in safe condition? | OSHA |  |  |  |  |  |
| Are supplies and equipment stored properly – heavy items waist high and light items high or low; stable; easy to identify. | OSHA |  |  |  |  |  |
| Floors mopped leaving dry path | OSHA |  |  |  |  |  |
| Wet walking / working surfaces have appropriate signs placed immediately warning of wet surface. | OSHA |  |  |  |  |  |
| Wet signs removed after floor dry. |  |  |  |  |  |  |
| Cleaning supplies: a. labeled b. securely fastened or stored in appropriate cabinet/closet c. no food or food containers (cups, lunch bags, etc) stored with food | OSHA |  |  |  |  |  |
| Personal Protective Equipment a. stored clean, ready to use b. close to hazards c. staff trained in use d. staff use when potential for exposure occurs e. Safety goggles, gloves, face shield, mask, hearing protection (circle what was inspected) | OSHA |  |  |  |  |  |
| Emergency eyewash station logs current (weekly test and monthly inspection) | ANSI |  |  |  |  |  |
| Safety topic discussed at this month’s staff/team meeting: _________________ |  |  |  |  | Indentify topics your service would like information? |  |

| EMS (CD/LD) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| All cleaning tools stored with sharps and points down or with safety guard. |  |  |  |  |  |  |
| Housekeeping closets neat, orderly, and no trash. (Should conduct physical inspection of 1/6th of closets each month so all are inspected by end of six months.) |  |  |  |  |  |  |
| Housekeeping carts have advance warning signs for strip/wax cleaning activities. | (List closets inspected in this block.) |  |  |  |  |  |
| N&FS ONLY |  |  |  |  |  |  |
| N&FS- only: Pulper disconnected when not in use. |  |  |  |  |  |  |
| Food waste and trash in covered container. |  |  |  |  |  |  |
| Hot pads available & Used. |  |  |  |  |  |  |
| CAUTION sign placed on combo-oven during cleaning process. |  |  |  |  |  |  |
| Knives stored in racks. |  |  |  |  |  |  |
| Knives washed according to N&FS Policy. |  |  |  |  |  |  |
| Are tuggers being used to move food carts? |  |  |  |  |  |  |

### Attachment J:  HAND HYGIENE FOR HEALTH CARE WORKERS Memorandum No. 111-09

HAND HYGIENE FOR HEALTH CARE WORKERS

PURPOSE:  To provide guidance for establishing the basic requirements for hand hygiene.

POLICY:  All staff will practice good hand hygiene and staff on patient care units will monitor staff practice by direct observations to improve hand hygiene techniques.  Infection Control staff will make additional observations (during tour of duty hours).

RESPONSIBILITY: Every health care worker (HCW) in the medical center is responsible for proper hand hygiene.

Director.  Holding all HCWs accountable to perform proper hand hygiene.

Chiefs of services, supervisors, managers, and HCWs. Acting as role models, instructors, and monitors to ensure proper hand hygiene.  All HCWs are to be held accountable for proper hand hygiene and use of personal protective equipment (PPE). Chiefs of services, supervisors and managers must hold everyone accountable. Failure on the part of any HCW to properly and appropriately conduct proper hand hygiene will be dealt with in accordance with the theory of progressive discipline as defined by Human Resources.

PROCEDURES

All HCWs in direct patient contact areas, i.e., inpatient rooms, outpatient clinics, etc., as well as those who may have direct patient contact in other settings, such as radiology technicians, phlebotomists, etc., are required to:

Use an alcohol-based hand rub or antimicrobial soap and water to routinely decontaminate their hands before and after having direct contact with a patient.  A single act of using an alcohol-based hand rub or an antimicrobial soap and water after one patient and before the next patient suffices to decontaminate a HCW’s hands, if the hands are not subject to recontamination in-between patients (as in talking on the telephone, handling objects, pushing elevator buttons, turning doorknobs, etc).

Procedures for hand hygiene are as follows Note: follow manufacturer’s instructions for all products:

| Soap and Water | Alcohol based hand rub | Surgical Hand Antisepsis |
| --- | --- | --- |
| 1. Wet hands with water. (NOTE: Hot water increases dermatitis risk) | 1. Apply recommended amount of product to the palm. | 1. Remove rings, watches and bracelets before beginning the surgical hand scrub. Apply to dry hands. |
| 2. Apply soap. | 2. Rub hands together to cover all surfaces of hands and fingers. | 2. Remove debris from underneath fingernails using a nail cleaner under running water |
| 3. Rub all surfaces of hand, cuticles and between fingers for at least 15 seconds. | 3. Rub until hands are dry. | 3. Surgical hand antisepsis using either an antimicrobial soap or an alcohol- based hand rub with persistent activity is recommended before |

|  |  | donning sterile gloves when performing surgical procedures |
| --- | --- | --- |
| 4. Wash around each wrist and wash under rings. | 4. Do not touch electrical sockets or switches or engage in activities that may generate static electricity until hands are dry. | 4. When performing surgical hand antisepsis using an antimicrobial soap, scrub hands and forearms for the length of time recommended by the manufacturer, usually 2–6 minutes. |
| 5. Rinse hands and under rings. Keep hands and clothing away from sink. | 5. Do not use alcohol rubs with patients that have C. diff. | 5. When using an alcohol-based surgical hand-scrub product with persistent activity, follow the manufacturer’s instructions. Before applying the alcohol solution, pre- wash hands and forearms with a non- antimicrobial soap and dry hands and forearms completely. |
| 6. Dry hands by patting dry with a disposable towel and use towel to turn the water off. |  | 6. After application of the alcohol- based product as recommended, allow hands and forearms to dry thoroughly before donning sterile gloves |

If hands are not visibly soiled, an alcohol-based hand rub may be used for routinely decontaminating hands.

When hands are visibly dirty or contaminated with proteinaceous material or are visibly soiled with blood or other body fluids, wash hands with antimicrobial soap and water.  Avoid using hot water as repeated exposures to hot water can increase the risk of dermatitis.

New products for handwashing and hand care must be reviewed and approved by the Skin Care Committee and Infection Control Committee prior to submission for final approval to the Product and Technology Standardization Team/Commodity Standardization Program.

Alcohol rubs are not to be used when Bacillus anthracis (Anthrax) or Clostridium difficile (C. diff) infections are suspected or proven.  Hands must be washed with soap and water to remove these bacteria and the spores that form from the bacteria when in a dormant state in the environment.  Alcohol based hand rubs are not effective against spores. Hands are always to be dried using a method that will not re- contaminate them, such as using individual or motion-activated paper towel dispensers.

Use an alcohol-based hand rub or antimicrobial soap and water.

Before inserting or handling any invasive device for patient care, whether or not gloves are used. Note:  Non-latex gloves are available for personnel with latex sensitivity.

Before and after donning or removing sterile or non-sterile gloves.

After contact with a patient’s intact skin (e.g., when taking a pulse or blood pressure or lifting a patient).

After contact with body fluids or excretions, mucous membranes, nonintact skin, and wound dressings if hands are not visibly soiled.

If moving from a contaminated body site to another body site during care of the same patient.

Before handling medication.

After contact with inanimate surfaces and objects (including medical equipment) in the immediate vicinity of the patient.

VHA food service workers are to practice appropriate hand hygiene as they perform their duties. They are to use soap and water rather than alcohol-based hand rubs in food preparation settings.  Those workers delivering food to inpatients are not required to routinely decontaminate their hands before and after entering each patient room. However, if during the course of delivering the food, the food service worker touches the patient, the bed linens, or objects in the room, the food service worker must decontaminate their hands before exiting the room.  Hands must be decontaminated when entering and leaving a patient care ward area (see VHA Handbook 1109.04, for additional information).

The facility Environment Management Service staff refers to guidance on hand hygiene and related practices in the Environmental Services Procedure Guide provided by the Environmental Programs Service within the VHA Office of the Deputy Under Secretary for Operations and Management.

Wear gloves when contact with blood or other potentially infectious materials, mucous membranes, and nonintact skin is anticipated.

Remove gloves after caring for a patient.

Do not wear the same pair of gloves for the care of more than one patient, and do not wash gloves between uses with different patients.

If gloves become visibly soiled, or if performing patient care on a contaminated site, remove or change gloves and perform hand hygiene before moving to another body site on the same patient, a device, or the environment.

During wound care remove gloves and perform hand hygiene after removing a dirty dressing. Don clean or sterile gloves to apply clean/sterile dressing. Perform hand hygiene after removing gloves.

Use an alcohol -based hand rub or antimicrobial soap and water to decontaminate hands after removing gloves.

Decontaminate hands after contact with inanimate objects (including medical equipment) in the immediate vicinity of the patient.

Any hand lotion brought in for individual use by those employees required to use latex gloves must not be used if petroleum, Vaseline, mineral oil, or other oil emollients are listed on the label. The employee must check the label and not use the lotion at work if those ingredients are listed.

All facility staff are to wash hands with a non-antimicrobial or antimicrobial soap and water whenever hands are visibly soiled or contaminated with body fluids, before eating, and after using the toilet.

For water loss issues.

In the event of a water main break where tap water is available but not drinkable: If hands are contaminated or visibly soiled or if the patient has a spore forming bacteria such as C. diff., wash hands in tap water with soap and water, rinse, dry, then follow with use of alcohol rub. Bottled water may also be used.  If hands are not contaminated or not visibly soiled, use the alcohol-based hand rubs.

In the event of no tap water availability, bottled water must be used to wash hands with soap and water if hands are contaminated or visibly soiled or if the patient has a spore forming bacteria such as C. diff.

All HCWs, who provide direct, hands-on care to patients, must not wear artificial fingernails or extenders; this includes non-supervisory and supervisory personnel who regularly or occasionally provide direct, hands-on care to patients. Natural tip nails are to be less than ¼ inch long.

Appropriate supplies will be provided.

An alcohol-based hand rub is available at the entrance to each patient room or at the bedside, as well as other locations such as clinics, emergency rooms, Community Living Centers, etc.  Note: Alcohol-based hand rubs contain at least 60 percent ethyl alcohol and may present an abuse risk in certain patient care areas, such as inpatient psychiatric or substance use treatment units.

Antimicrobial soap must be available in all patient care areas where soap is provided (i.e., at all sinks with a soap dispenser).  Note:  Antimicrobial-impregnated wipes (i.e., towelettes) are not as effective as alcohol-based hand rubs, or washing hands with an antimicrobial soap and water for reducing bacterial counts on the hands of HCWs, they are not a substitute for using an alcohol-based hand rub or an antimicrobial soap.

Appropriate labeling; both antimicrobial and non-antimicrobial soap must clearly and unambiguously label the dispensers to ensure that all users know which dispenser is providing antimicrobial soap and which dispenser is providing non-antimicrobial soap.

Pocket-sized containers of alcohol-based hand rub will be available to all HCWs. Note:  This

does not imply a requirement for all heath care workers to carry pocket-sized alcohol hand rubs.

Hand lotions or creams to minimize irritant contact dermatitis will be readily available. Products designed for health care applications that do not reduce the effectiveness of other hand hygiene products, such as antimicrobial compounds, e.g., Chlorhexidine Gluconate (CHG) will be provided. Some lotions are specifically advertised as "CHG compliant."   Hand lotions or creams must be compatible with gloves being used in the facility. The petroleum, Vaseline, mineral oil or other oil emollients in many lotions can adversely affect the integrity of latex gloves.

Soap is not to be added to partially empty dispensers.  "Topping off" soap dispensers can lead to bacterial contamination.  This means that soap needs to be dispensed from disposable bladders or other containers that prevent old and new soap from mixing.

Care must be taken in installing and storing alcohol-based hand rubs consistent with fire safety requirements.  The Life Safety Code (April 28, 2004) permits the installation and use of alcohol-based hand rubs in patient rooms, suites, and corridors of healthcare facilities.  In all locations alcohol-based hand rub dispensers must not be located over, or adjacent to (adjacent is considered to be no closer than 12 inches from the center of the dispenser to either side), ignition sources (including electrical receptacles and switches).  Corridors must have at least 6 feet of clear width with hand rub dispensers spaced at least 4 feet apart.  Alcohol-based hand rub dispensers may not be installed in carpeted corridors unless the corridor is sprinkler protected. Dispensers may not project more than 6 inches into corridor egress width.  Supplies of alcohol-based hand rub products must be stored in cabinets or areas approved for flammable materials consistent with applicable regulations and standards. Alcohol rubs may not exceed a total volume of 32 wall dispensers in a smoke compartment. Consideration needs to be given to installing dispensers at a height that ensures that they can be used by staff, patients and visitors who are in wheelchairs.

Improving hand hygiene is an institutional priority and administrative and financial support must be provided, as appropriate. Financial support includes providing adequate supplies of alcohol hand rubs (wall mounted and pocket-sized), antimicrobial soaps, gloves (regular and sterile), and lotion.

Input will be solicited from employees regarding the feel, fragrance, and skin tolerance of products, such as soap, alcohol hand rub, and gloves.

Information or educational materials on the provisions of this directive will be provided to all healthcare workers.

HCWs’ adherence to required hand hygiene practices will be monitored.  Units are to document at least 30 observations per month on the forms provided by Infection Control. Observations by staff, patients and visitors may be used for reporting purposes. HCWs will be provided with information regarding their performance.  This includes reporting information on the observed percent compliance with hand hygiene practices.

REFERENCES:  VHA Directive 2011-007 Required Hand Hygiene Practices, 16 Feb 2011. Joint Commission Hand Hygiene Project: Best Practices from Hospitals Participating in the Joint Commission Center for Transforming Healthcare Project, November 2010. Joint Commission

Center for Transforming Healthcare, Storyboard on Hand Hygiene, 2010.   Joint Commission National Patient Safety Goals available at www.jc.org/accredited+organizations/patient+safety/npsg.htm. Gupta A, Della-Latta P, Todd B, San Gabriel P, Haas J, Wu F, Rubenstein D, Saiman L. “Outbreak of Extended-spectrum Beta-lactamase- producing Klebsiella Pneumoniae in a Neonatal Intensive Care Unit Linked to Artificial Nails,” Infection Control and Hospital Epidemiology. 2004 Mar;25 (3):210-5.  VA Under Secretary for Health December 15, 2003 memorandum on Hand hygiene practices available at vaww.ncps.med.va.gov/Hand Hygiene/index.html and http://vaww.ncps.med.va.gov/hand_hygiene/memo_from_undersec.pdf.  McNeil SA, Foster CL, Hedderwick SA, Kauffman CA. “Effect of Hand Cleansing with Antimicrobial Soap or Alcohol- based Gel on Microbial Colonization of Artificial Fingernails Worn by Health Care Workers,” Clinical Infectious Diseases. 2001 Feb 1;32 (3):367-72.  NOTE:  Electronically published January 18, 2001, at www.journals.uchicago.edu/CID/journal/issues/v32n3/000140/000140.html. Parry MF, Grant B, Yukna M, Adler-Klein D, McLeod GX, Taddonio R, Rosenstein C. “Candida  Osteomyelitis and Diskitis after Spinal Surgery: An Outbreak that Implicates Artificial Nail Use,” Clinical Infectious Diseases. 2001 Feb 1;32 (3):352-7.  NOTE:  Electronically published January 24, 2001, at www.journals.uchicago.edu/CID/journal/issues/v32n3/000094/000094.html. Centers for Disease Control and Prevention web site with Guidelines for Hand Hygiene in Health-Care Settings and related materials.  Available at www.cdc.gov/handhygiene/. Summary of CDC Hand Hygiene Recommendations Required by JC (Reorganized and Summarized for VA Use) American Hospital Association information on amendment to the Life Safety Code available at www.hospitalconnect.com/ashe/currentevent/abhi.html.  Edel E, Houston S, Kennedy V, LaRocco

M. “Impact of a 5-minute Scrub on the Microbial Flora Found on Artificial, Polished, or Natural Fingernails of Operating Room Personnel,” Nursing Research. 1998 Jan-Feb;47 (1):54-9.

FOLLOW-UP RESPONSIBILITY: Infection Control.

RECERTIFICATION:  On or before June 13, 2017.

DeWayne Hamlin

Director

### Attachment K: Ergonomic Symptoms Survey

Ergonomics Symptoms Survey

Employee name _______________________________________Date___________________

Job name or workstation ______________________________________________________

Shift ____________ Hours _________ worked/week _________ Time on THIS job ______

Have you had any pain or discomfort during the last year?   Yes   No

If you answered yes to previous question, check areas where you have experienced pain or discomfort:

 Neck    Shoulder    Elbow/forearm    Hand/wrist    Fingers    Upper back

 Lower back    Thigh/knee    Lower leg    Ankle/foot

Put a check by the word(s) that best describe your problem.

 Aching    Burning    Cramping    Loss of color    Numbness (asleep)    Pain

 Swelling    Stiffness    Tingling    Weakness

 Other _____________________________________________________________________

When did you first notice the problem?  ____________________________________________

How long does each episode last?  ________________________________________________

What do you think caused the problem? ____________________________________________

Have you had this problem in the past 7 days? _______________________________________

Have you had medical treatment for this problem? ____________________________________

If yes, where did you receive treatment? ____________________________________________

Describe the treatment _________________________________________________________

Did the treatment help?_________________________________________________________

How much time have you lost in the last year because of this problem? ______________ days

Please comment on what you think would improve your symptoms.______________________

____________________________________________________________________________
____________________________________________________________________________

____________________________________________________________________________

### Attachment L: Instructions for ASISTS Incident Reporting

FIVE EASY STEPS TO SUCCESSFUL INCIDENT REPORTING USING THE

ASISTS ACCIDENT REPORTING SOFWARE PACKAGE

Completing the "Accident/Illness Report (2162)" :

Once an employee reports a work related accident or illness to their supervisor the following steps outline how to report and record the incident.

- Investigate the incident.  Find out the 'who, what, when, where and how' details.  Discuss with employee and other affected people possible ways to correct the incident's cause to minimize future injuries.

- Activate the ASISTS program by using the  button.  It is on the Application Server menu for PC users or on the Programs menu for thin client users.  Sign on as you would a VISTA program.

- Use the Supervisor menu to print a copy of the Employee Bill of Rights and give it to the employee.

- Use the Supervisor menu to 'create incident/illness report' (2162) this initial form, often called the 'stub record' provides information for all succeeding reports.  Remember to use drop down menus and lists in all ASISTS forms by clicking on the down arrows.    This will make your life much easier.  Especially when you start completing forms by selecting a case and it looks like a blank form.  All fields marked with a * must be filled out.

- Once the stub record is complete and saved move on to 'Complete/Validate/Sign Incident Report (2162)'.  It is easiest to search for  as it will only give you the cases you need to work on anyway.  Remember to use the down arrow  to select your report.  Fill out the report's pages using the button as you finish each section.

- "Corrective Action" is the most important field in the report.  Please give some thought to possible ways to improve safety issues associated with the incident.  Record any steps taken by anyone to correct this situation or minimize risks.  Once the form is complete, just click the  button and use your VISTA e-signature to verify the document.

- Print the Accident/Illness Report (2162) and give the copy to the Laboratory Safety Officer.

- The employee does not have to file a CA1/2.

- If the employee wants to complete the CA1/2(Federal Notice of Traumatic Injury and Claim of Continuation of Pay/Compensation/Notice of Occupational Disease and Claim for Compensation), have the employee sign into ASISTS.  The employee enters data for the CA-1 or CA-2 using the Complete/Validate/Sign CA-1 option for injury and Complete/Validate/Sign CA-2 option for illness.  When the employee signs their portion of the CA-1 or CA-2, this triggers a bulletin to the supervisor, union representatives, and workers’ compensation personnel notifying them of the requirement to complete the form and file with the Department of Labor within 2-3 working days.  The CA1/2 is a separate form from the above incident investigation.

- The Supervisor signs on and completes the Supervisor section of the same CA-1/2.  Once the form is complete, just click the  button and use your VISTA e-signature to verify the document.  When the supervisor signs the CA-1 or CA-2 using the Complete/Validate/Sign CA-1 option for injury and Complete/Validate/Sign CA-2 option for illness, a bulletin is sent to the OOPS WCP mail group and also to the supervisor.

- Once the supervisor signs the case, the original case is no longer available for edit by either the employee or the supervisor.  To edit the claim, the safety official or the workers’ compensation personnel must create an amendment.

- If an employee is incapacitated and cannot electronically sign the claim, the workers’ compensation personnel may sign for the employee via the Electronically Sign for Employee option.

- The Supervisor prints the CA 1/2

- The Supervisor signs the hard copy with blue ink.

- The employee signs the hard copy with blue ink.

- Copy the signed copy and give the copy to the Laboratory Safety Officer.

- Mail the original signed hard copy to the WC (Worker’s Comp) office (001-WC-LD).

- If the employee needs to go to Employee Health or ED if after hours,send them. The employee may opt to go to their private physician.  The ED or a private physician will give the employee a form to describe duty restrictions (if any).  A copy of this form must be given to the employee’s supervisor when the employee returns to the laboratory.

Helpful folks if you need aid through the process:

Judy Stivers, Employee Health Nurse – x: 5270 p: 713

Teresa Abney, OWCP Assistant (CA-1) – x: 3852

### Attachment M: P&LMS New Employee Safety Walkthrough Checklist

New P&LMS Employee Safety Walkthrough Checklist

The following items are reviewed during fire drills with staff.  If the topic is not pertinent,  note N/A in the Date Reviewed column. All attendees must sign the reverse side of the form.

| TOPIC | DATE REVIEWED |
| --- | --- |
| 1. Walk your primary evacuation route. |  |
| 2. Walk your secondary evacuation route. |  |
| 3. Show the location of extinguisher(s) |  |
| 4. Show the closest 2 pull stations. |  |
| 5. Show the location of Emergency Operation Plan (EOP) - red trifold. |  |
| 6. Show the location of the MSDS book. |  |
| 7. Show the location of any evacuation equipment (evacusleds; evacuchairs; etc) | N/A |
| 8. Show the location spill kits (if used). |  |
| 9. Show any special hazardous waste containers and review what goes in each (if any) |  |
| 10. Review the drill requirements and that the procedures are actually performed. |  |
| 11. Review what to do with the responders who arrive during a fire emergency. |  |
| 12. Review what your department does in a tornado watch/warning. |  |
| 13. Location of eyewashes |  |
| 14. Location of safety shower |  |
| 15. Location of safety manuals (Lab Safety, Chemical Hygiene, Infection Control, etc.) |  |
| 16. Telephone number of police – 4444 or 4445 |  |
| 17. Report a fire – 4500 (CDD), 3600 (LDD) |  |

Signatures indicating completion of walkthrough: