Kaiser Permanente, South San Francisco

Northern California Region

=============== ===============

THIS DOCUMENT CONTAINS CONFIDENTIAL INFORMATION. Its use is restricted to employees with a need to know and third parties with a need to know and who have signed a non-disclosure agreement.

|  | Work Instruction |
| --- | --- |

| Title: | Maximum Dilution Verification | WI Number SSFWI-0304 Revision: 1 |
| --- | --- | --- |

| Department: Chemistry Area: SSF-KFH Laboratory 1200 El Camino Real South San Francisco CA 94080 | Approved & Released Work Instruction | Implementation Date: 11/14/2012 |
| --- | --- | --- |

| Type of Document: Analytic | Review Period - 700 Days |
| --- | --- |

1.	PURPOSE

For analytes that may have results falling outside the limits of AMR (Analytical

measurement Range), the laboratory procedure specifies the maximum

dilution that may be performed to obtain a linear reportable numeric result.

2.	INTENDED USE

2.1	For each analyte, the laboratory protocol should define the

maximum dilution that falls within the AMR and that can be

subsequently corrected by the dilution factor to obtain a

reportable numeric result.

2.2	Analytes for which a dilution protocol is unable to bring the

activity or concentration into the AMR should be reported as

"greater than" the highest estimated values

2.3	Maximum dilution is reviewed every two years by the laboratory

medical director who is responsible for establishing the maximum

allowable dilution of samples that will yield a credible laboratory

result for clinical use.

2.4	In a mixing procedure also termed as the "method of standard

additions", a known quantity of the analyte is mixed with the

unknown , and the concentration of the mixture is measured.

If equal volumes of the two samples are used, the result is

multiplied by two, the concentration of the known subtracted,

and the concentration of the unknown is the difference. e.g. PHTN.

2.5	For some analytes, an acceptable dilution protocol may not exist

because dilution would alter the analyte or the matrix causing

erroneous results such as Na+, K+, and Cl-.

2.6	For some analytes there may be no clinical relevance to reporting a

numeric result greater than a stated value such as BNP, Trop I and

BHCG.

3.	PROCEDURE

3.1 	Thaw patient sample that exceeds the AMR.

3.2	Depending on the assay, program the highest dilution recommended

by NCAL LQC.

3.3	Verification of linear response:

3.3.1	Program two dilutions for Fusion "on board" dilutions.

3.3.2	Compare results obtained from two dilutions.

3.3.3	If results agree within 10%, linear response is demonstrated.

3.3.3.1	If initial results do not agree within 10%, the

maximum dilution may be too high. Repeat

with a lower maximum dilution.

3.4	Verification of instrument dilution:

3.4.1	Make a manual dilution that corresponds to the maximum

dilution specified under 3.3.3.

3.4.2	Compare instrument dilution result with manually diluted

result.

3.4.3	If results are within 10% accuracy, instrument dilution is

verified.

3.4.3.1	If results do not agree within 10%, document

corrective action on back of the form

and repeat the process.

3.5	Maximum Dilution Verification Worksheet.

4.	REFERENCE

4.1	CHM.13720 Maximum Dilution, Chemistry and Toxicology Checklist,

p 18, Jan 4, 2012.

Associated Documents:

Document Revision History:

| Revision: 1 | Date Created: 01/07/2011 Date of Last Revision: 11/14/2012 | Last Approval Date: 11/14/2012 |
| --- | --- | --- |

| Document Author: Crisanta G Pajarillo/CA/KAIPERM |  |
| --- | --- |

Reason for Change:

| Revision: | Sec/Para Changed | Change Made: | Date |
| --- | --- | --- | --- |
| 1 | N/A | Initial Issue of Document | 10/12/2012 |

Approvals:

| Name: Arleen Sacay/CA/KAIPERM Title: Laboratory Administrative Director | Nov 14, 2012 05:35:04 PM PST - Approved by: Arleen Sacay/CA/KAIPERM |
| --- | --- |

| Name: James H Bainton/CA/KAIPERM Title: Laboratory Medical Director | Nov 14, 2012 07:35:36 PM PST - Approved by: James H Bainton/CA/KAIPERM |
| --- | --- |