Kaiser Permanente, South San Francisco

Northern California Region

=============== ===============

THIS DOCUMENT CONTAINS CONFIDENTIAL INFORMATION. Its use is restricted to employees with a need to know and third parties with a need to know and who have signed a non-disclosure agreement.

|  | Work Instruction |
| --- | --- |

| Title: | MedToxScan Urine Drug Screen | WI Number SSFWI-0362 Revision: 1 |
| --- | --- | --- |

| Department: Chemistry Area: SSF-KFH Laboratory 1200 El Camino Real South San Francisco CA 94080 | Approved & Released Work Instruction | Implementation Date: 12/04/2012 |
| --- | --- | --- |

| Type of Document: Analytic | Review Period - 700 Days |
| --- | --- |

1. METHOD

MEDTOXScan Drugs of Abuse Test System consists of the MEDTOXScanTest Devices

and the MEDTOXScanReader.  The MEDTOXScanTest Devices are one-step 			immunochromatographic tests for the rapid, qualitative detection of one or more of the

following in human urine:  Amphetamines, Barbiturates, Benzodiazepines, Cocaine, 		Methadone, Methamphetamine, Opiates, Oxycodone, Phencyclidine, and THC

(Cannabinoids), or their metabolites.  MEDTOXScanTest Devices can only be used with the 	MEDTOXScanReader.  The device cannot be visually read.  The MEDTOXScanReader is

an instrument used to interpret and report the results of the MEDTOXScanTest Device.

1.1	MEDTOXScanDrugs of Abuse Test System is for in vitro diagnostic use and is 			intended for prescription use only.  It is not intended for use in point-of-care settings.

1.2	MEDTOXScanDrugs of Abuse Test System detects drug classes at the following

cutoff concentrations:

1.2.1	AMP Amphetamine (d-Amphetamine)		500 ng/mL

1.2.2	BAR Barbiturates (Butalbital)			200 ng/mL

1.2.3	BZO Benzodiazepines (Nordiazepam)		150 ng/mL

1.2.4	COC Cocaine 					150 ng/mL

(Benzoylecgonine)

1.2.5	MAMP Methamphetamine			500 ng/mL

(d-Methamphetamine)

1.2.6	MTD Methadone (Methadone)			200 ng/mL

1.2.7	OPI Opiates (Morphine)				2000 ng/mL

1.2.8	OXY Oxycodone (Oxycodone)			100 ng/mL

1.2.9	PCP Phencyclidine (Phencyclidine)		25 ng/mL

1.2.10	THC Cannabinoids (11-nor-9-carboxy-9-THC)	50 ng/mL

1.3.	MEDTOXScanDrugs of abuse test system provides only a preliminary analytical test 		result.  A more specific alternate chemical method must be used in order to obtain a 		confirmed analytical result.

1.4	Gas chromatography / mass spectrometry (GC/MS), high performance liquid 			chromatography (HPLC) or liquid chromatography / tandem mass spectrometry 			(LC/MS/MS) are the preferred confirmatory methods.

1.5	Clinical consideration and professional judgment should be applied to any drug of

abuse test result, particularly when preliminary positive results are obtained.

2. CLINICAL SIGNIFICANCE

2.1	MEDTOXScanTest Devices utilize a one-step, solid-phase immunoassay technology

that includes the MEDTOXScanReader for a convenient automated result. This test

system may be used to screen urine samples for one or more of the following drug 			classes prior to confirmatory testing:

2.1.1	The amphetamines are a group of drugs that are central nervous system

stimulants.  This group includes amphetamine and methamphetamine.

Amphetamine (d-amphetamine) is detected on the Test Device only at the

(AMP) position, methamphetamine (MAMP) is detected at the (MAMP) 				position.

2.1.2	Barbiturates (BAR) are a group of structurally related prescription drugs that 			are used to reduce restlessness and emotional tension, induce sleep and to

treat certain convulsive disorders.

2.1.3	Benzodiazepines (BZO), a group of structurally related central nervous

system depressants, are primarily used to reduce anxiety and induce sleep.

2.1.4	Cocaine (COC) is a central nervous system stimulant.  Its primary

metabolite is benzoylecgonine.

2.1.5	Methadone (MTD) is a synthetic opioid used clinically as a maintenance

drug for opiate abusers and for pain management.

2.1.6	Opiates (OPI) are a class of natural and semi-synthetic sedative narcotic

drugs that include morphine, codeine and heroin.

2.1.7	Oxycodone (OXY) (Oxycontin, Percodan, Percocet) is a semi synthetic 				narcotic analgesic that is prescribed for moderately severe pain.  It is

available in both standard and sustained release oral formulations.

Oxycodone is metabolized to Oxymorphone and Noroxycodone.

2.1.8	Phencyclidine (PCP) is a hallucinogenic drug.

2.1.9	Marijuana (THC) is a hallucinogenic drug derived from the hemp plant.  				Marijuana contains a number of active ingredients collectively known as 				Cannabinoids.

2.2	Many factors influence the length of time required for drugs to be

metabolized and excreted in the urine.  A variety of factors influence the time

period during which drug metabolites are detected in urine: these include the

rate of urine production, the volume of fluid consumption, the amount of drug taken,

the urine pH, and the length of time over which drug was consumed.  Drinking

large volumes of liquid or using diuretics to increase urine volume will lower the

drug concentration in the urine and may decrease the detection period.  Lower 			detection levels may increase the detection time window.

2.3	Although the detection period for these drugs varies widely depending upon the 			compound taken, dose and route of administration and individual rates of

metabolism, some general times have been established and are listed below.1-5

| DRUG | DETECTION PERIOD |
| --- | --- |
| Amphetamine Acid Conditions Alkaline Condition | 1-3 days 3-10 days |
| Barbiturates Short-Acting Long-Acting | Up to 6 days Up to 16 days |
| Benzodiazepines | 1-12 days |
| Cocaine metabolite | Up to 5 days 1 to 3 days typical |
| Methadone | 1-3 days |
| Methamphetamine Acid Conditions Alkaline Conditions | 1-3 days 3-10 days |
| Opiates Heroin Morphine Codeine | 1 day 1-3 days 1-3 days |
| Oxycodone | 1-3 days |
| PCP Single Use Chronic Use | 1-8 days Up to 4 weeks |
| THC Single Use Chronic Use | 1-7 days Less than 30 days typical |

3. PRINCIPLE OF THE PROCEDURE

3.1	The MEDTOXScanDrugs of Abuse Test System includes the one-step, competitive, 		membrane-based immunochromatographic MEDTOXScanTest Device and the

MEDTOXScanReader, which interprets and reports the test results automatically.

A single urine sample can be evaluated for the presence of each of the classes of

drugs specified in a single MEDTOXScanTest Device. The MEDTOXScanTest

Device includes antibody-colloidal gold, drug-conjugates and a control line.

3.1.1	antibody-Colloidal gold  Mouse monoclonal antibodies were developed

that bind specifically to the drug class being tested.The individual

monoclonal antibodies were adsorbed to colloidal gold and dried onto the

test device.

3.1.2	drug-conjugates  Drugs from each class to be tested were individually 				conjugated to bovine serum albumin (BSA) or IgG.  Each drug conjugate

is immobilized on a test line at a designated position on the membrane strip.

3.1.3	Control LINE  Each test strip has anti-mouse antibody immobilized at the 			Control (C) position of the membrane strip.  The anti-mouse antibody will

bind excess antibody-colloidal gold, indicating that the reagents are working

properly.

3.2	When the urine sample is placed in the sample well of a test strip, the dried 			antibody-colloidal gold under the sample pad dissolves and the urine wicks up

the white strips carrying the reddish-purple antibody-colloidal gold with it.

The MEDTOXScan Drugs of Abuse Test System will detect specific classes of drugs

in urine because drug(s) in the urine and the drug(s) conjugated to the protein

compete to bind to the antibody-colloidal gold.  A test line will form when drug in the 		sample is below the detection threshold (negative result).

3.3	The MEDTOXScanReader scans the test device and utilizes a contact imaging

sensor 	(CIS) to capture relative line intensities.  Software algorithms and barcodes

are used to identify the test device, the drug tests associated with the test device

and whether the presence or absence of a line is associated with a negative or

positive result, respectively. The results of the scans are displayed on the

MEDTOXScanReader screen or can be printed.

3.3.1	Negative Samples

When no drug(s) is present in the urine sample, the reddish purple 				antibody-colloidal gold solutions migrate along the strip and bind to the 				respective drug conjugate(s) immobilized on the membrane.  Each strip has

up to 4 drug test lines labeled T1 – T4.  The binding of the antibody-colloidal

gold to the drug conjugate generates a line at the corresponding test

(T) position on the strip. The MEDTOXScanReader will scan each test

position and if a line is detected it will return “NEG” on the display screen

(or print out) next to the abbreviation for the drug test, indicating a

negative result.

3.3.2	Positive Samples

When drug(s) is present in the urine sample the antibody-colloidal gold

binds to the drug(s) before it migrates along the strip.  When

the antibody-colloidal gold binds to the drug(s) in the urine, it cannot

bind to the drug conjugate immobilized on the membrane and no

line is generated at the drug-specific position in the result window.

The MEDTOXScanReader will scan each test position and if no line is

detected it will return “POS” on the display screen (or print out) next

to the abbreviation for the drug test, indicating a preliminary positive result.

3.3.3	Control Line (Valid or Invalid results)

Each test strip has an internal procedural control.  A line must form

at the Control (C) position in the result window to indicate that

sufficient sample was applied and that the reagents are migrating properly.

If a Control line does not form, the test is invalid.  The MEDTOXScan

Reader scans each control line and returns “VALID” to the right of the drug

test result to confirm that the control line was detected.  If no control

line is detected it will return “INVALID” on the display screen (or print out)

next to the abbreviation for the invalid drug test, and no result will be given

for that drug test.

4. SAMPLE COLLECTION AND PREPARATION

4.1	The urine sample should be collected in a clean, dry container.

Approximately 75 µL is required for each sample well. Collection of 30 mL

of urine is more than sufficient for initial and subsequent testing.

4.2	No preservatives should be added.

4.3	Urine may be tested immediately following collection. If it is necessary to

store the urine, store under refrigeration at 2 to 8°C (36 to 46° F) for no

more than two days.

4.4	Urine may be frozen at -20°C (-4° F) or colder for storage. Stored urine must

be brought to ambient temperature (18 to 25°C/64 to 77°F) and mixed well

to assure a homogeneous sample prior to testing.

5. REAGENTS AND MATERIALS / STORAGE CONDITIONS

5.1	MEDTOXScanDrugs of Abuse Test System kit contains:

5.1.1	MEDTOXScanTest Devices for use with the

MEDTOXScanReader.

5.1.2	Each test device has all the reagents necessary to test one

urine sample for one or more drugs simultaneously on the

MEDTOXScanReader.

5.1.3	Each test device holds one or more test strips

composed of a membrane strip coated with drug conjugate

and a pad coated with antibody-colloidal gold in a protein

matrix.

5.1.4	Kit Contents

5.1.4.1	Twenty-five (25) test devices in individual foil

packages

5.1.4.2	Twenty-five (25) disposable pipette tips

5.1.4.3	One Quick Reference guide

5.1.5	Storage Conditions

5.1.5.1	The kit, in its original packaging (box),

should be stored at 2-8°C once received

from Material Services.

5.1.5.2	If removed from the refrigerated state (box),

should be stored at <25 °C.

5.1.5.3	Good until the expiration date on the label.

5.2	MEDTOXScanReader Contents

5.2.1	Positive and Negative QC Test Devices

5.2.2	Cleaning Cassette

5.2.3	Micro Pipettor (75 µl)

5.2.4	User Manual

5.2.5	Quick Set Up Guide

6. QUALITY CONTROL

6.1	The purpose of quality control is to ensure accuracy and reliability of results

and to detect errors. QC monitors the performance of the MEDTOXScan

Test Devices and the MEDTOXScanReader that uses a combination of

internal controls and external controls.

6.2	Internal controls ensure that the test is working and that the test are

performed correctly. A control line (internal control) is included on each

MEDTOXScanTest strip.  Whether or not drug is present in the sample,

a line must form at the Control (C) position on the test strip to show

that enough sample volume was used and that the reagents are

migrating properly.  If a Control line does not form, the test is invalid.

The Control line consists of immobilized anti-mouse antibody that reacts

with the antibody-colloidal gold as it passes this region of the

membrane.  Formation of a line detectable by the MEDTOXScanReader

verifies the Control line antibody-antigen reaction occurred.

6.3	External controls are urine-based control materials that contain the drugs

to be tested at concentrations above the cutoff (positive control) or

contain no drug (negative control).

6.3.1	External Controls

6.3.1.1	MEDTOX Negative Control, 5 mL (REF# 101183)

6.3.1.2	MEDTOX Positive Control, 5 mL (REF# 102138)

6.3.1.3	Store frozen once received from Material Services

6.3.1.4	Once opened, store refrigerated 2-8°C.

6.3.2	Run external controls routinely or as needed for any of  the

following reasons:

6.3.2.1	every Monday, done by AM shift at 9 AM.

6.3.2.2	when you open a new lot of devices

6.3.2.3	when you open a new box

6.3.2.4	if you suspect that the reader or test device is not

working properly

6.3.2.5	if you have had a repeated unexpected test result

6.3.2.6	if you suspect that the test devices have been

stored improperly.

6.3.3	Should control results indicate a problem with the MEDTOXScan

Drugs of Abuse Test System, please follow the instructions in the

Troubleshooting Section.

7. PROCEDURAL STEPS

7.1	Open one pouch for each sample to be tested and mark the MEDTOXScan

Test Device with the patient or sample identification (ID).  Make sure you

only mark within the ID box on the test device. (You may notice a

reddish-purple color in the sample well.  This is normal, do not discard

the test).

7.2	Drop 75 µl of urine to each sample well (indicated by an arrow).

7.3	Place the test device in the MEDTOXScanReader cassette drawer

and insert.

7.4	The MEDTOXScanReader will read the barcode on the test device and 				determine its part number and test configuration.

7.5	Scan Lot# of the cartridge.

7.6	Enter User ID# by typing your initial.

7.7	Enter Specimen ID# by scanning the accession label.

7.8	The MEDTOXScanReader will begin timing the assay once it detects

the barcode and results will be displayed after the scan and analysis

are complete.

8. READING and INTERPRETATION OF TEST RESULTS

8.1	The MEDTOXScanReader will automatically read the control and test

lines at the correct test position and display the test results for each drug.

The MEDTOXScanReader displays the results as either “NEG” for a

negative result, “POS” for a preliminary positive result, or “INVALID”

for an invalid result.  “VALID” will be displayed if valid results are obtained.

8.2	Valid:	The control line must be present for the test to be valid.

8.3	NEG:	A NEGATIVE test result for a specific drug indicates that

the sample does not contain the drug/drug metabolite above the cutoff level.

8.4	POS:	A preliminary POSITIVE test result for a specific drug indicates

that the sample may contain drug/drug metabolite near or above the

cutoff level.  It does not indicate the level of intoxication or the

specific concentration of drug in the urine sample.

8.4.1	Positive samples should be sent to a Regional Laboratory for

definitive testing.

8.5	Invalid:	The control line must be present for the test to be valid.

The absence of a control line indicates the test is invalid.  The urine

sample should be retested on a new test device.

9. CALIBRATION

9.1	No external calibration is required to perform the assay.

9.2	Each device is internally calibrated at specific threshold for each of the drugs being

tested.

9.3	The  internally calibrated threshold for each drug are as follows:

| Amphetamines | AMP | 500 ng/mL |
| --- | --- | --- |
| Barbiturates | BAR | 200 ng/mL |
| Benzodiazepines | BZO | 150 ng/mL |
| Cocaine | COC | 150 ng/mL |
| Methamphetamine | MAMP | 500 ng/mL |
| Methadone | MTD | 200 ng/mL |
| Opiates | OPI | 2000 ng/mL |
| Phencyclidine | PCP | 25 ng/mL |
| Tetrahydrocannabinoids | THC | 50 ng/mL |

10. RESULTING IN C. MILLENNIUM

10.1	The aforementioned threshold concentrations are used to separate a negative result

from a presumptive positive result. MEDTOXScanReader displays the results as

either "POS" if the result is at or above the cut-off value or "NEG" if the result is

below the cut-off value.

10.1.1	Urine drug screen with results above the cut-off values are resulted as

"T'FOLLOW".

10.1.2	Samples with "T'FOLLOW" results are sent to Regional Laboratory for

additional testing, including screening and confirmation if indicated.

11. CONFIRMATORY TEST

11.1	Special Handling:

11.1.1	Collect urine in a 90 mL plastic, leak proof container

11.1.2	Look up the provider's name that ordered the original drug screen and use the

same provider to order the confirmation test.

11.1.3	COE using the same accession number, change REPORTING PRIORITY to

RR and PROCEDURE to RM COLLECT.

11.1.4	Order specific confirmation procedure for the drug to be confirmed as

follows:

11.1.4.1		Order AMPHET CON for amphetamine/

methamphetamine confirmation.

11.1.4.2		Order BARB CON for barbiturates confirmation.

11.1.4.3		Order BENZO CON for benzodiazepines confirmation.

11.1.4.4		Order COCA CON for cocaine confirmation.

11.1.4.5		Order OPIATE CON for opiate confirmation.

11.1.4.6		Order PCP CON for phencyclidine confirmation.

11.1.4.7		Order THC CON for tetrahydrocannabinoids

confirmation.

11.1.4.8		Order U METHADON for methadone confirmation.

11.2	Test Procedures at Regional Lab:

11.2.1	All urine samples received at Regional Lab for confirmations will be

re-screened by EMIT and DRI immunoassays first before proceeding to

confirmation.

11.2.2	Confirmations will be performed by GCMS procedures for amphetamine/

methamphetamine, barbiturates, cocaine, opiates 2000, THC and PCP.

11.2.3	Benzodiazepines and methadone will be confirmed by LCMSMS.

11.3	Reporting:

11.3.1	Confirmed positive drugs will be reported as "DETECTED".

11.3.2	Confirmed negative drugs will be reported as "UNDETECTED".

11.4	Regional Laboratory Dwell Time:

11.4.1	Setup Time:

11.4.1.1		Confirmation will be performed the day after samples

received at Regional Lab.

11.4.1.2		GCMS confirmation is available Tuesday-Saturday,

Holidays excluded.

11.4.1.3		LCMSMS confirmation is available Tuesday-Saturday,

Holidays excluded.

11.4.1.4		LCMSMS confirmation for methadone and tricyclics

is available three times a week, holidays excluded.

10.4.2	Reporting Time:

10.4.2.1		Confirmation results will be available 24-72 hours after

setup.

12. TROUBLESHOOTING

12.1 	Use the QC Test Devices provided with the MEDTOXScanReader to detect errors

associated with the MEDTOXScanReader and a contaminated contact imaging

sensor (CIS) and to verify that the CIS cleaning procedure using the MEDTOXScan

Cleaning Cassette effectively removed any contamination (dirt, dust or sample).

12.2	The QC Test Devices function as an optical performance system check for the

MEDTOXScanReader only, not for the MEDTOXScanDrugs of Abuse Test

System, and they are not intended to replace the need for the external controls.

12.2.1	The QC Test Devices have been designed to simulate the end points that

are generated in the MEDTOXScanTest Device when external positive

and negative QC controls are run.

12.2.2	The QC Test Devices consist of artificial control lines and test lines

(negative) or artificial control lines and no test lines (positive) printed on a

membrane and placed in the MEDTOXScanTest Device plastic housing.

12.2.3	The QC Test Devices are not intended to evaluate all components of the

test system from specimen preparation through generation of results.

12.2.4	They are intended to function as a troubleshooting device to determine

that the reader optics are functioning correctly.

12.3	Run the QC Test Devices

12.3.1	if you suspect the MEDTOXScanReader is not functioning properly

12.3.2	if you suspect the CIS (optics) is dirty

12.3.3.	if the MEDTOXScanReader has been dropped or damaged.

13. MAINTENANCE

13.1	Routine Maintenance

13.1.1	No maintenance other than periodic external cleaning.

13.1.2	Use isopropyl alcohol wipes to clean the outside of the reader

and the device drawer.

13.2	Cleaning Cassette Procedure

This procedure is done

13.2.1	if you get incorrect results running the QC Test devices

13.2.2	if reader displays "CONTAMINATED SENSOR" error message.

13.2.3	effective for cleaning contamination such as dirt, dust or sample

accumulated on the contact imaging sensor (CIS).

13.2.4	Procedural steps

13.2.4.1		DISCONNECT reader from power supply.

13.2.4.2		Prepare Cleaning Cassette by rinsing under tap

water. Apply one drop of mild liquid hand soap/

detergent to the pad. Rinse thoroughly to remove

any soap remnants.

13.2.4.3		Place moist cleaning cassette in the reader drawer.

While holding down the EJECT button, slowly

move the drawer in and out 10 times. Use the "Tape

tail" attached to the Cleaning Cassette to pull

the drawer out while holding down the EJECT button.

13.2.4.4		If visible dirt is observed on the Cleaning Cassette

pad, wash, rinse and blot the Cleaning Cassette

again. Place the cassette in the drawer of the reader

and repeat step 13.2.4.3.

13.2.4.5		Re-connect power supply and power on the reader.

13.2.4.6		If the WELCOME screen appears, then run the QC

test devices to confirm that any contamination has

been removed from the sensor and the reader is

functioning properly.

13.2.4.7		If the "CONTAMINATED SENSOR" error message

displays again after the first cleaning procedure,

repeat steps 13.2.4.1 to 13.2.4.5

13.2.4.8		If the reader still displays the error message after

the Cleaning Cassette procedure has been conducted

twice, call Technical Support at 1-800-832-3244.

14. INTERFERENCE

14.1	pH

The MEDTOXScanDrugs of Abuse Test System was assayed with three

negative clinical samples with pH values of 4.0, 7.0 and 9.0 ± 0.1.  Each

sample was assayed in triplicate.  The pH samples were fortified with drug

concentrations that were the maximum level to give a strong negative

(95% or greater negative) result , and the minimum level above the cut-off

to give a strong positive (95% or greater positive) result.  All three pH

samples gave negative results when fortified to the maximum strong negative

level for each drug, and all gave positive results when fortified to the minimum

strong positive level for each drug.

14.2	Specific Gravity

The MEDTOXScanDrugs of Abuse Test System was assayed with three

samples with specific gravity values of 1.003, 1.015 and 1.030 ± 0.001.

Each sample was assayed in triplicate.  The specific gravity samples were

fortified with drug concentrations as described above for pH to give strong

negative and strong positive results.  All three specific gravity samples gave

negative results when fortified to the maximum strong negative level for each

drug, and all gave positive results when fortified to the minimum strong positive

level for each drug.

15. LIMITATION OF THE PROCEDURE

15.1	The MEDTOXScanDrugs of Abuse Test System is only use with

unadulterated human urine samples. Urine samples that are either extremely

acidic (below pH 4.0) or basic (above pH 9.0) may produce erroneous results.

If adulteration is suspected, obtain an additional specimen and re-test.

15.2	A presumptive positive result for any drug does not indicate the level of

intoxication, administration route or concentration of that drug in the urine

specimen.

15.3	A negative result may not necessarily indicate drug-free urine.  Negative results

can be obtained when drug is present but below the cut-off level of the test.

15.4	Place MEDTOXScanTest Devices in MEDTOXScanReader immediately after

adding the sample. Once the test device has been read in the MEDTOXScan

Reader, it must not be reinserted for a repeat reading, as the ten minute timing

will begin again.  If a repeat reading is required, rerun the sample on a fresh

test cassette.

15.5	The MEDTOXScanDrugs of Abuse Test System is not intended for use in

point-of-care settings.

15.6	There is a possibility that other substances and/or factors, e.g. technical or

procedural errors, may interfere with the test and cause false results.

15.7	Gas Chromatography/Mass Spectroscopy is the recommended confirmatory

method for most drugs.  HPLC or LC/MS/MS is the preferred confirmatory

method for Benzodiazepines.  Any of the drugs being tested for in the

MEDTOXScan Drugs of Abuse Test System may give a preliminary positive

result if ingested at prescribed therapeutic doses.

15.8	The MEDTOXScanDrugs of Abuse Test System cannot distinguish between

abused drugs and certain prescribed medications.  A positive test may be obtained

from certain foods or food supplements.

15.9	The PROFILE-V MEDTOXScanTest Devices must be used only with the

MEDTOXScanReader.  They cannot be visually read.

16. REFERENCES

16.1	Blum, K. Handbook of Abusable Drugs. Gardener Press, Inc. New York, New York,

1984. pp. 305-349.

16.2	DeCresce, R.P., Lifshitz, M.S., Mazura, A.C. and Tilson, J.E.

Drug Testing in the Workplace.  ASCP Press. American Society of Clinical

Pathologists. Chicago, Illinois. 1989. pp. 105-109.

16.3	Baselt, R.C. Disposition of Toxic Drugs and Chemicals in Man.  Seventh Edition.

Biomedical Publications. Foster City, California, 2004.

16.4	White, R.M. and Black, M.L. Pain Management Testing Reference. AACC Press.

Washington, DC. 2007.

16.5	Cary, P.L. The Marijuana Detection Window: Determining the Length of Time

Cannabinoids will Remain Detectable in Urine Following Smoking: A Critical Review

of Relevant Research and Cannabinoid Detection Guidance for Drug Courts,

Drug Court Review.  Volume V:1. 2005, pp. 23 – 58.

16.6	MedToxScan Reader User Manual.

16.7	Regional Lab Work Instruction -RLWI-2187

17. Associated Documents

17.1

17.2

17.3

Associated Documents:

Document Revision History:

| Revision: 1 | Date Created: 11/21/2012 Date of Last Revision: 12/04/2012 | Last Approval Date: 12/04/2012 |
| --- | --- | --- |

| Document Author: Crisanta G Pajarillo/CA/KAIPERM |  |
| --- | --- |

Reason for Change:

| Revision: | Sec/Para Changed | Change Made: | Date |
| --- | --- | --- | --- |
| 1 | N/A | Initial Issue of Document | 11-29-2012 |

Approvals:

| Name: Arleen Sacay/CA/KAIPERM Title: Laboratory Administrative Director | Dec 3, 2012 02:05:13 PM PST - Approved by: Arleen Sacay/CA/KAIPERM |
| --- | --- |

| Name: James H Bainton/CA/KAIPERM Title: Laboratory Medical Director | Dec 4, 2012 08:10:07 AM PST - Approved by: James H Bainton/CA/KAIPERM |
| --- | --- |