## Slide 1: i-STAT/p-Web Interface
Go-Live Information
Prepared by Angie Thayer, BSMT (ASCP),
Clinical Laboratory Compliance, Quality, and Safety
09/16/15

## Slide 2: Information
The i-STAT/p-Web interface will have new processes and will include a learning curve for all areas. Therefore, this information may be updated as we begin to use this new interface.

## Slide 3: Process Flow
Physician places order in Wake One for i-STAT test
CSN for that specific date of service or current admission CSN should be entered by testing staff member into i-STAT handheld
After testing is completed, analyzer should be downloaded. Results flow to p-Web. P-Web searches it’s ADT for a CSN match. If match found, results post to Wake One.
Billing occurs when results hit Wake One. Billing occurs based on results that hit Wake One, Not on the physician order for the test.

## Slide 4: Significant Changes for Main Campus
Mechanism for how results get electronically transferred from the i-STAT handheld to Wake One will be different. It is a new interface.
Comment codes can be used by testing staff member to hold results from Wake One or Cause no-billing for testing
Staff members will need to edit their own results in Wake One, in the event of an error—such as sample misidentification, questionable results that were not flagged with a comment code at the time of testing, or inappropriate temperature corrected blood gas results.
The name of the person performing testing will now appear with the results in Wake One.
If liquid QC fails, a comment code is required
  —Use comment code 5—recheck/confirm---repeat liquid QC

## Slide 5: Citrate dialysis testing
Must enter C in field 1 of the i-STAT analyzer to identify the sample as a citrate dialysate sample, rather than a true patient sample.
Results will post in Wake One with the comment: Whole Blood Dialysis Circuit

## Slide 6: Contact Serial Number (CSN)
Do not add zeros to the CSN
You must use the current visit CSN. Verify the information Prior to scanning/entering the ID into the i-STAT handheld
It is not okay to enter the medical record number with added zeros as the patient ID

## Slide 7: Changes that WILL impact your test site—Comment Codes
We will use comment codes to better ‘automate’ result hold or posting to the patient record—MUST be entered at the time of testing for effect.
COMMENT CODES ---FOR ISTAT (same as glucose)
    0 translates to Procedure Error
    Results flagged with comment code 0 will NOT post to the patient electronic healthcare record.
    This code should be used when the staff member does not believe the result and feels an error was made in testing.
    5 translates to Recheck/confirm
    Results flagged with comment code 5 will post to the patient electronic healthcare record but the patient will not be billed for testing.
    6 translates to To notify Provider
COMMENT CODES--NEW FOR i-STAT (not currently defined for glucose)
  123 translates to laboratory confirmatory testing to follow-
  Results flagged with comment code 123 will post to the patient electronic healthcare record but the patient will not be billed for testing.

## Slide 8: Changes that WILL impact your test site—Star Out results
We will change the way we manage star out results
--To stop a star out result from posting to the patient record, comment code 0 will need to be entered into the analyzer by the testing staff member
If comment code 0 is not entered into the i-STAT handheld, results will NOT be held in p-Web. They will post automatically to Wake One.
The star out results will report as INSTRUMENT ERROR and additional NOCHARGE code will be sent to Wake One in the interface message.
When the NOCHARGE code is seen in Wake One, ALL tests on that particular cartridge will NOT bill. They will appear as a $0 charge.
There is a Wake One report that will include all results that posted as INSTRUMENT ERROR. For quality monitoring, this will allow us to track the number of star out results that occur.

## Slide 9: i-STAT Star Outs: How do I handle these?
If the testing staff member obtains a star-out test (***) on the handheld,
1) They should decide if they are going to repeat the full cartridge.
    A) If they will repeat the full cartridge, they should enter comment code 0 (zero) in the i-STAT handheld.
    B) When p-Web receives the results, and sees comment code 0 (zero), results will go to the PROCEDURE ERROR bucket and NOT post to Wake One. The results will remain in this bucket until they are reviewed and acknowledged by the laboratory. It is not intended that these results will ever be sent to the patient record (Wake One)
2) If the staff member cannot repeat testing on another cartridge for some reason and desires to use any results that actually reported a number value,
    A) they will NOT enter a comment code and proceed as normal.
    B) When p-Web receives the results, they will be automatically sent to Wake One via the interface. The star-out tests will report as INSTRUMENT ERROR
    C) INSTRUMENT ERROR in the interface message will also cause the NOCHARGE message, which will prevent any test results associated with this cartridge from billing to the patient. The test will appear on the Wake One Revenue and Usage report as $0
    D) We will monitor these star-out results as a quality indicator using the Wake One Instrument Error report.

## Slide 10: Changes that WILL impact your test site: Liquid QC Value Assignment Ranges
All analyzers will be programmed with electronic QC value assignments for the liquid QC. The analyzer will note whether or not the liquid QC passes or fails. If a value fails, a comment code will be required. However, patient testing lockout does Not occur.
Another reminder: Staff MUST ensure QC results are within acceptable limits PRIOR to proceeding with patient testing.
They should NEVER REPORT PATIENT RESULTS IF QC HAS FAILED!
If QC results are not within expected limits, then patient results may also be affected.
This has a direct impact on safe patient care. Inappropriate patient treatment may occur if inaccurate laboratory results are reported.

## Slide 11: Comment Codes: Liquid QC
Liquid QC that is not within acceptable tolerance limits will require a comment code
Main campus will use comment code 5—recheck/confirm
This has been put into place to create a STOP point for the testing staff member.

## Slide 12: Liquid QC Important Notes/Reminders
ALL failed liquid QC results must have corrective/follow-up action, according to your site’s established protocol.
i-STAT will NOT lock out a user from performing patient testing if liquid QC failures occur.
Staff Must ensure liquid QC values are within acceptable range or “PASS”
    If QC fails – Patient testing should STOP!!
    The testing staff member should begin troubleshooting procedures
Patient results may not be accurate if QC does not perform as expected –
  NEVER report patient results on a failed quality control event.
If a staff member proceeds with patient testing on a failed QC event, then the test site is out of compliance. Patient safety is at risk.

## Slide 13: Changes that WILL impact your test site: Units of Measure—Notify Providers Prior to Go-Live
Units of measure is changing for certain analytes. Changed to match Core Lab
—mEq/L changing to mmol/L for sodium, potassium, chloride, TCO2, anion gap
Follow-up with your site’s providers to ensure they are aware of these changes Prior to go-live.

## Slide 14: Changes that may impact your test site: CPB Hematocrit correction
CPB hematocrit option will not be active for any test site. Will be deactivated from analyzers.

## Slide 15: Changes that WILL impact your test site: Order of Screen Display
The i-STAT analyzer display will be slightly different. The order of items –sample type, temperature, field 1,2,3 ---may be different for some sites.

## Slide 16: Changes that may impact your test site: Test Select
All sites will now need to select the tests they want to run—even if only testing a single test cartridge

## Slide 17: Changes that WILL impact your test site: Information Reported to Wake One
Testing personnel name, i-STAT analyzer serial number, and cartridge lot number will now report with each test result in Wake One.

## Slide 18: i-STAT Operator ID
Employee ID---user specific
Serves as identification of testing personnel and will be reported with patient results in Wake One

## Slide 19: Editing of results once they have already posted to Wake One
MUST be done in Wake One
All staff members will have to complete some form of Wake One training for editing interfaced results.
This access is needed if a sample misidentification occurs or other result modification is needed.

## Slide 20: Correcting a sample misidentification
Notify providers of affected patients
In Wake One, open the patient chart---select More Activities
Go into Enter/Edit Results
  —Click Filter
  —Select Point of Care Testing Docked Device
  —Radio Button ALL
  —Enter Date range—Click Accept—
  Highlight the results that need edit--Double click to open edit screen—ENSURE YOU SELECT THE CORRECT TEST RESULTS FOR EDIT!!!--VERY IMPORTANT
At Component Value (the results), replace All i-STAT values with “XXX”. If a comment is present, XXX out the comment, as well.
In ONE OF THE COMMENT sections for each set of results, enter this text: INVALID RESULT, PATIENT SAMPLE MISIDENTIFICATION, NOTIFIED enter the full name of the provider that was notified with credentials and the date and time of notification BY the name of the person that notified the provider
Click Accept.
File an RL6 Report. Email RL6 number to the Laboratory Compliance, Quality, and Safety office: athayer@wakehealth.edu, rddyer@wakehealth.edu, engrego@wakehealth.edu

## Slide 21: Correcting erroneous temperature corrected blood gas results in Wake One
Notify provider of affected patient
In Wake One, open the patient chart---select More Activities
Go into Enter/Edit Results
  —Click Filter
  —Select Point of Care Testing Docked Device
  —Radio Button ALL
  —Enter Date range—Click Accept—
  Highlight the results that need edit--Double click to open edit screen—ENSURE YOU SELECT THE CORRECT TEST RESULTS FOR EDIT!!!--VERY IMPORTANT
At Component Value (the results), replace the erroneous patient temperature entered into the i-STAT handheld and the temperature corrected pH, pCO2, and pO2 values with “XXX”.
In the patient temperature COMMENT section, enter this text: INVALID PATIENT TEMPERATURE, NOTIFIED enter the full name of the provider that was notified with credentials and the date and time of notification BY the name of the person that notified the provider
Click Accept.
File an RL6 Report. Email RL6 number to the Laboratory Compliance, Quality, and Safety office: athayer@wakehealth.edu, rddyer@wakehealth.edu, engrego@wakehealth.edu

## Slide 22: Flagging questionable i-STAT results if posted to Wake One
Notify provider of affected patient
In Wake One, open the patient chart---select More Activities
Go into Enter/Edit Results
  —Click Filter
  —Select Point of Care Testing Docked Device
  —Radio Button ALL
  —Enter Date range—Click Accept—
  Highlight the results that need edit--Double click to open edit screen—ENSURE YOU SELECT THE CORRECT TEST RESULTS FOR EDIT!!!--VERY IMPORTANT
At Component Value (the results), replace ALL questionable results with “XXX”.
In the first COMMENT section, enter this text: QUESTIONABLE I-STAT RESULTS, NOTIFIED enter the full name of the provider that was notified with credentials and the date and time of notification BY the name of the person that notified the provider
Click Accept.
File an RL6 Report. Email RL6 number to the Laboratory Compliance, Quality, and Safety office: athayer@wakehealth.edu, rddyer@wakehealth.edu, engrego@wakehealth.edu

## Slide 23: Flagging tests that were not ordered by the physician
Notify provider of affected patient
In Wake One, open the patient chart---select More Activities
Go into Enter/Edit Results
  —Click Filter
  —Select Point of Care Testing Docked Device
  —Radio Button ALL
  —Enter Date range—Click Accept—
  Highlight the results that need edit--Double click to open edit screen—ENSURE YOU SELECT THE CORRECT TEST RESULTS FOR EDIT!!!--VERY IMPORTANT
DO NOT REMOVE ANY TEST RESULTS
In the first COMMENT section, enter this text: TESTING PERFORMED WITHOUT PHYSICIAN ORDER—CREDIT PENDING, NOTIFIED enter the full name of the provider that was notified with credentials and the date and time of notification BY the name of the person that notified the provider
Click Accept.
File an RL6 Report. Email RL6 number to the Laboratory Compliance, Quality, and Safety office: athayer@wakehealth.edu, rddyer@wakehealth.edu, engrego@wakehealth.edu

## Slide 24: For DHP Phlebotomy, Family Practice Lab, and Davie Lab staff ONLY! Documentation of critical value notification
Notify provider of affected patient
In Wake One, open the patient chart---select More Activities
Go into Enter/Edit Results
  —Click Filter
  —Select Point of Care Testing Docked Device
  —Radio Button ALL
  —Enter Date range—Click Accept—
  Highlight the results that need edit--Double click to open edit screen—ENSURE YOU SELECT THE CORRECT TEST RESULTS FOR EDIT!!!--VERY IMPORTANT
In the COMMENT section for the specific test with the critical value, enter this text: CRITICAL VALUE CALLED TO AND READ BACK BY, enter the full name of the provider that was notified with credentials and the date and time of notification BY the name of the person that notified the provider
Click Accept.

## Slide 25: Questions?
Contact the Laboratory Compliance, Quality, and Safety Office
3-4285—Melanie Haire
3-4136—Angie Thayer
3-4137—Ray Dyer
3-0377---Liz Gregory
istatpwebgloliveinformationmaincampus091615