## Slide 1: i-STAT Creatinine Education Module
Prepared by Angie Thayer, BSMT (ASCP), Clinical Lab POCT Coordinator
Updated by Liz Gregory 11/11/2015

## Slide 2: i-STAT Support
Angie Thayer 713-4136
Ray Dyer 713-4137
Liz Gregory 713-0377
OR Blood Gas Lab 716-2181 (M-F 7am-7pm)
Refer to the i-STAT System Manual and WFBMC specific i-STAT policies
WFBMC POCT website

## Slide 3: i-STAT Policies/Procedures/Guidelines
Posted on the WFBMC Intranet
Department—Point of Care Testing—Policies
The i-STAT System Manual (Abbott POC website) can be found under Resources-- Resource Links—Manuals

## Slide 4: i-STAT Operator ID
User specific
Serves as identification of testing personnel
Do NOT share operator ID’s
Do NOT enter your operator ID into analyzer and allow use by another individual

## Slide 5: Training and Competency Requirements
Must be documented
Initial training
Initial competency evaluation
6 month competency evaluation-NEW users
On-going Annual competency evaluation—must meet all 6 points of competency assessment, as defined by CLIA
Failure to maintain updated training and competency will result in loss of testing privileges
Competency observation may only be completed by an authorized staff member

## Slide 6: Safety
Always wear gloves when handling analyzer
        Including patient testing and performing QC
Disinfect when contaminated with blood AND between EACH patient / Follow WFBMC policies
DO NOT lean over the cartridge while filling with sample
A safety shield is recommended
Place gauze or tissue over snap when closing

## Slide 7: i-STAT Components
Analyzer
Cartridge
Simulator (Electronic Quality Control)
Liquid Quality Control Materials
Downloader
Recharger

## Slide 8: i-STAT Analyzer (Handheld)
Volt Meter
Testing occurs on the test cartridge bio-sensors
Rechargeable batteries are utilized—care should be taken to keep batteries charged—do not discard rechargeable batteries

## Slide 9: i-STAT Analyzer
Will turn off after 2 minutes of non-use
Can be turned on by pressing the on/off keypad—circle with line
If operator and patient information has been entered, the analyzer will remain on for 15 minutes

## Slide 10: Analyzer display
‘Cartridge/Simulator Locked’
      Cartridge or simulator locked in analyzer
      DO NOT remove when this message is displayed
Flashing battery icon or ‘Low Battery’ indicates battery voltage is low and batteries need to be re-charged

## Slide 11: Menu Options
Analyzer Status —allows viewing of battery voltage - recommend maintaining charge >8 volts
Data Review —allows review/print of data stored in analyzer
Quality Tests —quality control checks performed utilizing this menu
Other Menu options —refer to i-STAT System Manual for additional information

## Slide 12: Analyzer Download
Use the Downloader located in each site
Analyzers should be downloaded after each case
It is the responsibility of testing personnel to download i-STAT analyzers
You can’t “over-download”

## Slide 13: Analyzer Download
The downloader will display “Communication in Progress”, indicating successful download of data
An error message will display on the analyzer if results do not successfully transmit. If an error occurs:
      Troubleshoot the downloader
      Check for a green power light
      Check all cord/cable connections to the down loader
      Try unplugging and re-plugging the cables
      If troubleshooting does not resolve problem, report problem to the Help Desk

## Slide 14: Re-charger
Can charge battery in compartment and battery in analyzer
Rotate analyzers on chargers to ensure they stay adequately charged

## Slide 15: i-STAT Supplies
Cartridges and paper can be obtained from the OR Blood Gas Lab M-F 7am-7pm
Care should be taken to get only the amount of cartridges which will be used within the 2 week room temperature expiration date
Replacement analyzers can be obtained from the Clinical Laboratory Point-of-Care Testing Office

## Slide 16: i-STAT Cartridges
Testing occurs in the cartridge
Cartridges have bio-sensors that “measure” the analytes tested
There are different cartridge configurations available from i-STAT

## Slide 17: Cartridge Storage
Cartridges are stored refrigerated 2 to 8˚C until the manufacturer’s expiration date.
Do NOT use past manufacturer’s expiration date.
Can be stored at room temperature (18 to 30˚C) per manufacturer instructions. Creatinine cartridges have a 14 day expiration date at room temperature.
DO NOT use past the room temperature expiration date!
If a small supply of cartridges are kept refrigerated at the POC site, the refrigerator temperature should have 24/7/365 temperature monitoring. The refrigerator temperature must be maintained at 2 to 8˚C.
If the temperature is ever outside of 2 to 8˚C, contact the Clinical Laboratory Point-of-Care Testing Office for assistance.
Cartridges should NOT be returned to the refrigerator once they have been stored at room temperature.

## Slide 18: Emergency Release of Cartridges
An i-STAT Emergency Cartridge Release form must be completed should i-STAT cartridges be needed outside of M-F 7am-7pm.
The site picking up the cartridges is responsible for room temperature expiration dating of the cartridges.
NEVER take cartridges that are marked with “Do Not Use”.

## Slide 19: Cartridge Handling - DO NOT’s
DO NOT use “quick heating”—holding close to body or near a warm surface
DO NOT store or expose cartridges to extreme cold or heat, such as freezing temperatures or storing above lights or computers
Do NOT re-refrigerate cartridges once they have warmed to room temperature
Do NOT touch bio-sensors
Do NOT pre-rupture silver calibrant disc
Do NOT re-use cartridges
Do NOT leave cartridge exposed to air and moisture
Do NOT use expired cartridges for patient testing

## Slide 20: Cartridge Handling
Individual cartridges can be used after 5 minutes at room temperature
A box of 25 cartridges must sit at room temperature for 1 hour prior to use
Enter operator and patient ID information prior to collecting the sample
Open by tear symbol
Use cartridge IMMEDIATELY after opening
Fill to blue triangle—leave “blood dome”
Insert into analyzer immediately after filling with sample
Do not move analyzer after cartridge has been inserted

## Slide 21: Simulators
Electronic Quality control that is used to validate the i-STAT analyzer
      Pass or Fail
QC results are automatically documented when the i-STAT analyzer is downloaded

## Slide 22: Simulators - Internal
An INTERNAL electronic simulator is performed automatically by the analyzer every 8 hours of use with each cartridge type.
The internal simulator is performed when a patient cartridge is inserted.
Results will not be given if the simulator fails.
If the simulator passes, patient results will be displayed.

## Slide 23: Simulators - External
The EXTERNAL electronic simulator is the same as the internal simulator but is an external device and can be tested upon demand.

## Slide 24: Simulators - External
The external simulator should be used:
        If the analyzer is dropped
        If the internal simulator fails
        If an error code occurs that indicates the simulator should be tested
        If analyzer performance is in question
        Prior to and after performing the ceramic cartridge cleaning procedure

## Slide 25: Liquid Quality Control
Used to validate performance of test cartridges
Consists of multiple levels
        Test cartridge specific
Should be performed:
      On each new cartridge shipment per cartridge type per cartridge lot # PRIOR to patient use—by Clinical Lab staff
      Monthly—by testing site staff
      If cartridge/analyzer performance is in question

## Slide 26: Liquid QC Important Notes/Reminders
All analyzers will be programmed with electronic QC value assignments for the liquid QC. The analyzer will note whether or not the liquid QC passes or fails. If a value fails, a comment code will be required. However, patient testing lockout does NOT occur.
If liquid QC fails, a comment code is required:
      Use comment code 5—recheck/confirm---repeat liquid QC
      This has been put into place to create a STOP point for the testing staff member
Staff MUST ensure QC results are within acceptable limits PRIOR to proceeding with patient testing.

## Slide 27: Staff should NEVER REPORT PATIENT RESULTS IF QC HAS FAILED!
ALL failed liquid QC results must have corrective/follow-up action, according to your site’s established protocol.
i-STAT will NOT lock out a user from performing patient testing if liquid QC failures occur.
Staff Must ensure liquid QC values are within acceptable range or “PASS”
      If QC fails – Patient testing should STOP!!
      The testing staff member should begin troubleshooting procedures.
Patient results may not be accurate if QC does not perform as expected –
      NEVER report patient results on a failed quality control event.
If a staff member proceeds with patient testing on a failed QC event, then the test site is out of compliance. Patient safety is at risk.

## Slide 28: Process Flow
The testing staff member is responsible for ensuring that there is a documented physician order or documented protocol before i-STAT testing is performed.
CSN for that specific date of service or current admission CSN should be entered by testing staff member into i-STAT handheld.
After testing is completed, analyzer should be downloaded. Results flow to p-Web. P-Web searches it’s ADT for a CSN match. If match found, results post to Wake One.
Operator’s name, i-STAT analyzer serial number, and cartridge lot number will report with each test result in Wake One.
Billing occurs when results hit Wake One. Billing occurs based on results that hit Wake One, not on the physician order for the test.

## Slide 29: Patient identity
Follow WFBMC policy for verification of patient identity— Use 2 patient identifiers at the patient bedside and 4 identifiers (name, DOB, MRN, CSN) in sites which perform procedural cases where patient armband is not accessible.
If scanning a patient armband or label not on the patient, include patient ID verification in time out procedure.
Ensure sample identity throughout entire testing and reporting process
Press and hold the SCAN key pad
Scan patient ID into the analyzer DIRECTLY from the patient armband on the patient or in the presence of the patient
Verify patient ID again on analyzer display
CAUTION: Laser light—do not stare into beam or point at anyone

## Slide 30: Patient Identity-Misidentifications
Any mis-identified samples/results should be reported to the Clinical Lab immediately.
In the event of an error, operators will need to edit their own results in Wake One – for example, sample misidentification or questionable results that were not flagged with a comment code at the time of testing.
It is the responsibility of the staff member who creates the error to notify appropriate patient care staff (MD) of the sample misidentification.

## Slide 31: Sample Requirements Creatinine Cartridge
ONLY whole blood should be tested on i-STAT
Non-anti-coagulated whole blood for creatinine testing should be tested within 3 minutes of collection
Heparin is the ONLY acceptable anti-coagulant which may be used with the i-STAT creatinine cartridge
    MUST use the BD Vacutainer lithium heparin green top 4ml tube— FILL TO CAPACITY or results may be adversely affected
    Heparinized samples may be tested up to 30 minutes after collection

## Slide 32: In-Dwelling Line Sample Collection
  Always adequately ‘clear’ line prior to collecting a sample for testing on i-STAT.
  Inaccurate results will be given if samples are contaminated.
  Back flush line with sufficient amount of blood to remove intravenous solution, heparin, or medications that may contaminate the sample.
  i-STAT Recommendation: five to six times the volume of the catheter, connectors, and needle should be collected as ‘waste’ for creatinine samples.
  Accurate results depend on an adequate back flush to eliminate the possibility of sample contamination with IV fluids.
  Caution should be taken when collecting from lines which have had fluids that could adhere to the sides of the tubing. These lines may be difficult to adequately ‘clear’.

## Slide 33: CREATININE - Sample Considerations
ALWAYS use a well mixed sample
Mix sample for 15 seconds
Always squirt out the first drop of blood from syringe samples
      To check for clots and to get rid of any micro air bubbles
NEVER run a sample that has or has had a clot. Inaccurate results may be obtained!

## Slide 34: Factors affecting results
Moving/carrying the analyzer during sample testing
Vibration of the analyzer
Tilting the analyzer during testing - the analyzer MUST remain flat during sample testing
Refer to the i-STAT System Manual—Cartridge Test Information (CTI) sheets for other factors which can adversely affect results

## Slide 35: Result Considerations
Samples diluted with IV fluids will give inaccurate results
Results should be carefully evaluated
Contamination may be indicated by:
        Falsely decreased creatinine value

## Slide 36: Result considerations
Unexpected and unexplained results should be repeated by another test method.
Problems should be reported to:
        Angie Thayer 3-4136
        Ray Dyer 3-4137
        Liz Gregory 3-0377

## Slide 37: When Results Not Given
Cartridge error
    Refer to Technical Bulletin --- ‘Analyzer Coded Messages’
    Do NOT continue with testing until the cause of the error is clearly understood
    If the same error code occurs multiple times with no apparent cause, notify the Clinical Laboratory POCT Coordinator
    Consider using a different analyzer and a different batch of test cartridges
Star out (***)
    Occurs when specific bio-sensor is compromised or there is an interfering substance in the sample
Out of instrument range (>x value or <x value)
    Result is out of reportable range for the analyzer

## Slide 38: Analyzer Error Messages: Codes 20, 23, 43
Be aware that codes 20, 23, 43, and others indicate that the ceramic cartridge conditioning may need to be performed.
Every site should have one of these cartridges.
Instructions are in the i-STAT Technical Bulletin—Analyzer Coded Messages.
This cartridge should NOT be used as preventive maintenance. It should only be used when a quality check code indicates.
Routine use can damage the internal mechanism of the i-STAT handheld.

## Slide 39: Analyzer Error Messages: Codes 79 and 80
Codes 79 and 80 may indicate that the thermal probes are dirty and need cleaning.
Refer to “i-STAT Thermal Probe Cleaning” procedure on the POCT website.
        Go to INTRANET – DEPARTMENTS – POCT – EDUCATION TOOLS AND COMPETENCY TESTING – i-STAT –“i-STAT TROUBLESHOOTING: ERROR CODES 79 AND 80”

## Slide 40: Use of Comment Codes
    We will use comment codes to better ‘automate’ result hold or posting to the patient record—MUST be entered at the time of testing for effect.
COMMENT CODES --- FOR ISTAT (same as glucose)
    0 translates to --- Procedure Error
        Results flagged with comment code 0 will NOT post to the patient electronic healthcare record.
        This code should be used when the staff member does not believe the result and feels an error was made in testing.
    5 translates to --- Recheck/confirm
        Results flagged with comment code 5 will post to the patient electronic healthcare record but the patient will not be billed for testing.
    6 translates to --- To notify Provider
    123 translates to --- Lab Confirmation to Follow
        Results flagged with comment code 123 will post to the patient electronic healthcare record but the patient will not be billed for testing.

## Slide 41: Star Out Results
To stop a star out result from posting to the patient record, comment code 0 will need to be entered into the analyzer by the testing staff member.
Retest the sample and if star out values continue, send the sample to the Clinical lab for creatinine testing.
If comment code 0 is not entered into the i-STAT handheld, the result will NOT be held in p-Web. It will post automatically to Wake One.
The star out result will report as INSTRUMENT ERROR and an additional NO CHARGE code will be sent to Wake One in the interface message.
There is a Wake One report that will include all results that posted as INSTRUMENT ERROR. For quality monitoring, this will allow us to track the number of star out results that occur.

## Slide 42: Creatinine GFR Results
Estimated Glomerular Filtration Rate (GFR) values will be reported when results are downloaded into the electronic medical record.
The values may be accessed via Wake One.
The i-STAT handheld will NOT report GFR values.

## Slide 43: Result Recall
Turn on analyzer
Select MENU
Select 2 DATA REVIEW
Select desired data
PATIENT DATA requires entry of patient and operator ID information

## Slide 44: i-STAT Maintenance
Recharge batteries
Disinfect outside of analyzer between patients
Do not get moisture inside of the analyzer
Perform ceramic cartridge pin conditioning when indicated by analyzer error codes (scheduled conditioning is not recommended)
      Each site has a ceramic cartridge and instructions for performing this procedure

## Slide 45: Editing of results once they have already posted to Wake One
MUST be done in Wake One
All staff members will have to complete some form of Wake One training for editing interfaced results
This access is needed if a sample misidentification occurs or other result modification is needed

## Slide 46: Correcting a Sample Misidentification
Notify providers of affected patients
In Wake One, open the patient chart---select More Activities
Go into Enter/Edit Results
    —Click Filter
    —Select Point of Care Testing Docked Device
    —Radio Button ALL
    —Enter Date range—Click Accept—
    Highlight the results that need edit--Double click to open edit screen—ENSURE YOU SELECT THE CORRECT TEST RESULTS FOR EDIT!!!--VERY IMPORTANT
At Component Value (the results), replace All i-STAT values with “XXX”. If a comment is present, “XXX” out the comment, as well.
In ONE OF THE COMMENT sections for each set of results, enter this text: INVALID RESULT, PATIENT SAMPLE MISIDENTIFICATION, NOTIFIED enter the full name of the provider that was notified with credentials and the date and time of notification BY the name of the person that notified the provider
Click Accept.
File an RL6 Report. Email RL6 number to the Laboratory Compliance, Quality, and Safety office: athayer@wakehealth.edu, rddyer@wakehealth.edu, engregor@wakehealth.edu

## Slide 47: Flagging questionable i-STAT results if posted to Wake One
Notify provider of affected patient
In Wake One, open the patient chart---select More Activities
Go into Enter/Edit Results
    —Click Filter
    —Select Point of Care Testing Docked Device
    —Radio Button ALL
    —Enter Date range—Click Accept—
    Highlight the results that need edit--Double click to open edit screen—ENSURE YOU SELECT THE CORRECT TEST RESULTS FOR EDIT!!!--VERY IMPORTANT
At Component Value (the results), replace ALL questionable results with “XXX”.
In the first COMMENT section, enter this text: QUESTIONABLE I-STAT RESULTS, NOTIFIED enter the full name of the provider that was notified with credentials and the date and time of notification BY the name of the person that notified the provider
Click Accept.
File an RL6 Report. Email RL6 number to the Laboratory Compliance, Quality, and Safety office: athayer@wakehealth.edu, rddyer@wakehealth.edu, engregor@wakehealth.edu

## Slide 48: Flagging tests not ordered by the physician
Notify provider of affected patient
In Wake One, open the patient chart---select More Activities
Go into Enter/Edit Results
  —Click Filter
  —Select Point of Care Testing Docked Device
  —Radio Button ALL
  —Enter Date range—Click Accept—
  Highlight the results that need edit--Double click to open edit screen—ENSURE YOU SELECT THE CORRECT TEST RESULTS FOR EDIT!!!--VERY IMPORTANT
DO NOT REMOVE ANY TEST RESULTS
In the first COMMENT section, enter this text: TESTING PERFORMED WITHOUT PHYSICIAN ORDER—CREDIT PENDING, NOTIFIED enter the full name of the provider that was notified with credentials and the date and time of notification BY the name of the person that notified the provider
Click Accept.
File an RL6 Report. Email RL6 number to the Laboratory Compliance, Quality, and Safety office: athayer@wakehealth.edu, rddyer@wakehealth.edu, engregor@wakehealth.edu

## Slide 49: RL6 Report
Ensures the patient receives appropriate credit for tests billed to their account
Allows review of edited i-STAT results
Tracks i-STAT occurrences for quality assurance and quality improvement purposes

## Slide 50: College of American Pathologists (CAP)
If there are concerns regarding result quality and the concerns are not adequately addressed by Lab or WFBMC management, CAP can be notified at 1-866-236-7212.

## Slide 51: Questions?
Contact the Laboratory Compliance, Safety, and Quality Assurance Office
      3-4285—Melanie Haire
      3-4136—Angie Thayer
      3-4137—Ray Dyer
      3-0377---Liz Gregory

## Slide 52: Congratulations
This concludes the i-STAT Creatinine Education power point presentation.
Please note that this power point does not replace the i-STAT policies/procedures/guidelines.
The presentation only gives highlights of i-STAT information.
Please complete the online i-STAT exam within 1 week of training.

## Slide 53: Hands On Experience
Change batteries and check battery voltage
Test External electronic simulator
Test/Discuss Liquid QC
Test patient sample
Recall results
Discuss Menu Options
Download of results
        IstatCreatinineEducationModule111115