## Slide 1: Point of Care Testing
i-STAT G3, CG4, EG7, Creatinine Education Module
Prepared by Angie Thayer, BSMT (ASCP), Clinical Lab POCT Coordinator
Updated by Liz Gregory 12/22/15

## Slide 2: i-STAT Support
Angie Thayer 713-4136
Ray Dyer 713-4137
Liz Gregory 713-0377
OR Blood Gas Lab 716-2181 (M-F 7am-7pm)
Refer to the i-STAT System Manual and WFBMC specific i-STAT policies
WFBMC POCT web site

## Slide 3: i-STAT Policies/Procedures/Guidelines
Posted on the WFBMC Intranet
Department—Point of Care Testing—Policies
The i-STAT System Manual (Abbott POC website) can be found under Resources-- Resource Links—Manuals

## Slide 4: i-STAT Operator ID
User specific
Serves as identification of testing personnel
Do NOT share operator ID’s
Do NOT enter your operator ID into analyzer and allow use by another individual

## Slide 5: Training and Competency Requirement
Must be documented
Initial training
Initial competency evaluation
6 month competency evaluation-NEW users
On-going Annual competency evaluation—must meet all 6 points of competency assessment, as defined by CLIA
Failure to maintain updated training and competency will result in loss of testing privileges
Competency observation may only be completed by an authorized staff member

## Slide 6: Safety
Always wear gloves when handling analyzer
        Including patient testing and performing QC
Disinfect when contaminated with blood AND between EACH patient / Follow WFBMC policies
Only use auto-disabling single-use finger stick devices
        Never use finger stick device for more than one person
DO NOT lean over the cartridge while filling
A safety shield is recommended
Place gauze or tissue over snap when closing

## Slide 7: i-STAT Components
Analyzer
Cartridge
Simulator (Electronic Quality Control)
Liquid Quality Control Materials
Downloader
Recharger
Printer

## Slide 8: Analyzer

## Slide 9: i-STAT Analyzer (Handheld)
Volt Meter
Testing occurs on the test cartridge bio-sensors
Rechargeable batteries are utilized—care should be taken to keep batteries charged—do not discard rechargeable batteries

## Slide 10: i-STAT Analyzer
Will turn off after 2 minutes of non-use
Can be turned on by pressing the on/off keypad—circle with line
If operator, patient, and cartridge information has been entered, the analyzer will remain on for 15 minutes

## Slide 11: Analyzer Display
‘Cartridge/Simulator Locked’
      Cartridge or simulator locked in analyzer
      DO NOT remove when this message is displayed
Flashing battery icon or ‘Low Battery’ indicates battery voltage is low and batteries need to be re-charged

## Slide 12: Menu Options
Analyzer Status
      Allows viewing of battery voltage - recommend maintaining charge above 8 volts
Data Review
      Allows review/print of data stored in analyzer
Quality Tests
      Quality control checks are performed utilizing this menu
Other Menu options
      Refer to i-STAT System Manual for additional information

## Slide 13: Analyzer Download
Use the Downloader located in each site
The downloader is specifically labeled as the downloader
Analyzers should be downloaded after each patient test so electronic medical record information is current
It is the responsibility of testing personnel to download i-STAT analyzers
You can’t “over-download”

## Slide 14: Analyzer Download
The downloader will display “Communication in Progress”, indicating successful download of data
An error message will display on the analyzer if results do not successfully transmit. If an error occurs:
      Troubleshoot the downloader
      Check for a green power light
      Check all cord/cable connections to the down loader
      Try unplugging and re-plugging the cables
      If troubleshooting does not resolve problem, report problem to the Help Desk

## Slide 15: Re-charger
Can charge battery in compartment and battery in analyzer
Rotate analyzers on chargers to ensure they stay adequately charged

## Slide 16: Printing Results
Printer uses thermal printer
Results fade over time—can’t be used as permanent chartable record
Printer operates using rechargeable batteries—MUST remain plugged into outlet when not in use so batteries can stay charged
Verify patient ID on printout against patient sample ID prior to reporting results

## Slide 17: Cartridges

## Slide 18: i-STAT Supplies
Cartridges and paper can be obtained from the OR Blood Gas Lab M-F 7am-7pm
Care should be taken to get only the amount of cartridges which will be used within the room temperature expiration dating period
Replacement analyzers or printers can be obtained from the Clinical Laboratory Point-of-Care Testing Office

## Slide 19: i-STAT Cartridges
Testing occurs in the cartridge
Cartridges have bio-sensors that “measure” the analytes tested
There are different cartridge configurations available from i-STAT

## Slide 20: i-STAT Cartridges
E=electrolyte
G=gas
C=Chemistry
#=Number of measured parameters
Everything above line is measured
Everything below line is calculated

## Slide 21: Cartridge Storage
Cartridges are stored refrigerated 2 to 8˚C until the manufacturer’s expiration date.
Do NOT use past manufacturer’s expiration date.
Can be stored at room temperature (18 to 30˚C) per manufacturer instructions. G3, CG4, and EG7 cartridges have a 2 month expiration date and Creatinine cartridges have a 14 day expiration date at room temperature.
DO NOT use past the room temperature expiration date!
If a small supply of cartridges are kept refrigerated at the POC site, the refrigerator temperature should have 24/7/365 temperature monitoring. The refrigerator temperature must be maintained at 2 to 8˚C.
If the temperature is ever outside of 2 to 8˚C, contact the Clinical Laboratory Point-of-Care Testing Office for assistance.
Cartridges should NOT be returned to the refrigerator once they have been stored at room temperature.

## Slide 22: Emergency Release of Cartridges
An i-STAT Emergency Cartridge Release Form must be completed should i-STAT cartridges be needed outside of M-F 7am-7pm
The site picking up the cartridges is responsible for room temperature dating of the cartridges
NEVER take cartridges that are marked with “Do Not Use”

## Slide 23: Cartridge Handling - DO NOT’s
DO NOT use “quick heating”—holding close to body or near a warm surface
DO NOT store or expose cartridges to extreme cold or heat such as freezing or stored above lights or computers
Do NOT re-refrigerate cartridges once they have warmed to room temperature
Do NOT touch bio-sensors
Do NOT pre-rupture silver calibrant disc
Do NOT re-use cartridges
Do NOT leave exposed to air and moisture
Do NOT use expired cartridges for patient testing

## Slide 24: Cartridge Handling
Individual cartridges can be used after 5 minutes at room temperature
A box of 25 cartridges must sit at room temperature for 1 hour prior to use
Enter operator and patient ID information prior to collecting the sample
Open by tear symbol
Use cartridge IMMEDIATELY after opening
Fill to blue triangle—leave “blood dome”
Insert into analyzer immediately after filling with sample
Do not move analyzer after cartridge has been inserted

## Slide 25: Quality Controls (QC) and Maintenance
  Simulators
  Liquid QC
  Maintenance

## Slide 26: Simulators
Electronic Quality control that is used to validate the i-STAT analyzer—Pass or Fail
QC results are automatically documented when the i-STAT analyzer is downloaded

## Slide 27: Simulators - Internal
An INTERNAL electronic simulator is performed automatically by the analyzer every 8 hours of use with each cartridge type.
The internal simulator is performed when a patient cartridge is inserted.
Results will not be given if the simulator fails.
If the simulator passes, patient results will be displayed.

## Slide 28: Simulators - External
The EXTERNAL electronic simulator is the same as the internal simulator but is an external device and can be tested upon demand.

## Slide 29: Simulators - External
The external simulator should be used:
        If the analyzer is dropped
        If the internal simulator fails
        If an error code occurs that indicates the simulator should be tested.
        If analyzer performance is in question
        Prior to and after performing the ceramic cartridge cleaning procedure

## Slide 30: Liquid Quality Control
Used to validate performance of test cartridges
Consists of multiple levels—test specific
Should be performed:
      On each new cartridge shipment, per cartridge type, per cartridge lot #, PRIOR to patient use—by Clinical Lab staff
      Monthly—by testing site staff
      If cartridge/analyzer performance is in question

## Slide 31: Liquid QC Important Notes/Reminders
All analyzers will be programmed with electronic QC value assignments for the liquid QC. The analyzer will note whether or not the liquid QC passes or fails. If a value fails, a comment code will be required. However, patient testing lockout does NOT occur.
If liquid QC fails, a comment code is required:
      Use comment code 5—recheck/confirm---repeat liquid QC
      This has been put into place to create a STOP point for the testing staff member
Staff MUST ensure QC results are within acceptable limits PRIOR to proceeding with patient testing.

## Slide 32: Staff should NEVER REPORT PATIENT RESULTS IF QC HAS FAILED!
ALL failed liquid QC results must have corrective/follow-up action, according to your site’s established protocol.
i-STAT will NOT lock out a user from performing patient testing if liquid QC failures occur.
Staff Must ensure liquid QC values are within acceptable range or “PASS”
      If QC fails – Patient testing should STOP!!
      The testing staff member should begin troubleshooting procedures.
Patient results may not be accurate if QC does not perform as expected –
      NEVER report patient results on a failed quality control event.
If a staff member proceeds with patient testing on a failed QC event, then the test site is out of compliance. Patient safety is at risk.

## Slide 33: i-STAT Maintenance
Recharge batteries
Disinfect outside of analyzer between patients
      Do not get moisture inside of the analyzer
Perform ceramic cartridge pin conditioning when indicated by analyzer error codes. (scheduled conditioning is not recommended)
      Each site has a ceramic cartridge and instructions for performing this procedure

## Slide 34: Patient and Sample Identification

## Slide 35: Process Flow
The testing staff member is responsible for ensuring that there is a documented physician order or documented protocol before i-STAT testing is performed.
CSN for that specific date of service or current admission CSN should be entered by testing staff member into i-STAT handheld.
After testing is completed, analyzer should be downloaded. Results flow to p-Web. P-Web searches it’s ADT for a CSN match. If match found, results post to Wake One.
Operator’s name, i-STAT analyzer serial number, and cartridge lot number will report with each test result in Wake One.
Billing occurs when results hit Wake One. Billing occurs based on results that hit Wake One, not on the physician order for the test.

## Slide 36: Patient identity
Follow WFBMC policy for verification of patient identity— Use 2 patient identifiers at the patient bedside and 4 identifiers (name, DOB, MRN, CSN) in sites which perform procedural cases where patient armband is not accessible.
If scanning a patient armband or label not on the patient, include patient ID verification in time out procedure.
Ensure sample identity throughout entire testing and reporting process.
Press and hold the SCAN key pad.
Scan patient ID into the analyzer DIRECTLY from the patient armband on the patient or in the presence of the patient.
Verify patient ID again on analyzer display.
        CAUTION: Laser light—do not stare into beam or point at anyone.

## Slide 37: Patient Identity - Misidentifications
Any misidentified samples/results should be reported to the Clinical Laboratory immediately.
In the event of an error, operators will need to edit their own results in Wake One – for example, sample misidentification or questionable results that were not flagged with a comment code at the time of testing.
It is the responsibility of the staff member who creates the error to notify appropriate patient care providers of the sample misidentification.

## Slide 38: Sample Information

## Slide 39: Sample Requirements
ONLY whole blood should be tested on the i-STAT
Arterial, venous or capillary is acceptable
Syringe or capillary collection containers may be used
Non-anti-coagulated whole blood and capillary samples should be tested within 3 minutes of collection
Lactate samples should be tested immediately after collection
Heparin is the ONLY acceptable anti-coagulant which may be used with the i-STAT – samples should be tested within 10 minutes of collection for blood gas and electrolytes
For Creatinine cartridge:
        MUST use the BD Vacutainer lithium heparin green top 4ml tube— FILL TO CAPACITY or results may be adversely affected
        Heparinized samples may be tested up to 30 minutes after collection

## Slide 40: In-Dwelling Line Sample Collection
  Always adequately ‘clear’ line prior to collecting a sample for testing on i-STAT.
  Inaccurate results will be given if samples are contaminated.
  Back flush line with sufficient amount of blood to remove intravenous solution, heparin, or medications that may contaminate the sample.
  i-STAT Recommendation: five to six times the volume of the catheter, connectors, and needle should be collected as ‘waste’ for i-STAT samples.
  Accurate results depend on an adequate back flush to eliminate the possibility of sample contamination with IV fluids.
  Caution should be taken when collecting from lines which have had fluids that could adhere to the sides of the tubing. These lines may be difficult to adequately ‘clear’.

## Slide 41: Sample Considerations
ALWAYS use a well mixed sample
Hematocrit is most affected by an improperly mixed sample
Mix syringe or capillary samples for 15 seconds
Always squirt out the first drop of blood from syringe or capillary samples
      To check for clots and to get rid of any micro air bubbles
NEVER run a sample that has or has had a clot. Inaccurate results may be obtained!

## Slide 42: Sample Considerations
Any sample collected for pH, pCO2, pO2, iCa, or TCO2 should not be contaminated with air.
Samples should be capped immediately after collection and kept airtight.
If not kept airtight, results will be adversely affected.
Ionized calcium requires balanced heparin.
Ionized calcium results will be adversely affected if a balanced heparin collection container is NOT used

## Slide 43: Information Entered into i-STAT
A temperature corrected blood gas can be obtained
      Make sure any entries made at this prompt are true and correct
      Erroneous temperature entries should be corrected in the electronic medical record
The FIO2 should always be entered
        Can be entered as percent, liters, or decimal entries
The sample type should always be entered
        Arterial, Venous, Capillary, Cord, Mixed

## Slide 44: Results
  Factors Affecting Results
  Result Considerations
  Error Codes
  Comment Codes
  Critical Results
  Enter/Edit of Results – Wake One
  RL6

## Slide 45: Factors Affecting Results
Moving/carrying the analyzer during sample testing
Vibration of the analyzer
Tilting the analyzer during testing. The analyzer MUST remain flat during sample testing.
Refer to the i-STAT System Manual—Cartridge Test Information (CTI) sheets for other factors which can adversely affect results

## Slide 46: Effects of Hemolysis on Potassium Result
Hemolysis is the destruction of red blood cells, caused by disruption of the cell membrane and resulting in the release of hemoglobin.
When the red blood cell membrane is ruptured, potassium is released from the red cell into the blood sample.
To detect hemolysis, the sample must be spun in a centrifuge - the presence of pink or red color in the serum/plasma (top layer) indicates the presence of hemolysis.
Hemolysis adversely affects i-STAT potassium results.
        Potassium value is increased

## Slide 47: Result Considerations
The i-STAT tests whole blood. Hemolysis can NOT be determined.
DO NOT report/treat high potassium results until verified via an alternate method.
Critical values should be repeated and validated. Consider repeat testing via an alternate method.
Use correct critical value documentation of notification. Follow WFBMC policy.
Samples diluted with IV fluids will give inaccurate results
      Drop in hemoglobin and hematocrit
      Falsely elevated sodium and decreased potassium if contaminated with saline
      Blood gas results will be erroneous—low pCO2 may be a possibility
    Results should be carefully evaluated

## Slide 48: Result considerations
Unexpected and unexplained results should be repeated by another test method.
Problems should be reported to:
        Angie Thayer 3-4136
        Ray Dyer 3-4137
        Liz Gregory 3-0377

## Slide 49: When Results Not Given
Cartridge error
      Refer to Technical Bulletin --- ‘Analyzer Coded Messages’
      Do NOT continue with testing until the cause of the error is clearly understood
      If the same error code occurs multiple times with no apparent cause, notify the Clinical Laboratory POCT Coordinator
      Consider using a different analyzer and a different batch of test cartridges
Star out (***)
      Occurs when specific bio-sensor is compromised or there is an interfering substance in the sample
Out of instrument range (>x value or <x value)
      Result is out of reportable range for the analyzer
Suppressed results (<>)
      Displays for calculated parameters, if the calculation is dependent on a result that is out of reportable range.

## Slide 50: Analyzer Error Messages: Codes 20, 23, 43
Be aware that codes 20, 23, 43, and others indicate that the ceramic cartridge conditioning may need to be performed.
Every site should have one of these cartridges.
Instructions are in the i-STAT Technical Bulletin—Analyzer Coded Messages.
This cartridge should NOT be used as preventive maintenance. It should only be used when a quality check code indicates.
Routine use can damage the internal mechanism of the i-STAT handheld.

## Slide 51: Analyzer Error Messages: Codes 79 and 80
Codes 79 and 80 may indicate that the thermal probes are dirty and need cleaning.
Refer to “i-STAT Thermal Probe Cleaning” procedure on the POCT website.
        Go to INTRANET – DEPARTMENTS – POCT –Resources–Resource Links –Internal Resources -“i-STAT TROUBLESHOOTING: ERROR CODES 79 AND 80”

## Slide 52: Use of Comment Codes
    We will use comment codes to better ‘automate’ result hold or posting to the patient record—MUST be entered at the time of testing for effect.
COMMENT CODES --- FOR ISTAT (same as glucose)
    0 translates to --- Procedure Error
        Results flagged with comment code 0 will NOT post to the patient electronic healthcare record.
        This code should be used when the staff member does not believe the result and feels an error was made in testing.
    5 translates to --- Recheck/confirm
        Results flagged with comment code 5 will post to the patient electronic healthcare record but the patient will not be billed for testing.
    6 translates to --- To notify Provider
    123 translates to --- Lab Confirmation to Follow
        Results flagged with comment code 123 will post to the patient electronic healthcare record but the patient will not be billed for testing.

## Slide 53: Star Out Results
To stop a star out result from posting to the patient record, comment code 0 will need to be entered into the analyzer by the testing staff member.
Retest the sample and if star out values continue, send the sample to the Clinical lab for creatinine testing.
If comment code 0 is not entered into the i-STAT handheld, the result will NOT be held in p-Web. It will post automatically to Wake One.
The star out result will report as INSTRUMENT ERROR and an additional NO CHARGE code will be sent to Wake One in the interface message.
There is a Wake One report that will include all results that posted as INSTRUMENT ERROR. For quality monitoring, this will allow us to track the number of star out results that occur.

## Slide 54: Estimated Glomerular Filtration Rate (eGFR)
Estimated Glomerular Filtration Rate (eGFR) values will be reported when Creatinine results are downloaded into the electronic medical record.
The values may be accessed via Wake One.
The i-STAT handheld will NOT report eGFR values.

## Slide 55: Critical Values
Critical values should be evaluated with patient’s clinical symptoms and appropriate repeat testing performed
Any critical value not consistent with patient symptoms should be repeated by an alternate test method
WFBMC policy must be followed for reporting critical values
Follow “Read Back” procedures
The i-STAT policy/procedure has defined i-STAT critical values
The patient physician must be notified of these critical values
Documentation of notification must be included in the patient record--include—result, date and time of notification, name of MD notified, and name of staff member notifying MD

## Slide 56: Result Recall
Turn on analyzer
Select MENU
Select 2 DATA REVIEW
Select desired data
PATIENT DATA requires entry of patient and operator ID information

## Slide 57: Editing of results once they have already posted to Wake One
MUST be done in Wake One
All staff members will have to complete some form of Wake One training for editing interfaced results
This access is needed if a sample misidentification occurs or other result modification is needed

## Slide 58: Correcting a Sample Misidentification
Notify providers of affected patients
In Wake One, open the patient chart---select More Activities
Go into Enter/Edit Results
  —Click Filter
  —Select Point of Care Testing Docked Device
  —Radio Button ALL
  —Enter Date range—Click Accept—
  Highlight the results that need edit--Double click to open edit screen—ENSURE YOU SELECT THE CORRECT TEST RESULTS FOR EDIT!!!--VERY IMPORTANT
At Component Value (the results), replace All i-STAT values with “XXX”. If a comment is present, “XXX” out the comment, as well.
In ONE OF THE COMMENT sections for each set of results, enter this text: INVALID RESULT, PATIENT SAMPLE MISIDENTIFICATION, NOTIFIED enter the full name of the provider that was notified with credentials and the date and time of notification BY the name of the person that notified the provider
Click Accept.
File an RL6 Report. Email RL6 number to the Laboratory Compliance, Quality, and Safety office: athayer@wakehealth.edu, rddyer@wakehealth.edu, engregor@wakehealth.edu

## Slide 59: Correcting erroneous temperature corrected blood gas results in Wake One
Notify provider of affected patient
In Wake One, open the patient chart---select More Activities
Go into Enter/Edit Results
  —Click Filter
  —Select Point of Care Testing Docked Device
  —Radio Button ALL
  —Enter Date range—Click Accept—
  Highlight the results that need edit--Double click to open edit screen—ENSURE YOU SELECT THE CORRECT TEST RESULTS FOR EDIT!!!--VERY IMPORTANT
At Component Value (the results), replace the erroneous patient temperature entered into the i-STAT handheld and the temperature corrected pH, pCO2, and pO2 values with “XXX”.
In the patient temperature COMMENT section, enter this text: INVALID PATIENT TEMPERATURE, NOTIFIED enter the full name of the provider that was notified with credentials and the date and time of notification BY the name of the person that notified the provider
Click Accept.
File an RL6 Report. Email RL6 number to the Laboratory Compliance, Quality, and Safety office: athayer@wakehealth.edu, rddyer@wakehealth.edu, engregor@wakehealth.edu

## Slide 60: Flagging questionable i-STAT results if posted to Wake One
Notify provider of affected patient
In Wake One, open the patient chart---select More Activities
Go into Enter/Edit Results
  —Click Filter
  —Select Point of Care Testing Docked Device
  —Radio Button ALL
  —Enter Date range—Click Accept—
  Highlight the results that need edit--Double click to open edit screen—ENSURE YOU SELECT THE CORRECT TEST RESULTS FOR EDIT!!!--VERY IMPORTANT
At Component Value (the results), replace ALL questionable results with “XXX”.
In the first COMMENT section, enter this text: QUESTIONABLE I-STAT RESULTS, NOTIFIED enter the full name of the provider that was notified with credentials and the date and time of notification BY the name of the person that notified the provider
Click Accept.
File an RL6 Report. Email RL6 number to the Laboratory Compliance, Quality, and Safety office: athayer@wakehealth.edu, rddyer@wakehealth.edu, engregor@wakehealth.edu

## Slide 61: Flagging tests not ordered by the physician
Notify provider of affected patient
In Wake One, open the patient chart---select More Activities
Go into Enter/Edit Results
  —Click Filter
  —Select Point of Care Testing Docked Device
  —Radio Button ALL
  —Enter Date range—Click Accept—
  Highlight the results that need edit--Double click to open edit screen—ENSURE YOU SELECT THE CORRECT TEST RESULTS FOR EDIT!!!--VERY IMPORTANT
DO NOT REMOVE ANY TEST RESULTS
In the first COMMENT section, enter this text: TESTING PERFORMED WITHOUT PHYSICIAN ORDER—CREDIT PENDING, NOTIFIED enter the full name of the provider that was notified with credentials and the date and time of notification BY the name of the person that notified the provider
Click Accept.
File an RL6 Report. Email RL6 number to the Laboratory Compliance, Quality, and Safety office: athayer@wakehealth.edu, rddyer@wakehealth.edu, engregor@wakehealth.edu

## Slide 62: RL6 Report
Allows review of edited i-STAT results
Tracks i-STAT occurrences for quality assurance and quality improvement purposes

## Slide 63: College of American Pathologists (CAP)
If there are concerns regarding result quality and the concerns are not adequately addressed by Lab or WFBMC management, CAP can be notified at 1-866-236-7212.

## Slide 64: Congratulations
This concludes the i-STAT G3, CG4, EG7, and Creatinine education module.
Please note that this power point does not replace the i-STAT policies and procedures.
The presentation only gives highlights of i-STAT information.
Please complete the i-STAT written exam within 1 week of training.

## Slide 65: Hands On Experience
Change batteries and check battery voltage
Test External electronic simulator
Test/Discuss Liquid QC
Test patient sample
Recall results
Discuss Menu Options
Download of results
        istatG3,CG4,EG7,andCreatinineEducationModule122215