- General Procedure Statement:

  - Scope:	       To provide laboratory testing personnel with instructions for

performing laboratory procedures as deemed appropriate by

industry practices and regulatory agencies to assist in quality

patient care.

  - Responsible Department/Party/Parties:

    - Procedure owner: 	   Clinical Core Laboratory Management-Chemistry

    - Procedure:		   Clinical Core Laboratory Personnel

    - Procedure prepared by:    Greg Horton

    - Supervision:		   Clinical Core Laboratory Management-Chemistry

Clinical Core Laboratory Specialist and Designees

Medical Director Clinical Chemistry

- Implementation:  	   Clinical Core Laboratory Management-Chemistry

Clinical Core Laboratory Specialist and Designees

Medical Director Clinical Chemistry

- Definitions:  N/A

- Procedure:

This procedure is valid for the following chemistry analyzers:

| - Siemens Centaur XP |  |
| --- | --- |

# Title: Determining Cardiac Troponin I (TnI-Ultra) in Serum or Plasma using the ADVIA Centaur TnI-Ultra assay on the ADVIA Centaur and ADVIA Centaur XP Systems

# Principle of the Test

The ADVIA Centaur TnI-Ultra assay is a three-site sandwich immunoassay using direct chemiluminometric technology. An ancillary reagent is included to reduce nonspecific binding. The Binary Lite reagent includes a polyclonal goat anti-troponin I antibody labeled with acridinium ester and two biotinylated mouse monoclonal anti-troponin I antibodies. The solid phase reagent is magnetic latex particles conjugated with streptavidin. All reagents are contained within the ReadyPack.

The antibodies in the Binary Lite reagent bind to troponin I in the sample. The biotin contained in the immune complex then binds to the streptavidin-labeled magnetic particles.

The system automatically performs the following actions:

- dispenses 100 µL of sample into a cuvette

- dispenses 100 µL of Binary Lite Reagent plus 50 µL of ancillary reagent and incubates for 2.75 minutes at 37°C

- dispenses 150 µL of Solid Phase Reagent and incubates for 5.0 minutes at 37°C

- separates, aspirates, and washes the cuvettes with Wash 1

- dispenses 300 µL each of Acid Reagent and Base Reagent to initiate the chemiluminescent reaction

- reports results according to the selected option, as described in the system operating instructions or in the online help system

- A direct relationship exists between the amount of troponin I present in the patient sample and the amount of relative light units (RLUs) detected by the system.

# Clinical Application and Usefulness

- For in vitro diagnostic use in the quantitative determination of cardiac troponin I (cTnI) in serum, heparinized plasma, or EDTA plasma using the ADVIA Centaur and ADVIA Centaur XP systems.

- Cardiac troponin I determinations aid in the diagnosis of acute myocardial infarction and in the risk stratification of patients with non-ST segment elevation acute coronary syndromes with respect to relative risk of mortality, myocardial infarction (MI) or increased probability of ischemic events requiring urgent revascularization procedures.

# PREANALYTIC PROCESS

# Specimen Collection and Handling

- Specimen Collection

|  |  | - BIOHAZARD - All products or objects that come in contact with human or animal body fluids should be handled, before and after cleaning, as if capable of transmitting infectious diseases. Wear facial protection, gloves, and protective clothing. |
| --- | --- | --- |

- Serum, heparinized plasma, and EDTA plasma are the recommended sample types for this assay.

NOTE: It is not recommended that either plasma or serum samples from the same patient be used interchangeably with this test.

Evaluation of heparinized and EDTA plasma samples resulted in a mean 1% and 4% decrease in observed values, respectively, compared to serum. A comparison of values for pairs of serum and plasma specimens yielded the following Passing & Bablok equations:

heparinized plasma troponin I = 0.99 (serum troponin I) + 0.03 ng/mL (µg/L); n = 53

EDTA plasma troponin I = 0.96 (serum troponin I) + 0.02 ng/mL (µg/L); n = 31

- Collect all blood samples observing universal precautions for venipuncture.

- Allow samples to clot adequately before centrifugation.

- This assay requires 100 µL of sample for a single determination. Additional volume is required for onboard dilutions.

- Before placing samples on the system, ensure that samples are free of fibrin or other particulate

- matter and that samples are free of bubbles.

Additional instructions for patient preparation as designated by this laboratory: None

Additional specimen type information: Heparinized plasma is used in this laboratory.

Additional handling conditions as designated by this laboratory:

All specimens are labeled with a unique barcode identifier. Specimens received in a standard 13x75 vacutainer tube are processed by the Beckman Power Processor and delivered to attached instruments robotically. Those samples which are of insufficient volume or are of any tube type other than a 13x75 vacutainer are centrifuged and processed manually.

- Specimen Storage and Stability

- Keep tubes stoppered and upright at all times.

- Do not use samples that have been stored at room temperature for longer than 4 hours.

- Tightly cap and refrigerate specimens at 2° to 8°C if the assay is not completed within 4 hours.

- Freeze samples at or below -20°C if the sample is not assayed within 24 hours.

- Freeze samples only once and mix thoroughly after thawing. Frozen specimens can remain frozen up to 1 month in non-frost free freezers.

- Centrifuge frozen samples at 2200 x g for 10 minutes before analysis.

Additional storage conditions as designated by this laboratory: None

- Specimen Rejection Criteria

- See above.

- Specimen Referral Criteria

- Follow manufacturer’s criteria.

# Reagents

- Storage and Stability

- Store the reagents upright at 2 – 8°C.

- Binary Lite Reagent, Solid Phase, and Ancillary Reagent stable until the expiration date on the pack label, or for 28 days onboard the system.

- Calibrators stable until the expiration date on the vial label, or for one day reconstituted at 2 – 8°C, or for 30 days reconstituted at < -20°C, or for 4 hours onboard the system.

- Wash 1 stable until the expiration date on the pack label, or for one month onboard the system.

- CAUTION:

- Discard the primary reagent packs at the end of the onboard stability interval.

- Do not use reagents beyond the expiration date.

- Location of reagents within this laboratory:

- Refrigerator #14	Ambient Storeroom

- Ingredients

- Reagent ingredients for the ADVIA Centaur TnI-Ultra assay are as follows:

| - Reagent | - Volume | - Ingredients |
| --- | --- | --- |
| - Binary Lite Reagent | - 10.0 mL/reagent pack | - polyclonal goat anti-cTnI antibody (0.15 µg/mL) labeled with acridinium ester, 2 monoclonal mouse anti-cTnI antibodies labeled with biotin (2.0 µg/mL), in buffer with stabilizers and preservatives |
| - Solid Phase | - 15.0 mL/reagent pack | - latex magnetic particles in buffer with stabilizers and preservatives |
| - Ancillary Agent | - 5.0 mL/reagent pack | - non-magnetic latex particles in buffer with sodium azide (< 0.1%) |
| - Calibrators | - 2.0 mL/vial | - after reconstitution, low or high levels of bovine cTnI in goat serum with preservatives and stabilizers |
| - Wash 1 | - 1500 mL/pack 2500 mL/pack | - phosphate buffered saline with sodium azide (< 0.1%) and surfactant |

- CAUTION: This device contains material of animal origin and should be handled as a potential carrier and transmitter of disease.

- NOTE: Sodium azide can react with copper and lead plumbing to form explosive metal azides. On disposal, flush reagents with a large volume of water to prevent the buildup of azides, if disposal into a drain is in compliance with federal, state, and local requirements.

# Reagents Special Preparation

- Mix all primary reagent packs by hand before loading them onto the system. Visually inspect the bottom of the reagent pack to ensure that all particles are dispersed and resuspended.

# Calibration

- The ADVIA Centaur TnI-Ultra assay is standardized to an internal standard manufactured using highly purified material. Assigned values for calibrators are traceable to this standardization. The ADVIA Centaur TnI-Ultra assay is traceable to NIST SRM 2921.

- For detailed information about scheduling a calibration, refer to Scheduling Calibrators.(Contained within this procedure).

- Two-point Calibration Interval

- Use Calibrator UL to perform two-point calibrations.

NOTE: The Calibrator UL calibrators provided in this kit are matched to the Solid Phase and Binary Lite Reagent. Do not mix Calibrator UL lots with different lots of Solid Phase and Binary Lite Reagent.

- Perform a two-point calibration every 28 days and:

- when changing lot numbers of primary reagent packs

- when replacing system components

- when quality control results are repeatedly out of range

Defining Calibrator Values for Two-point Calibration

Use the barcode reader to enter the calibrator values from the Calibrator Assigned Value card onto the system.

- At the workspace, select Calibration.

- Select Calibrator Definition.

- Select Scan Data.

- Use the handheld barcode reader to scan the barcodes (from top to bottom) on the Calibrator Assigned Value card.

- Ensure that the calibrator values are correct. After you select Save, you cannot add or delete a test from a calibrator definition.

- Select Save.

Preparing Calibrators

- Add 2.0 mL of reagent water into each calibrator vial using a volumetric or precision pipet.

- Let the calibrators stand undisturbed for 15 to 20 minutes at room temperature (20 to 30°C) to allow the lyophilized material to dissolve.

- Gently swirl and invert the vials until homogeneous.

- The calibrator can be aliquoted into sample cups that are sealed tightly. Freeze for up to 30 days and thaw only once. To ensure complete homogeneity, thawed calibrators must be gently mixed and inverted.

Defining the Master Curve

The ADVIA Centaur TnI-Ultra assay requires a Master Curve calibration when using a new lot number of Binary Lite Reagent and Solid Phase. For each new lot number of Binary Lite Reagent and Solid Phase, use the barcode reader or keyboard to enter the Master Curve values on the system. The Master Curve card contains the Master Curve values. For detailed information about entering calibration values, refer to the system operating instructions or to the online help system.

- At the workspace, select Calibration.

- Select Master Curve Definition.

- Select Scan Data.

- Use the handheld barcode reader to scan the barcodes on the Master Curve card.

- IMPORTANT: Ensure you are scanning the side of the Master Curve card labeled ADVIA Centaur.

- NOTE: The table on the Master Curve card represents the fields as they are displayed in the ADVIA Centaur System only. Use the illustrations in the Using the Master Curve Card and Calibration – Master Curve Definition Window Customer Bulletin, part number 078D0560-01 Rev. A, to match the fields in the Master Curve card with the fields as they are displayed in the ADVIA Centaur XP System.

- Select Save.

Quality Control (QC)

Follow government regulations or accreditation requirements for quality control frequency.

For detailed QC procedural information, refer to the system operating instructions or to the online help system.

QC Materials

Siemens Healthcare Diagnostics recommends the use of commercially available quality control materials with at least 2 levels (low and high). A satisfactory level of performance is achieved when the analyte values obtained are within the Acceptable Control Range for the system or within your range, as determined by an appropriate internal laboratory quality control scheme.

To monitor system performance and chart trends, as a minimum requirement, two levels of quality control material should be assayed on each day that samples are analyzed. Quality control samples should also be assayed when performing a two-point calibration. Treat all quality control samples the same as patient samples.

Quality Control information specific to this laboratory:

BIORAD Cardiac Markers Plus Control 1&3, Reference 181,183.  LIS codes: MBCT1,MBCT3

BIORAD Liquichek Cardiac Markers Plus Control LT, Level 1B, Reference 27105. LIS code: CRL

Storage: -20⁰C to -70⁰C until expiration date on material. Once thawed, 20 days when stored tightly capped at 2⁰C to 8⁰C.

Location of controls in this laboratory:

Frozen: Revco freezer		Refrigerated: Refrigerator #4

QC Frequency

Analyze all levels of quality control material each time a two-point calibration is performed.

Analyze all levels of quality control material on each day that samples are analyzed for each reagent pack on the instrument (the one currently in use and the back-up).

To order QC on multiple reagent packs on the Centaur:

Worklist>>>Schedule>>Control>>>TnIUltra>>>choose MBCT1, MBCT3, CRL>>>Reagent options

>>>choose reagent lot (if more than one onboard)>>>use the dropdown for Pack ID>>> Individual

packs have IDs embedded in a portion of their barcode>>>choose the 1st one>>>repeat the process

and choose the second one (if more than one onboard).

Troubleshooting Out-of-Range QC Values

If the quality control results do not fall within the Expected Values or within the laboratory’s established values, do not report results. Take the following actions:

- Verify that the materials are not expired.

- Verify that required maintenance was performed.

- Verify that the assay was performed according to the instructions for use.

- Rerun the assay with fresh quality control samples.

- If necessary, contact your local technical support provider or distributor for assistance.

- Quality Control Monitoring / Failure Resolution within this Laboratory:

- Quality Control is verified using Westgard Rules programmed into the LIS system, Beaker. Corrective action to resolve control failures includes repeat analysis, recalibration, new reagent and/or control material, lot change discrepancies and if necessary field service. Once quality control is deemed acceptable, specimen processing may proceed.

ANALYTIC PROCESS

# Instrument Operation and System Description

The ADVIA Centaur and ADVIA Centaur XP systems are automated, direct chemiluminescent immunoassay analyzers that offer optimal productivity and efficiency. No-pause reloading of reagents, samples, and supplies means that the system is always ready to process samples.

Sample racks are loaded in the sample entry queue and the sample start button is pressed to activate the test sequence. The sample entry queue moves the sample racks to the inprocess queue, where the sample is aspirated and dispensed it into a cuvette in the incubation ring.

Reagents are dispensed into the cuvette, the reaction mixture is incubated, and then the cuvette is moved to the wash station where the magnetic particles are washed. Acid Reagent is dispensed into the cuvette and then the cuvette is moved into the luminometer. The addition of Base Reagent causes the chemiluminescent reaction to occur. The PMT measures the chemical light reaction that takes place.

A. Processing Start and Fast Stop

If the sample entry queue is in standby, press the sample start button to activate the sample entry queue. The sample entry queue moves the sample racks one at a time to the inprocess queue.

There are two fast stop buttons on the system. Use these buttons if there is an obstruction that could cause serious system or physical damage.

B. Start-up

- Start the system (if required) and sign in.

- Load required reagents and consumables.

- Schedule (if needed) and load samples and controls.

- Press the sample start button.

C. Managing Supplies

While the system is running, you can manage the following supplies without interrupting the run.

| - Primary reagents | - Tips |
| --- | --- |
| - Ancillary reagents | - Cuvette waste |
| - System fluids | - Tip tray waste |
| - Water | - Tip waste |
| - Cuvettes | - Liquid waste |

D. Scheduling a Run/Entering a Worklist

You can schedule a run through a laboratory information system (LIS) or you can manually create a worklist. For detailed information about scheduling a run, refer to the Operator's Guide.

LIS Supplied Worklist

The system can obtain worklist entries from an LIS by manually requesting the worklist (Receive Requests from LIS) or by allowing the system to automatically request worklist entries when a sample is added to the inprocess queue (System Automatically Queries Host for Worklist).

Manual (Operator-Initiated) Worklist

Scheduling Tests for an Individual Patient Specimen

- At the workspace, select Worklist.

- Select Schedule.

- Select Patient.

- Select Schedule by SID or Schedule by Rack ID to indicate how you want the system to identify the sample.

- Enter the SID or the Rack ID and then press Enter. Alternatively, select Scan Data to scan the SID using the hand-held barcode reader.

- If you want to process the sample before routine samples, select Stat.

- Select or enter the tests or the profiles for the sample.

- Optionally, select Demographics, enter the demographic information, and select Continue.

- Select Save.

Scheduling Batches

Groups of samples with the same tests can be scheduled as a batch. The system performs the requested tests on all samples loaded in the starting rack, ending rack, and all racks in between.

- At the workspace, select Worklist.

- Select Schedule.

- Select Batch.

- Enter the rack number of the starting rack in Starting Rack.

- Enter the rack number of the ending rack in Ending Rack.

- Select or enter the tests or profiles.

- Select Save.

Scheduling Calibrators

Calibrate an assay in the following instances:

- the calibration interval expires (use the Calibration – Summary window to view calibration status)

- a calibration is invalid

- a new lot of assay reagents is used

- controls are repeatedly out of range

- Enter the calibrator information provided on the ADVIA Centaur Calibrator Assigned Value card.

- Enter the Master Curve information from the Master Curve card provided with the primary reagent packs.

- Load the reagent lot you want to calibrate.

- At the workspace, select Worklist.

- Select Schedule.

- Select Calibrator.

- Select Schedule by SID or Schedule by Rack ID to indicate how you want the system to identify the calibrator.

- Select or enter a test. The calibrators previously defined for the test are automatically displayed.

- NOTE: Calibrators do not display in the following situations:

- •	there are no calibrators defined

- •	there are no Master Curves defined

- •	there is no reagent onboard

- •	the calibration material for defined lots is expired

- Select a calibrator.

- Select a reagent lot.

- If you schedule by rack, enter the Rack ID in Rack.

- Select Save.

- Repeat steps 7 through 12 to schedule additional tests for a calibrator.

- Label two sample cups with calibrator barcode labels: one for the low and another for the high.

- Gently mix the calibrators and dispense at least 1.0 mL of each calibrator into the appropriate sample cups.

NOTE: Sample cups are marked at 1.0 mL intervals to assist in filling. The volume depends on how many assays are being calibrated using these calibrators and the number of calibrator replicates.

- Load the sample cups in a rack. See Loading Sample Racks.

NOTE: The low calibrator cup position must precede the high calibrator cup position.

Calibration failure resolution within this laboratory:

Corrective action to resolve calibration failures includes  recalibration, new reagent and/or control/calibration material, lot change discrepancy investigation, and if necessary, field service.

E. Loading Sample Racks

NOTE: You can also load samples using your laboratory automation system (LAS).

|  |  | BIOHAZARD All products or objects that come in contact with human or animal body fluids should be handled, before and after cleaning, as if capable of transmitting infectious diseases. Wear facial protection, gloves, and protective clothing. |
| --- | --- | --- |

- If you use sample cups, label them with the appropriate barcode labels and dispense the sample into the labeled sample cups.

- Move the tube-type selector to the position for the tube type you are loading in the rack or for multiple tube types.

- NOTE: When you use an ADVIA Centaur rack, you can only load one tube type in the rack.

- Place the sample tubes or cups in the rack.

- CAUTION: Do not load capped sample tubes. Ensure that all caps are removed from the sample tubes before loading the rack onto the system. Leaving the caps on the tubes can cause mechanical damage to the system.

- Ensure that the barcode labels are clearly visible above or between the slots in the rack.

- Load the rack in the sample entry queue. Ensure that the notch in the rack rests on the raised area of the sample entry queue.

CAUTION: Do not manually push a rack into the inprocess queue and do not manually place sample racks near the entrance to the inprocess queue without using the raised area of the sample entry queue as a positioning guide.

CAUTION: Do not press the sample start button until you close the primary reagent door and ensure that all reagent barcodes were read correctly.

F. Loading STAT Samples

- STAT samples can be placed in a rack and loaded in the Stat entry. The system assigns STAT priority to these samples and moves the rack to the inprocess queue as soon as space is available.

- Samples can also be scheduled as STAT samples at the LIS or at the Worklist – Schedule window and loaded in the sample entry queue. STAT samples loaded in the sample entry queue are processed as soon as the system scans the barcode label and loads the sample rack into the inprocess queue.

G. Loading Reagents

You can monitor the volumes of primary and ancillary reagents using the Worklist – Reagent Inventory window.

CAUTION: Mix all reagent packs by hand before loading them onto the system. It is important to minimize foaming. For detailed information about preparing reagents for use, refer to the Operator's Guide.

Loading Primary Reagents

- Mix the primary reagent pack by hand. Visually inspect the bottom of the reagent pack to ensure that all particles are dispersed and re-suspended.

- Open the primary reagent door.

- Identify a location for the pack by using the arrows as a placement guide.

- Pull the reagent holder towards you.

- Load the pack on the reagent holder.

- Press down on the top of the pack to ensure it is correctly seated on the reagent holder and then push the reagent holder back firmly until you hear the pack lock in place.

- Close the primary reagent door.

- Press the sample start button for the system barcode scanner to rescan the reagent barcodes and acknowledge any reagent changes.

Loading Ancillary Reagents

- Mix the ancillary reagent pack by hand.

- Load the pack in the ancillary reagent entry with the barcoded side facing outward.

An LED next to the ancillary reagent entry turns green to indicate that the pack is loading.

NOTE: Load ADVIA Centaur Multi-Diluent 11 in the ancillary reagent area if automatic dilution of a sample is required.

# POSTANALYTIC PROCESS

# Reporting Results

As with all in vitro diagnostic assays, each laboratory should determine its own reference range(s) for the diagnostic evaluation of patient results.

Reference Interval used by this laboratory:

[0.000-0.040] NG/ML           >0.040 INDICATES MYOCARDIAL DAMAGE.  THE DIAGNOSIS OF MYOCARDIAL INFARCTION REQUIRES CLINICAL CORRELATION.COMMENT:  Elevated troponin may also be due to pulmonary emboli, aortic dissection, heart failure, renal failure, inflammation, trauma, toxins, etc.  or demand ischemia in the setting

Expected Results

A reference range study was conducted using the ADVIA Centaur TnI-Ultra assay based on guidance from the Clinical and Laboratory Standards Institute (CLSI) Protocol C28-A2. The study, which used 1845 fresh serum, lithium heparin plasma, and EDTA plasma samples from 648 apparently healthy individuals ranging from 17–91 years of age, demonstrated a 99th percentile of 0.04 ng/mL (μg/L). The potential range of results for the 99th percentile is 0.02 to 0.06 ng/mL (μg/L) for the ADVIA Centaur family of systems, dependent upon sample type, instrument, and reagent lot.

The precision curve for the ADVIA Centaur TnI-Ultra assay demonstrates a total coefficient of variation of 10% at a level of 0.03 ng/mL (μg/L). This level of imprecision shows a CV of ≤ 10% at the 99th percentile of the reference population indicating that the TnI-Ultra assay meets the definition of a high sensitivity Troponin method as described by the ESC/ACC recommendations.

Critical Values

Presently, troponin I is considered a critical value for the Emergency Department.

Reporting Protocol for Critical Values

Critical results are called to the charge nurse or physician and documented in the Beaker software.

Addition reporting information as designated by this laboratory:

Results within the analytical range are auto-validated by Beaker software whereas flagged results or those falling above the analytical range are not processed until action is taken by the technologist. This includes repeat runs, dilutions, etc.  Automatic printing of results is not routinely enabled since records are maintained electronically.

Any abnormal result generated for a patient in the Emergency Department is to be called and documented in Beaker. A pop-up notification stating “ELEVATED RESULT MUST BE CALLED TO EMERGENCY DEPT” will be displayed in Beaker. Call and document in the comment field. Type “.ctrb” to append the comment “Called to and readback by” adding name of person contacted, time AND date.

Units for Reporting Results

The system reports serum troponin I results in ng/mL (common units) or µg/L (SI units), depending on the units defined when setting up the assay. The conversion formula is 1.0 ng/mL = 1.0 µg/L. You can define the units and the number of decimal places for test results using the Worklist – Test Summary window.

Acceptable Results

Patient test results are acceptable and may be reported when:

- Calibration criteria has been satisfied.

- Quality control is within defined ranges

- Corrective Action

- Patient test results must be repeated and corrective action taken when:

- Quality control falls outside defined ranges.

- Upon discovery of potential instrument malfunction which may have affected patient results

# Procedure Notes

- Calculations

- For detailed information about how the system calculates results, refer to the system operating instructions or to the online help system.

- High Dose Hook Effect

Patient samples with high troponin I levels can cause a paradoxical decrease in the RLUs (high dose hook effect). In this assay, patient samples with troponin I levels as high as 1000 ng/mL (µg/L) will assay greater than 50 ng/mL (µg/L).

- Disposal

- Dispose of hazardous or biologically contaminated materials according to the practices of your institution. Discard all materials in a safe and acceptable manner, and in compliance with all federal, state, and local requirements.

- Method Limitations

- Reportable Range

- The reportable range of the ADVIA Centaur TnI-Ultra assay is 0.006 – 50 ng/mL (µg/L).

- Dilutions

- The assay range is 0.006 to 50 ng/mL (µg/L). Patient samples with troponin I levels greater than 50 ng/mL (µg/L) must be diluted and retested to obtain quantitative results.

- Patient samples can be automatically diluted by the system or prepared manually.

- For automatic dilutions, ensure that ADVIA Centaur Multi-Diluent 11 is loaded and set the system parameters as follows:

- Dilution point: ≤ 50 ng/mL (µg/L)

- Dilution factor: 2, 5, 10

- For detailed information about automatic dilutions, refer to the system operating instructions or to the online help system.

- Manually dilute the patient samples when patient results exceed the linearity of the assay using automatic dilution, or when laboratory protocol requires manual dilution.

- Use Multi-Diluent 11 to manually dilute patient samples, and then load the diluted sample in the sample rack, replacing the undiluted sample.

- Ensure that results are mathematically corrected for dilution. If a dilution factor is entered when scheduling the test, the system automatically calculates the result.

- Other Limitations

- Do not pour the calibrators back into the vials after calibration because evaporation could occur, which may affect performance.

- Dispose of any calibrator remaining in the sample cups after 4 hours.

- Do not refill calibrator sample cups when the contents are depleted. If required, dispense fresh calibrators.

- Heterophilic antibodies in human serum can react with reagent immunoglobulins, interfering with in vitro immunoassays. Patients routinely exposed to animals or to animal serum products can be prone to this interference and anomalous values may be observed. Additional information may be required for diagnosis.

- Whenever Allergy Screen or sIgE reagents are used on the same system as ADVIA Centaur TnI-Ultra reagents, ensure that sufficient Probe Wash 4 is available to eliminate any potential interference.

- Samples from patients receiving preparations of mouse monoclonal antibodies for therapy or diagnosis may contain Human Anti-Mouse Antibodies (HAMA). Such samples may show either falsely elevated or falsely depressed values when tested with this method.

- The following endogenous substances were added to human-based samples at the concentrations listed and evaluated for potential interference in the ADVIA Centaur TnI-Ultra assay. The results demonstrate a ≤ 10% interference from each substance:

| - Specimens that are . . . | - Demonstrate ≤ 10% change in results up to . . . |
| --- | --- |
| - hemolyzed | - 500 mg/dL of hemoglobin |
| - lipemic | - 1000 mg/dL of triglycerides |
| - icterus | - 20 mg/dL of conjugated bilirubin |
| - icterus | - 20 mg/dL of unconjugated bilirubin |
| - human IgG | - 12 g/dL of protein |
| - biotin | - 10 ng/mL of biotin |

- The biotin level in a normal population is approximately 0.3 to 1.0 ng/mL (1.2 to 4.3 nmol/L).25 Biotin could be higher if a biotin supplement is used, particularly in dialysis patients taking multiple doses of vitamins. In samples with biotin levels greater than 10 ng/mL (41 nmol/L), the troponin-I assay may demonstrate a 10% decrease in results.

- For additional information on performance characteristics, see the product information in the ADVIA Centaur TnI-Ultra product insert.

# Equipment and Supplies

- ADVIA Centaur TnI-Ultra ReadyPack

- ADVIA Centaur Wash 1

- ADVIA Centaur Multi-Diluent 11 (optional)

- ADVIA Centaur Acid Reagent (0.5% H202, 0.1N HNO3)

- ADVIA Centaur Base Reagent (0.25N NaOH and surfactant)

- ADVIA Centaur Cleaning Solution Concentrate (~52.5 g/L sodium hypochlorite)

- ADVIA Centaur Sample Cups and Caps

- ADVIA Centaur Cuvettes

- ADVIA Centaur Tips

- Reagent Water

- Review/Revision/Implementation:

  - Review Cycle: 2 years

  - Office of Record:	  Department of Clinical Core Laboratory-Chemistry

  - All new procedures and procedures that have major revisions must be signed by the Department Chairman.

  - All reviewed procedures and procedures with minor revisions can be signed by the designated section medical director

- Related Procedures:

- References, National Professional Organizations, etc.:

- Siemens Healthcare Diagnostics ADVIA Centaur TnI-Ultra Product Insert.

- Siemens Healthcare Diagnostics ADVIA Centaur Reference Manual.

- Siemens Healthcare Diagnostics ADVIA Centaur XP Operator's Guide.

- Clinical and Laboratory Standards Institute (CLSI). Clinical Laboratory Technical Procedure Manuals; Approved Guideline, GP02-A5, 2006.

# Technical Assistance

Siemens Healthcare Diagnostics Technical Care Center: 1-877-229-3711

Customer Service: 1-800-255-3232

Serial Number:

Customer Account Number:

- Attachments:

- Revised/Reviewed Dates and Signatures:

| Review/Revision Date | Review/Revision Description | Signature |
| --- | --- | --- |
| 06/17/219 | Added further QC instructions (pg.8) |  |