## Slide 1: Nova Stat Sensor Creatinine Testing
HPRLABSSC20

## Slide 2
How to Use this Course
Course navigation buttons are always present at the bottom of the screen unless designated on the screen.
Please follow navigation instructions on each page or click on the topics in the navigation menu to the left of the screen to move backward and forward. Where noted, click the forward button to move to the next screen, or back to go back to the previous screen.
This course has a quiz. You cannot use the forward or back buttons while you are in the quiz. Follow the instructions in the quiz to select your answer, and then click the Submit button.
Notes: For use with Articulate 2013
This explains how learner will navigate through the course and explains where and how buttons work in the module.

## Slide 3: Objectives
List the steps in the NOVA StatSensor procedure
Discuss test procedure documentation and reporting
Explain quality control measures

## Slide 4: Background
The Creatinine StatSensor is used for determining a patient’s creatinine/eGFR prior to using contrast.
Creatinine with eGFR is a more accurate and sensitive assessment of kidney function than creatinine alone.
StatSensor Creatinine/eGFR is used for capillary samples only

## Slide 5: Equipment and Supplies
Test Strips:
expire 90 days after opening, not to exceed the manufacturer’s expiration date
Test Strips:
are affected by heat, light, and humidity. Keep vial closed
Test Strips:
do not transfer from one bottle to another
Always use disposable lancet
Control Solutions:
expire 90 days after opening, not to exceed the manufacturer’s expiration date
All strips and control solutions must have open date and new expiration date on label – cover with scotch tape to keep from wiping off

## Slide 6: Quality Control
If quality controls are not checked every 24 hours of patient testing the meter will lock out
Two levels of quality control must be performed every 24 hours - QC lot numbers are scanned in to the StatSensor at the time of testing

## Slide 7: Quality Control Continued
A vial of test strips was left open or been exposed to extreme heat, humidity, or cold
Meter is dropped, or when troubleshooting the meter
Quality Controls must also be perfomed in the following situations:

## Slide 8: Patient Testing Procedure
Gloves are donned & patient identified
Tap the screen of the StatSensor & press OK/Login
Scan employee ID barcode on back of name badge (preferred) or manually enter employee ID & press accept;
Select eGFR from the Patient Test screen and press accept - results will NOT cross to the chart without eGFR
Notes: Blank template

## Slide 9: Patient Testing Procedure
From the “Enter Patient ID” screen, scan the patient arm band (preferred) or manually enter the CSN number, press accept
From the “Enter Strip Lot” screen, scan the barcode on the creatinine strip vial, press accept
If eGFR is selected, enter patient age/sex/race, press accept after each entry;
Select “Capillary” when prompted;
Insert the test strip into the meter
Notes: Blank template

## Slide 10: Specimen Collection
Cleanse the finger thoroughly with alcohol and allow to dry; perform a fingertip puncture using a stab technique
When the second blood drop appears, keep the meter horizontal and touch the end of the test strip to the blood until the well is full & the meter beeps
The “Apply Sample” screen should be displayed on the meter before proceeding
Wipe away the first drop of blood with a clean gauze pad
Notes: Blank template

## Slide 11: Specimen Collection
The test result will appear in 30 seconds;
Remove test strip and discard; If testing fails for any reason, repeat using a new test strip
DO NOT hold the meter upright when applying blood to the test strip. Doing so will allow blood to run down into the meter and damage it.
To accept result, press accept, then place the meter on the docking station to transmit results to the chart.
Notes: Blank template

## Slide 12: Patient Reference Ranges
0.60 – 1.10 mg/dl
  Results below 0.3 mg/dl will display “LO”
  Results above 12.0 mg/dl will display “HIGH”

## Slide 13: Cleaning Procedure
Meter must be cleaned between patients and prior to docking
When using disinfectant wipes, squeeze out excess solution prior to using
Do not expose the meter’s base connector and docking station to disinfectant wipes
After disinfection, use clean damp cloth or paper towel to wipe cleaner residue from scanner window and touch screen
Use appropriate disinfectant and adhere to required contact time– 3 minutes for alcohol free wipes, 4 minutes for bleach wipes

## Slide 14: Flow Errors
Flow errors can be caused by the following and require repeat with a new test strip
  Insufficient sample
  Sample applied incorrectly
  Removing strip from blood droplet before countdown is displayed on the screen

## Slide 15: Docking-Charging Station
When the meter is not in use, place it into the docking station to:
  Update patients/operators/lot numbers in the device
  Send patient results to the chart
  Recharge battery for the device
The meter must be docked a minimum of every 8 hours
Failure to dock meter will result in meter lock out

## Slide 16: Docking Station Indicator Lights
The left light is green if the station is connected to the network
The middle light flashes green if data is transferring
The right light is green when the battery is fully charged or amber when charging

## Slide 17: Points to Remember
Hold meter 4 to 5 inches away from barcode when scanning
Do NOT hold meter upright when blood sample or QC sample is applied to strip
Always pay attention to blue dialogue box at top of each screen to ensure you are on correct screen
Do NOT use the patient medical record number as the patient ID-Results will not cross into the medical record

## Slide 18: Thank you for taking time to complete “NOVA Stat Sensor”. This information can help us better care for our patients and their families. For more information about this course, e-mail Kelly.Leiner@wakehealth.edu