|  |  | COBE Quality Control BB.QC.1008.5 | Dept: | 324311 |
| --- | --- | --- | --- | --- |
|  |  |  | Dept Name | Blood Bank |
|  |  |  | Effective Date: | 3/6/2002 |
|  |  |  | Revised Date: | 4/16/2020 |
| Name & Title: CLIA Laboratory Medical Director |  |  | Contact: | Julie H. Simmons/ Christina S Warren |
| Signature: |  |  | Date: |  |

1.  General Procedure Statement:

- Purpose: To confirm that the Cobes are functioning properly when utilizing the washing and delycerolization protocols. Annually, each Cobe is checked to ensure at least 80% red blood cell recovery following the deglycerolization and washing procedures. Periodic checks are done on the hydraulic system.

B. Responsible Department/Scope:

i  Procedure owner/Implementer:  Julie H. Simmons/Christina S. Warren

ii. Procedure prepared by:  Julie Simmons

iii. Who performs procedure:     Department staff/management

C. Definitions:

Check Prime:    Procedure to confirm the hydraulic system has been properly primed before use.

%Recovery:      AABB requires 80% recovery of red blood cells following deglycerolization.

Priming Cushion:  Donut shaped foam kept in Cobe during prime and when Cobe not in use.

Degly:   Deglycerolized

Hue:     Compared color to the Haemonetics color comparator located on the Cobe.  Hue

corresponds to incremental levels of hgb.

CALL Stop:   Places Cobe in a hold status.

PC Stop:  Ends the Cobe program at the end of the supernatant out operation of the cycle

RC Stop: Ends the program at the end of the Agitate/wash-in operation

D. Sections:

I. Hydraulic System, Check Prime

II.  General Maintenance of Cobes

III. Percent Recovery for Deglycerolized (Degly) and Washed Red Blood Cells

IV. Verify Programs on Cobe 3, Cobe 5 and Cobe 6 Monthly.

E. Protocol:

- BioMed performs the quarterly timer checks and RPMs for the Cobes.

- Annual preventive maintenance is performed by BioMed and includes checking of the red cell detector.

- The hydraulic system is checked at periodic intervals (hydraulic prime for Cobe BMT and super prime for Cobe 3, 5 and 6) to ensure that the hydraulic system is properly primed prior to use by Blood Bank staff.  Refer to 4 and 5 below for when to check.

- The hydraulic system prime (hydraulic prime for Cobe BMT/ Super prime for Cobe 3, 5 and 6) is performed by Blood Bank and/or BioMed:

- if air is observed in the hydraulic system

- if the instrument has not been used in the last 48 hours. Refer to Cobe/IBM Component Prep worksheet to determine when last used.

- monthly for preventative maintenance for Cobe BMT. Record on the Weekly and Monthly Component Prep Checklist.

- The “Check Prime” is performed each shift before use and after hydraulic system prime by Blood Bank staff and recorded on the Cobe/IBM Component Prep worksheet:

- Weekly, the surface, rims, bowls, covers, red cell detectors and vertical posts of each Cobe are cleaned. This is documented on the Daily/Weekly/Monthly Component Prep Checklist.

- Monthly, the overflow bottle of each Cobe is checked. This is documented on the Daily/Weekly/Monthly Component Prep Checklist.

- Each Cobe is checked annually by Blood Bank for ≥80% recovery of red blood cells following the deglycerolization  and the washing procedure. This is documented on the Percent Recovery-Washed or Deglyced Unit Calculation From.  Refer to Attachment 3.

- Each unit that is deglycerolized is checked for hue of last Super Out.

- Acceptable hue is between 1 and 4 which correlates to ≤150 mg/dl of free hemoglobin.

- Unacceptable hue is greater than 4 and additional washes are required.

Refer to: CP: Deglycerolization

- Each deglycerolized red cell is checked for fragility and indirectly for glycerol removal by placing the deglyced red blood cells from a segment in 0.7% saline.

- Acceptable hue is 1 to 5 which correlates to ≤ 300mg/dl of free hemoglobin.

- Hue greater than 5 is unacceptable and credit should be requested from supplier.

- A segment is retained for a final inspection for hue at issue for both deglyced and washed red blood cells.

- Acceptable hue is 1 to 4 which correlates to ≤ 150 mg/dl of free hemoglobin.

- Unacceptable hue is greater than 4 and unit cannot be issued.

Refer to: CP: Deglycerolization/Washing Red Cells

2. Procedure:  I:  Hydraulic System, Check Prime

Chemical Risk Assessment:  Low

Biological Risk Assessment: Low

Protective Equipment:  Lab coat, gloves

Reagents: NA

Supplies: Cushion and disc

Equipment: Cobe

Specimen Requirements:  NA

- Check Prime

| STEPS | INSTRUCTIONS | CHANGE/ APPROVAL |
| --- | --- | --- |
| 1.0 | Insert priming cushion and disc if not already in place. - Priming cushion and disc are placed on top of COBE while in use. They should be kept inside centrifuge bowl when not in use. |  |
| 2.0 | Replace centrifuge bowl cover and alignment blocks. - COBE #3, COBE #5 and COBE #6 – close centrifuge latch |  |
| 3.0 | Turn power switch on - IBM COBE BMT: set controls for washing RBC Refer to Section A, Step 3 for diode pin settings. - COBE #3. COBE #5 and COBE #6: set program to 0 and dial controls for washing RBC NOTE: Super Out rate mL/min must be set to 450. |  |
| 4.0 | Press START/SPIN and allow to spin for 15 seconds. |  |
| 5.0 | Press SUPER OUT. 5.1 The audible alarm should sound and the EXCESS PRESSURE light should come on within 15 seconds. 5.2 If no alarm, repeat the hydraulic system prime in Section A and try again. 5.3 If repeat prime fails, take Cobe out of service and place a service request to BioMed. |  |
| 6.0 | Press STOP/RESET. |  |
| 7.0 | Wait for READY light to come on before processing. |  |
| 8.0 | Record check on COBE/IBM Component Preparation Worksheet. Refer to Attachment 2: COBE/IBM Component Preparation Worksheet |  |

2. Procedure: II.  General Maintenance of Cobes

Chemical Risk Assessment:  Low

Biological Risk Assessment: Low

Protective Equipment:  Lab coat, gloves

Reagents: NA

Supplies: Caviwipes, Alcohol pads, Cotton swabs

Equipment: Cobe

Specimen Requirements:  NA

| STEPS | INSTRUCTIONS | CHANGE/ APPROVAL |
| --- | --- | --- |
| 1.0 | Clean surfaces, rims of bowls and covers of each Cobe weekly. 1.1 Wipe thoroughly with a Cavi-Wipe. |  |
| 2.0 | Clean red cell detectors weekly. 2.1 Take an alcohol pad and wrap around cotton swab. 2.2 Insert between red cell detectors and wipe clean. |  |
| 3.0 | Clean vertical posts weekly. 3.1 Wipe clean with an alcohol pad. |  |
| 4.0 | Check reservoir bottle for liquid monthly. 4.1 Open the right panel door of the Cobe. 4.2 Locate the overflow bottle. 4.3 Check visually for the appearance of any liquid in the bottle and initial Weekly and Monthly Component Prep Checklist. Refer to Attachment 1: Weekly and Monthly CP Checklist 4.4 Check for:; Steps; Comments. Liquid Present; - Notify BioMed - Put sign on unit that it is out of use. - Record in maintenance book. - Notify management; - Complete Maintenance log when repaired by BioMed. - Notify management.. No Liquid Present; - Initial Checklist; - No additional actions. |  |

2. Procedure: III.  COBE RBC % RECOVERY FOR WASHING AND DEGLYCEROLIZED (DEGLY) RED CELLS

Chemical Risk Assessment:  Low

Biological Risk Assessment: Low

Protective Equipment:  Lab coat, gloves

Reagents: NA

Supplies: Sample site coupler, 12x75 mm plastic tube, 1 ml syringe, 18 g needle, blood bag, transfer bag, alcohol pad, cotton

swab

Equipment: Cobe

Specimen Requirements:  NA

Frequency: Deglycerolization/Washing:  Annually EACH Cobe is checked for RBC Percent Recovery

| STEPS | INSTRUCTIONS | CHANGE/ APPROVAL |
| --- | --- | --- |
| 1.0 | Select choice of units for procedure:: - WASH – choose packed RBCs from QC shelf with final status of DS (Discarded / Incinerated) that has NOT been opened or a Group O RBC that can be transfused. - DEGLY – any frozen unit that is being deglyced for patient transfusion may be used. |  |
| 2.0 | Open form, Cobe Percent Recovery QC 2.1 Location: G:\Lab_Shared\BloodBankStaff\BB Staff Information\3 – Forms 2.2 Attachment 3: (BB.form.1000) Cobe Percent Recovery QC |  |
| 3.0 | Enter the below on the Cobe Percent Recovery QC form: 3.1 Enter date 3.2 Enter tech initials 3.3 Enter Cobe ID 3.4 Enter Unit Number |  |
| 4.0 | Enter PRE-WEIGHT information: 4.1 If; Do (Bag with blood. Empty bag weighed later.). WASH; Weigh (gms) unit of RBCs on component scale and enter results. DEGLY; Weigh (gms) unit of thawed RBCs on component scale and enter results. |  |
| 5.0 | Mix unit to be washed well for at least 3 minutes. |  |
| 6.0 | Obtain PRE-HEMATOCRIT from thawed blood and or unit to be washed: - Enter unit with sampling site coupler. NOTE: If using Terumo Sampling Site coupler, hold coupler firmly in thawed bag when obtaining sample. Coupler is not as secure in the freezing bag as a liquid unit bag. - Label plastic 12x75 tubes with unit number and “PRE”. - Retrieve 1 ml sterile syringe and 18g needle. - Insert into rubber coupler site. - Draw back 1ml of red cells. 6.6 Drain into 12x75 mm test tube |  |
| 7.0 | Deliver PRE-HEMATOCRIT specimen to closest location that can perform testing. (i.e. Hematology on Main Floor, Stat Lab on First Floor, BMT Lab (future)). No requisition is required for this testing. 7.1 Ask tech to run (SPUN if degly unit) hematocrit. NOTE: Inform them if preHCT sample contains glycerol. They need to do a spun HCT. 7.2 Label print out with unit number, “PRE” and any dilution factor used. 7.3 Enter Pre-Hematocrit and dilution factor on form. 7.4 This may be taken with the post-Hematocrit sample. |  |
| 8.0 | Wash or Degly the unit depending on the QC needed. 8.1 For Wash and labeling, refer to procedure: Washing Red Blood Cells (BB.COMP.1006) 8.2 For Degly and labeling , refer to procedure: Deglycerolization – Red Blood Cells (BB.COMP.1010) |  |
| 9.0 | Obtain weight of empty blood bag (pre-preparation bag). 9.1 Weigh empty blood bag (original bag) on component scale 9.2 Enter weight (gm) on form. |  |
| 10.0 | Enter POST-WEIGHT information: 10.1 Weigh (gm) empty transfer bag that will be used to transfer contents of Cobe set on component scale and enter results. |  |
| 11.0 | Mix the blood in the Cobe bag well at least 3 minutes. |  |
| 12.0 | Obtain POST_HEMATOCRIT from unit bag: 12.1 Transfer from Cobe processing bag to at least a 300mL transfer pack by spiking into Cobe processing bag. 12.2 Label the transfer bag before separating it from the Cobe Processing set. a. Refer to Deglycerolization: Red Blood Cells 12.3 Seal off tubing and make 1-2 segments to remain attached to bag. Prepare one long segment to use for POST HEMATOCRIT. 12.4 Label plastic 12x75mm tube with unit number and “POST”. 12.5 Work behind a shield. Cut the bottom end of the segment. Hold over labeled tube. 12.6 Make a small cut in top of segment (behind shield). 12.7 Drain into 12x75 mm test tube. |  |
| 13.0 | Weigh the unit of blood AFTER transfer into the transfer bag. 13.1 Record weight on form. |  |
| 14.0 | Deliver POST-HEMATOCRIT and PRE-HEMATOCRIT (if not taken previously) specimens to closest location that can perform testing. (i.e. Hematology on Main Floor, Stat Lab on First Floor, BMT Lab (future)). No requisition is needed. 14.1 Ask tech to run (SPUN if Degly unit) hematocrit: 14.2 Label print out with unit number, “POST” and any dilution factor used. 14.3 Enter Post-Hematocrit and dilution factor on form |  |
| 15.0 | Enter HUE results into form. Refer to CP: Washing and/or Deglycing Procedures for instructions on reading hue. 15.1 Enter HUE of last wash (final spin) in the Hue-Last Wash section. 15.2 Enter HUE of Final Fragility Test in the Hue-Final Fragility Transfusion section. |  |
| 16.0 | Interpretation of Results on form 16.1 IF; Interpretation; Record on form. % RBC recovery 80% or higher And Hue of Last Wash is 1 thru 4 And Final Fragility Hue is 1-5; Acceptable; PASS. % RBC recovery less than 80% And / or Hue of Last Wash is 5 or higher And/or Final Fragility Hue is 6 or greater; Unacceptable Verify that SPUN HCT was performed if degly unit.; FAIL Notify management Repeat testing once Repeat failure: - take Cobe out of use. - Message on Cobe “Do not Use” date/initials - Notify Clinical Services to maintenance the Cobe. - Record problem on Maintenance form. - Do appropriate before and after validation for service.. 16.2 Calculations during Downtime are performed automatically on the form. a. In the event of a computer failure, all necessary formulas are included on the downtime form found in the forms drawer. Attachment 3: (Downtime) Percent Recovery-Unit Calculation Form. |  |
| 17.0 | Print form for review. |  |
| 18.0 | Submit for Management review. |  |
| 19.0 | Sign off on the appropriate summary checklist for either washed or deglycerolized recovery. |  |

2. Procedure: IV:  Verify Programs on Cobe 3, Cobe 5 and Cobe 6 or new additional Cobes with

electronic programs.

Chemical Risk Assessment:  Low

Biological Risk Assessment: Low

Protective Equipment:  Lab coat, gloves

Reagents: NA

Supplies: NA

Equipment: Cobe 3, Cobe 5, Cobe 6

Specimen Requirements:  NA

Frequency: Monthly

| STEPS | INSTRUCTIONS | CHANGE/ APPROVAL |
| --- | --- | --- |
| 1.0 | Cobes 3, 5 and 6 have an electronic program that controls the valves that open and close instead of the diode pins that are used for the older model Cobes. Fig. 1 Digital Display of Programs Program Cycle Time (min) Up/down arrows to move up/down Valve Select RCO (Red Cell over ride) - Cobe 3 has Programs 0 to 4. - Cobe 5 and Cobe 6 have Programs 0 to 8. |  |
| 2.0 | Verify each Program using the following steps: - Press the up/down arrows under program to the program number desired. - Press the up/down arrows under the cycle to move through the pre-set steps. - Start with 1 and move up comparing to the corresponding diagram. - As each number comes up verify that the step matches the correct spin time. - Verify that the correct valve is lit. - Verify RCO is lit when indicated. - Notify management if the Cycles do not match the Program and put a sign on the Cobe that is should not be used for that particular program. |  |
| 3.0 | Program 0: Check Prime 3.1 Select Program “0”. 3.2 Verify cycles Cycle; Time; RCO; STOP; VALVE. CALL; PC; RC; 1; 2; 3. 1; 1 minute; ○; ○; ○; ○; ○; ●; ○. 2; 1.5 minutes; ○; ○; ○; ○; ○; ○; ●. 3; 1 minute; ○; ○; ○; ○; ○; ●; ○. 4; 1 minute; ○; ○; ○; ○; ○; ●; ○. 5; 1minute; ○; ○; ○; ○; ○; ●; ○. 6; 1 minute; ○; ○; ○; ○; ○; ●; ○. 7; 1 minute; ○; ○; ○; ○; ○; ●; ○. |  |
| 4.0 | Program 1: Degly HiHCT 4.1 Follow the steps list in 2.0 and verify that the times and valves match the diagram below. Cycle; Time; RCO; STOP; VALVE. CALL; PC; RC; 1; 2; 3. 1; 2 minutes; ○; ○; ○; ○; ●; ○; ○. 2; 2 minutes; ○; ○; ○; ○; ○; ●; ○. 3; 1.5 minutes; ○; ○; ○; ○; ○; ○; ●. 4; 1.5 minutes; ○; ○; ○; ○; ○; ○; ●. 5; 1.5 minutes; ○; ●; ●; ○; ○; ○; ○. 6; 1 minute; ○; ○; ○; ○; ○; ●; ○. 7; 1 minute; ○; ○; ○; ○; ○; ●; ○. |  |
| 5.0 | Program 2: Extra Wash 5.1 Follow the steps listed in 2.0 and verify that the times and valves match the diagram below. Cycle; Time; RCO; STOP; VALVE. CALL; PC; RC; 1; 2; 3. 1; 1.5 minutes; ○; ○; ○; ○; ○; ○; ●. 2; 1.5 minutes; ○; ○; ○; ○; ○; ○; ●. 3; 1.5 minutes; ○; ●; ●; ○; ○; ○; ○. 4; 1 minute; ○; ○; ○; ○; ○; ●; ○. 5; 1 minute; ○; ○; ○; ○; ○; ●; ○. 6; 1 minute; ○; ○; ○; ○; ○; ●; ○. 7; 1 minute; ○; ○; ○; ○; ○; ●; ○. |  |
| 6.0 | Program 3: 1 Liter Wash Packed Cells 6.1 Follow the steps listed in 2.0 and verify that the times and valves match the diagram below. Cycle; Time; RCO; STOP; VALVE. CALL; PC; RC; 1; 2; 3. 1; 2.5 minutes; ○; ○; ○; ○; ○; ●; ○. 2; 1.5 minutes; ○; ○; ○; ○; ○; ●; ○. 3; 1.5 minutes; ○; ○; ●; ○; ○; ○; ○. 4; 1 minute; ○; ○; ○; ○; ○; ●; ○. 5; 1 minute; ○; ○; ○; ○; ○; ●; ○. 6; 1 minute; ○; ○; ○; ○; ○; ●; ○. 7; 1 minute; ○; ○; ○; ○; ○; ●; ○. |  |
| 7.0 | Program 4: Wash Platelets 7.1 Follow the steps listed in 2.0 and verify that the times and valves match the diagram below. Cycle; Time; RCO; STOP; VALVE. CALL; PC; RC; 1; 2; 3. 1; 15 minutes; ○; ○; ○; ○; ○; ○; ●. 2; 5 minutes; ○; ○; ○; ○; ○; ○; ●. 3; 2.5 minutes; ○; ○; ●; ○; ○; ○; ○. 4; 1 minute; ○; ○; ○; ○; ○; ●; ○. 5; 1 minute; ○; ○; ○; ○; ○; ●; ○. 6; 1 minute; ○; ○; ○; ○; ○; ●; ○. 7; 1 minute; ○; ○; ○; ○; ○; ●; ○. |  |
| 8.0 | Program 5: Volume Reduce Platelets 8.1 Follow the steps listed in 2.0 and verify that the times and valves match the diagram below. Cycle; Time; RCO; STOP; VALVE. CALL; PC; RC; 1; 2; 3. 1; 20 minutes; ○; ○; ●; ○; ○; ○; ○. 2; 1 minute; ○; ○; ○; ○; ○; ●; ○. 3; 1 minute; ○; ○; ○; ○; ○; ●; ○. 4; 1 minute; ○; ○; ○; ○; ○; ●; ○. 5; 1 minute; ○; ○; ○; ○; ○; ●; ○. 6; 1 minute; ○; ○; ○; ○; ○; ●; ○. 7; 1 minute; ○; ○; ○; ○; ○; ●; ○. |  |
| 9.0 | Program 6:Washing Pedi Units 9.1 Follow the steps listed in 2.0 and verify that the times and valves match the diagram below. Cycle; Time; RCO; STOP; VALVE. CALL; PC; RC; 1; 2; 3. 1; 2.5 minutes; ○; ○; ○; ○; ○; ●; ○. 2; 1.5 minutes; ○; ○; ●; ○; ○; ●; ○. 3; 1 minute; ○; ○; ○; ○; ○; ●; ○. 4; 1 minute; ○; ○; ○; ○; ○; ●; ○. 5; 1 minute; ○; ○; ○; ○; ○; ●; ○. 6; 1 minute; ○; ○; ○; ○; ○; ●; ○. 7; 1 minute; ○; ○; ○; ○; ○; ●; ○. |  |
| 10.0 | Program 7: 1 Liter Wash Whole Blood 10.1 Follow the steps listed in 2.0 and verify that the times and valves match the diagram below. Cycle; Time; RCO; STOP; VALVE. CALL; PC; RC; 1; 2; 3. 1; 5 minutes; ○; ○; ○; ○; ○; ●; ○. 2; 5 minutes; ○; ○; ○; ○; ○; ●; ○. 3; 1.5 minutes; ○; ○; ○; ○; ○; ●; ○. 4; 1.5 minutes; ○; ○; ●; ○; ○; ○; ○. 5; 1 minute; ○; ○; ○; ○; ○; ●; ○. 6; 1 minute; ○; ○; ○; ○; ○; ●; ○. 7; 1 minute; ○; ○; ○; ○; ○; ●; ○. |  |
| 11.0 | Program 8: 2 Liter Wash 11.1 Follow the steps listed in 2.0 and verify that the times and valves match the diagram below. Cycle; Time; RCO; STOP; VALVE. CALL; PC; RC; 1; 2; 3. 1; 2.5 minutes; ●; ○; ○; ○; ○; ●; ○. 2; 1.5 minutes; ●; ○; ○; ○; ○; ●; ○. 3; 1.5 minutes; ○; ○; ○; ○; ○; ○; ●. 4; 1.5 minutes; ○; ○; ○; ○; ○; ○; ●. 5; 1.5 minutes; ○; ○; ●; ○; ○; ○; ○. 6; 1 minute; ○; ○; ○; ○; ○; ●; ○. 7; 1 minute; ○; ○; ○; ○; ○; ●; ○. |  |

3.  Review/Revised/implemented:

All protocols must be reviewed as stated in the document change control.

All new protocols that have major revisions must be signed by the CLIA Director.

All reviewed protocols with minor revisions can be signed by the designated section medical

Director or designee.

4.        Related Policies/Procedures:

Cobe QC Training

5. 	References:

TECHNICAL MANUAL.  Bethesda, MD:  American Association of Blood Banks, updated periodically.

COBE 2991 Cell Processor Operations Manual

6. 	Attachments:

Attachment 1: Weekly and Monthly CP Checklist

Attachment 2: Cobe/IBM Component Prep Worksheet

Attachment 3: Percent Recovery-Washed Calculation Form/Deglycerolized Unit Calculation Form

7. 	Revised/Reviewed Dates and Signatures:

See Document Change Control

| Document Change Control |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Title: Cobe Quality Control Page 1 of 2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Previous title: Checking Prime and Cobe RBC % Recovery for Washing and Deglycerolized (Degly) Red Cells. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Written date |  |  |  |  |  |  |  |  | Written by: |  |  |  |  |  |  |  |  |  |
| Validation date |  |  |  |  |  |  |  |  | Validation by |  |  |  |  |  |  |  |  |  |
| Reviewed date |  |  |  |  |  |  |  |  | Reviewed by |  |  |  |  |  |  |  |  |  |
| Approved date |  |  |  |  |  |  |  |  | Approved by |  |  |  |  |  |  |  |  |  |
| Approved date |  |  |  |  |  |  |  |  | Approved by |  |  |  |  |  |  |  |  |  |
| Effective date in use |  |  |  |  | 10/11/07 |  |  |  | In use by |  |  |  | See archive history |  |  |  |  |  |
| Revisions |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Revised Date | By |  |  | MD Date |  | By |  | MD Date |  | By |  | Review Date |  |  | By |  | Effective Date | By |
| 12/4/12 | JF |  |  | 12/17/12 |  | EF |  |  |  |  |  | 12/17/12 |  |  | MRJ |  | 12/18/12 | MRJ |
| Validate Date | By |  |  | Revisions: - Excel form with calculations for RBC recovery available on the BBSTAFFINFORMATION tab. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 12/17/12 | KJ |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Revised Date |  | By |  | MD Date |  | By |  | MD Date |  | By |  | Review Date |  |  | By |  | Effective Date | By |
| 12/31/12 |  | JHS |  | NA |  | NA |  | 2/27/13 |  | EF |  | 2/26/13 |  |  | MRJ |  | 2/28/13 | JHS |
| Validate Date |  | By |  | Revisions: Added other locations that can perform Hematocrit testing. Noted that procedure should be rotated among Cobes. Transfer QC units to a transfer pack from Cobe processing set to obtain the post Hematocrit sample so that unit can be used for transfusion. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 1/9/13 |  | KJ |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Revised Date |  | By |  | MD Date |  | By |  | MD Date |  | By |  | Review Date |  |  | By |  | Effective Date | By |
| 6/18/15 |  | JHS |  | GP |  | 7/15/15 |  | 7/8/15 |  | EF |  | 7/6/15 |  |  | JHS |  | 7/15/15 | JHS |
| Validate Date |  | By |  | Revisions: Combined Checking Prime procedures and Cobe RBC % Recovery for washing and Deglycerolized (Degly) Red Cells into one procedure. Dr. Fadeyi approved checking Cobes annually for 80% recovery for only deglycing and added section to check for plasma removal for washing red cells. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 7/2/15 |  | WE |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Revised Date |  | By |  | MD Date |  | By |  | MD Date |  | By |  | Review Date |  |  | By |  | Effective Date | By |
| 2/20/18 |  | JHS |  | 2/28/18 |  | GP |  | 2/28/18 |  | EF |  | 2/22/18 |  |  | MRJ |  | 3/1/18 | JHS |
| Validate Date |  | By |  | Revisions: Modified the way the information is displayed for checking the programs on Cobe 3 and 5 in Section V to make it clearer. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 2/23/18 |  | SMA |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Locations |  |  |  |  |  |  | Out of Use: Date: |  |  |  |  |  |  | By |  |  |  |  |
|  |  |  |  |  |  |  | Reason |  |  |  |  |  |  |  |  |  |  |  |

Reviews: Record date/initials

| Date | Initials | Date | Initials | Date | Initials | Date | Initials |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 12/2/13 | JHS | 11/13/17 | JHS |  |  |  |  |
| 12/2/13 | GP |  |  |  |  |  |  |
| 11/9/15 | MRJ |  |  |  |  |  |  |

| Document Change Control |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Title: Cobe Quality Control Page 2 of 2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Previous title: Checking Prime and Cobe RBC % Recovery for Washing and Deglycerolized (Degly) Red Cells. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Written date |  |  |  |  |  |  |  |  | Written by: |  |  |  |  |  |  |  |  |  |
| Validation date |  |  |  |  |  |  |  |  | Validation by |  |  |  |  |  |  |  |  |  |
| Reviewed date |  |  |  |  |  |  |  |  | Reviewed by |  |  |  |  |  |  |  |  |  |
| Approved date |  |  |  |  |  |  |  |  | Approved by |  |  |  |  |  |  |  |  |  |
| Approved date |  |  |  |  |  |  |  |  | Approved by |  |  |  |  |  |  |  |  |  |
| Effective date in use |  |  |  |  | 10/11/07 |  |  |  | In use by |  |  |  | See archive history |  |  |  |  |  |
| Revisions |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Revised Date | By |  |  | MD Date |  | By |  | MD Date |  | By |  | Review Date |  |  | By |  | Effective Date | By |
| 1/31/19 | JHS |  |  |  |  |  |  | 2/26/19 |  | EF |  | 2/26/19 |  |  | JS |  | 2/28/19 | JS |
| Validate Date | By |  |  | Revisions: Doing 80% recovery on washing and deglycing in the same manner per an earlier version. Removed section IV that related to testing for protein. Added Cobe 6 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 2/26/19 | SMA |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Revised Date |  | By |  | MD Date |  | By |  | MD Date |  | By |  | Review Date |  |  | By |  | Effective Date | By |
| 4/16/20 |  | KP |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Validate Date |  | By |  | Revisions: removed super prime from definitions, Sections, and procedures. Removed Mary R. Jones. Added Christina S. Warren. Removed Hydraulic system Prime for Cobe BMT from Sections and procedures. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Revised Date |  | By |  | MD Date |  | By |  | MD Date |  | By |  | Review Date |  |  | By |  | Effective Date | By |
| Validate Date |  | By |  | Revisions: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Revised Date |  | By |  | MD Date |  | By |  | MD Date |  | By |  | Review Date |  |  | By |  | Effective Date | By |
| Validate Date |  | By |  | Revisions: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Locations |  |  |  |  |  |  | Out of Use: Date: |  |  |  |  |  |  | By |  |  |  |  |
|  |  |  |  |  |  |  | Reason |  |  |  |  |  |  |  |  |  |  |  |

Reviews: Record date/initials

| Date | Initials | Date | Initials | Date | Initials | Date | Initials |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 12/2/13 | JHS | 11/13/17 | JHS |  |  |  |  |
| 12/2/13 | GP |  |  |  |  |  |  |
| 11/9/15 | MRJ |  |  |  |  |  |  |