## Slide 1
i-STAT Testing – CG4
Prepared by Laboratory Compliance, Quality, Safety, and
Point-of-Care Testing Team Updated May 2020

## Slide 2: General Information
  Contact Information
  Policies/Procedures
  Training and Competency
  i-STAT Access
  Safety
  Components
  Testing Process Overview

## Slide 3: i-STAT Support: Contact Information
Kelly Leiner 781-2421 (High Point Medical Center)
Trish Sorensen 713-4136
Liz Gregory 713-0377
Ray Dyer 713-4137
OR Blood Gas Lab 716-2181 (M-F 7am-7pm)
Site-specific preceptors
Refer to the i-STAT System Manual and medical center specific i-STAT policies
WFBMC POCT web site

## Slide 4: i-STAT Policies/Procedures/Guidelines
Posted on the WFBMC Intranet—Go to: Departments—Point of Care Testing—Policies/Procedures/Guidelines
The i-STAT System Manual (Abbott POC website) can be found under Resources-- Resource Links—Manuals

## Slide 5: Only Authorized Users May Perform i-STAT Testing: Training and Competency Requirements
Initial training and evaluation for proper test performance
6 month competency assessment-NEW users
On-going Annual competency assessment
  Must meet all 6 points of competency assessment, as defined by CLIA— refer to page 1 of competency observation form
Failure to maintain competency will result in loss of testing privileges
Competency observation may only be completed by an authorized staff member
Written exams 80% correct or remediation will be required (All necessary exams must be complete PRIOR to competency observation)

## Slide 6: i-STAT Access (Operator ID)
  Training=i-STAT access
Employee ID=operator ID
Do NOT share your employee/operator ID
Do NOT enter your ID into analyzer and allow testing by another individual

## Slide 7: Safety
Always wear gloves
        Patient testing
        Performing/handling Quality Controls (QC)
        Disinfecting analyzer
Disinfect analyzers/printers
      When contaminated with blood AND
      Between EACH patient
      Follow medical center policies
DO NOT lean over the cartridge when filling
    A safety shield is recommended
Place gauze or tissue over snap when closing

## Slide 8: Components of the i-STAT System
Primary System Components
Cartridge
Handheld
Downloader
Printer

## Slide 9: Quick Overview: The i-STAT System Testing Process
Enter information into handheld
Obtain the blood specimen and fill cartridge
Insert cartridge; results will appear within minutes
Download results

## Slide 10: The i-STAT Analyzer
External Components
Battery Replacement/Charging
Barcode Scanner
Test Menu
Menu Options

## Slide 11: i-STAT Analyzer
Turns off after 2 minutes of non-use
Can be turned on by pressing the on/off keypad—circle with line
If operator, patient, and cartridge information has been entered, the analyzer will remain on for 15 minutes

## Slide 12: Handheld External Components
Battery
compartment
Infrared eye
Barcode scanner
Cartridge port

## Slide 13: Batteries: Hands On
9-volt lithium disposable batteries or battery pack
  HPMC and Main campus sites use the Abbott rechargeable battery pack
Change/charge the batteries:
Flashing battery icon
“Battery Low ” message
Do not discard rechargeable batteries
Green light will be on if battery in compartment is seated properly and charging
Battery
compartment
9-volt lithium disposable batteries
Abbott rechargeable battery pack
Battery
compartment

## Slide 14: The Barcode Scanner
To scan
Hold 3 to 12 inches away from the barcode
Press and hold down the SCAN key
Laser beam must cover the entire length of the barcode
Wait for the beep or for laser beam to disappear
Release the SCAN key
Barcode scanner
The barcode scanner is a class II laser device. Momentary exposure is not known to be harmful—however you should never point the scanner at anyone or look into the laser beam.

## Slide 15: Handheld Test Menu—Hands On
Test Menu
First menu to appear
Important Note:
This menu is used for patient
tests only!
Cartridge Test Menu options
1 = Last Result
2 = i-STAT Cartridge (for a patient test)

## Slide 16: Administration Menu Options Hands On
1--Analyzer Status
      Battery voltage
          Maintain charge above 8 volts
      Shows Total records and Unsent records
2--Data Review
      Review/print stored data
      #7 – List All Results
3--Quality Tests
      Quality control checks

## Slide 17: Cartridges
  General Information
  Labeling
  Storage
  Handling

## Slide 18: General Cartridge Information
Self-contained lab analyzers
Analysis, sample manipulation, and waste are confined within the cartridge
Single use
24 or 25 per box
Use IMMEDIATELY after opening
Requires only a few drops of blood
Cartridges contain biosensors that “measure” the analytes being tested.
Abbott offers different i-STAT cartridge configurations.
  Only use cartridges authorized for your test site.

## Slide 19: i-STAT Cartridge Labeling
Panel name
Analytes
Room temperature storage expiration date
Refrigerated expiration date
Lot number
Cartridge barcode
Measured
Calculated
Cartridges MUST be brought to room temperature
5 minutes for a single cartridge
1 hour for a box of cartridges
Do NOT manipulate the cartridge in any way to bring it to room temperature more quickly.

## Slide 20: Cartridge Storage
ALL refrigerated cartridges must be stored 2- 8˚C
  Storage environment MUST BE MAINTAINED and monitored 24/7/365
    Monitored by SPOT
If temperature is ever outside of 2 to 8˚C, contact the POCT office for assistance.
  Do NOT use the cartridges for patient testing, until consulting the POCT office.
Cartridges should NOT be returned to refrigerated storage once they reach temperatures above 8˚C.
Do NOT store or expose cartridges to extreme cold or heat such as freezing or stored above lights or computers

## Slide 21: Cartridge Storage (continued)
Room temperature (RT) storage
14 days or 2 months, depending upon cartridge type
Room temperature expiration should NEVER exceed manufacturer expiration date
RT 18°C - 30°C (64-86°F)
Do NOT use cartridges past manufacturer’s expiration date or past room temperature expiration date.

## Slide 22: i-STAT Cartridge Handling
Do NOT reuse cartridges
Open cartridge IMMEDIATELY prior to use.
Carefully tear open the cartridge pouch where indicated (do NOT leave cartridges exposed to air or moisture)
Hold the cartridge by the sides or the bottom
Do not touch the sensors or apply pressure to the center of the cartridge (be careful when putting room temperature expiration date on cartridges – write or apply sticker at line on bottom – not in the middle)

## Slide 23: i-STAT Cartridge Components

## Slide 24: Quality Controls (QC)
  Simulators
  Liquid QC

## Slide 25: i-STAT System Quality Control
i-STAT System Quality Control
Internal Electronic Simulator
External Electronic Simulator
Liquid Quality Controls

## Slide 26: Simulators
Electronic Quality Control
  Validates i-STAT analyzer
  Pass or Fail
QC results automatically documented when analyzer downloaded

## Slide 27: Internal Electronic Simulator
Internal checks ensure proper and accurate functioning
Performed automatically
  Every 8 hours of patient use
  With each cartridge type
If simulator passes, testing proceeds/patient results will be displayed
Results will not be given if the simulator fails.
If the “Electronic Simulator Fail” message appears,
  Do Not proceed with patient testing until issue is resolved.
  Test external electronic simulator

## Slide 28: External Electronic Simulator--Hands On
Same as the internal simulator but is an external device and can be tested upon demand.
Simulator is performed if:
Handheld is dropped
Internal Electronic Simulator failure
Error code occurs that indicates the simulator should be tested
Questionable analyzer performance
Ceramic cartridge cleaning procedure
External Electronic Simulator
Do not remove simulator when “Simulator Locked” is displayed

## Slide 29: Liquid Quality Control
Validates performance of test cartridges
Consists of multiple levels—test specific
Should be performed:
      Each new cartridge shipment or new cartridge lot #, PRIOR to patient use
      Monthly—by testing site staff
      Questionable performance cartridges/analyzer

## Slide 30: Liquid Quality Controls for CG4
3 levels
Monitors cartridge performance
Store 2°C to 8°C
Room temperature for 4 hours
Expiration Date

## Slide 31: Liquid QC Important Notes/Reminders
Analyzers programmed with QC ranges for the liquid QC.
  Pass or Fail will display
If liquid QC fails, a comment code and corrective action is required:
      Use comment code 5—recheck/confirm
      Repeat liquid QC
      Creates a STOP point for the testing staff member
      Patient testing lockout does NOT occur
Staff MUST ensure QC results are within acceptable limits PRIOR to proceeding with patient testing.

## Slide 32: Patient Identification
  Process Flow
  Patient Identity

## Slide 33: Process Flow
Ensure documented physician order or documented protocol before performing i-STAT testing.
Patient ID:
  CSN for that specific date of service or current admission CSN
After testing is completed, analyzer should be downloaded.
If ID match is found, results post to Wake One and billing occurs, regardless of provider order

## Slide 34: Patient identity
Follow medical center policy for verification of patient identity—
  2 patient identifiers at the patient side (name and Date of Birth (DOB))
Scan patient ID DIRECTLY from the patient armband attached to the patient
If an ID barcode is scanned that is not physically attached to the patient, 4 identifiers (name, DOB, MRN, CSN) MUST be verified.
  Include label/armband verification in the procedure time-out.
Ensure sample identity throughout entire testing and reporting process.
  Verify patient ID again on analyzer display.

## Slide 35: Sample Information
  Sample Collection/Sample Considerations
  Filling the cartridge
  Running a patient test

## Slide 36: Sample Collection/Considerations:
ONLY whole blood should be tested on the i-STAT
  Arterial
  Venous
Important notes:
  Make sure to follow medical center procedures for proper sample collection.
  Always correlate results with the status of the patient!
Consult Intended Use Statements for
APOC Products in the Cartridge
  and Test Information (CTI) sheets at:
www.abbottpointofcare.com
Within the i-STAT 1 System Manual

## Slide 37: In-Dwelling Line
  Adequately ‘clear’ line
    Inaccurate results will be given if samples are contaminated.
      Saline Contamination can cause:
        Falsely decreased pCO2,
  i-STAT Recommendation:
    Five to six times the volume of the catheter, connectors, and needle should be collected as ‘waste’

## Slide 38: Blood Gas Cartridges— CG4
Use a syringe
Seal sample with cap IMMEDIATELY after collection.
  Do not contaminate with air
    pO2 will be falsely increased if air contamination occurs
Heparinized samples, test within 10 minutes
Non-heparinized samples, test immediately
ALWAYS test CG4 cartridge immediately after collection--lactate will be falsely increased if testing is delayed

## Slide 39: Cartridge Filling Overview
Mix the sample thoroughly and gently for 15 seconds
Discard first few drops to check for clots and to get rid of any micro air bubbles
  NEVER test a sample that has or has had a clot. Inaccurate results may be obtained.
Fill to the fill mark and close the closure to seal
Insert IMMEDIATELY into the analyzer
Do not move analyzer after cartridge has been inserted
Fill mark
closure is closed

## Slide 40: Running a Patient Test: Overview of the Chart Page
Allows entry of patient data
  Sample type
    Enter the number next to the sample type
    Press the ENT key
  Patient temperature
    ONLY if requested by provider
  FIO2
  Free fields for additional info
    Free fields do not populate the electronic medical record

## Slide 41: Results
  Interpreting Results
  Comment Codes
  Result Considerations
  Quality Check Codes
  Printing Results
  Transmitting Results
  Enter/Edit of Results – Wake One
  Supporting Documents

## Slide 42: Interpreting Test Results: The Results Screen
Patient ID
Time and date of test
Type of cartridge
Test results

## Slide 43: Interpreting Test Results: Critical Value Alerts
Indicated by up or down arrows
Evaluate with patient’s clinical symptoms
  As appropriate, perform repeat testing
    Values not consistent with patient symptoms --repeat and validate by an alternate test method
      Use Main Lab)
Follow medical center policy for handling critical results

## Slide 44: Critical Values – found in i-STAT procedure on POCT website
The patient provider must be notified of critical values:
Documentation must include:
  Result
  Date/Time of notification
  Name of provider notified
  Name of notifying staff member
  Critical values must be notified within 15 minutes

## Slide 45: Interpreting Test Results: Out of Reportable Range Flags
Out of Reportable Range flag
Indicated with a <xxx or >xxx
Suppressed Result Flag <>
Result dependent on the out of reportable range result

## Slide 46: Interpreting Test Results: Star Outs (***)
*** appears in place of the test result
  Sensor on the cartridge is unable to report a result
Use comment code 0 (zero) to prevent result from posting to the patient record
  Failure to enter comment code 0 into the i-STAT handheld at the time of testing=results post to Wake One
    The star out result will report as INSTRUMENT ERROR.
ACTION REQUIRED:
  Retest the sample
  If star out values continue, recollect and send the sample to the Main lab for testing.

## Slide 47: Responding to Patient Tests: Comment Codes
Numeric codes
      ‘Automates’ result hold or posting of result to the patient record
      Affects patient billing for testing
Comment field--top of the first page of the results screen
  MUST be entered at the time of testing AND before the analyzer powers down
  Question: What can you do if you realize that you have misidentified a patient sample immediately after inserting the cartridge?
Code | Translation | Result Post to Wake One | Allows charge to patient | Use
0 (zero) | Procedure Error | NO | NO | Star out Questionable Results Error in testing
5 | Recheck/Confirm | Yes | NO | Result may be accurate and test will be repeated
6 | To Notify Provider | Yes | Yes | Documents intent to notify provider of result
123 | Lab Confirmation to Follow | Yes | NO | Use when sending repeat testing to the laboratory

## Slide 48: Result Considerations: Factors That May Affect Results
Avoid factors that may affect the sample prior to analysis
Always correlate results with the status of the patient
Do NOT Move/Carry the analyzer during sample testing
  NO Vibration or tilting analyzer
References:
Cartridge and Test Information Sheets (CTI Sheets)
  i‑STAT 1 System Manual
  www.abbottpointofcare.com

## Slide 49: Result Considerations: Cartridge and Test Information Sheet

## Slide 50: Quality Check Codes:
i-STAT Quality System
Numerous quality checks during patient test cycle
Failed quality check
  Cause message— pay attention to the code number
  Action message
  Quality Check Code
      Do NOT continue with testing until the cause of the error is clearly understood
      If the same error code occurs multiple times with no apparent cause, notify the POCT office
      Consider using a different analyzer and a different batch of test cartridges

## Slide 51: Quality Check Codes

## Slide 52: Quality Check Codes

## Slide 53: Printing Results with the i-STAT Portable Printer
Display results and align the infrared eyes of each device
Press the printer icon key on the handheld
Printer uses rechargeable batteries—MUST remain plugged into outlet and turned off when not in use
Verify patient ID on printout against patient sample ID

## Slide 54: Transmitting Results: Downloader
Transmit results after each test/case:
Turn handheld off
Place handheld in the Downloader
Look for display screen to show arrows and the message “Communication in Progress”
When transmission is complete:
Screen goes blank
Handheld turns off

## Slide 55: Transmitting Results: Troubleshooting
If the downloader does Not display “Communication in Progress” or an error message appears
      Check for a green power light
      Check all cord/cable connections to the down loader
      Try unplugging and re-plugging the cables
      If troubleshooting does not resolve problem, report problem to the Help Desk (Call 6-HELP)

## Slide 56: Editing of results once posted to Wake One
Refer to i-STAT Procedure or “Tips and Tricks: Procedure to correct i-STAT results in Wake One”
Situations that require editing of results:
  Sample Misidentification
    Also requires i-STAT resolution requisition found on POCT web site
  Erroneous temperature corrected blood gas
  Questionable i-STAT results if posted to Wake One
  Tests performed but not ordered by the physician
  Documentation of critical value notification if not done in another section of Wake One

## Slide 57: Supporting Document(s)
i-STAT Resolution Requisition
  (Located on POCT website – Forms and Records)

## Slide 58: Congratulations
This concludes the i-STAT education module.
Please note that this power point does not replace the i-STAT policies and procedures.
The presentation only gives highlights of i-STAT information.
Please complete the on-line i-STAT exams within 1 week of training.