## Slide 1
AVOX 1000E
Education Module
Point of Care Testing Department
Updated – July 2025

## Slide 2
Avox Support
Avox manufacturer Werfen 1-800-678-0710
Lab Point of Care Office –
  716-7248
  713-4136
Contact your site manager or specific preceptors
This presentation does not include all information regarding AVOX testing. It only highlights information.
For complete information, refer to the AVOX procedure and manufacturer manual in Policy Tech.

## Slide 3
Safety
Always wear gloves when handling analyzer or samples including patient testing and performing QC.
Disinfect when contaminated with blood AND between EVERY patient.
Routine external cleaning is recommended for infection control purposes.
When performing liquid QC, protect fingers by carefully breaking open glass ampule.

## Slide 4
Avox 1000E - Basic Principle
The AVOX 1000E test device is used to measure:
  Fractional oxyhemoglobin (%HbO2)
  Total hemoglobin (THb) concentration
Blood samples are injected into a sealed, single-use disposable cuvette that is inserted into the sample chamber.
The Avox 1000E illuminates the sample at multiple wavelengths, records the optical density of the sample at each of the wavelengths, and computes the results.

## Slide 5
Avox 1000E - Components
Avox analyzer
Cuvettes
Orange and yellow quality control filters
Liquid quality control materials

## Slide 6
Avox 1000E - Analyzer
To power ON the analyzer
  Press the Enter/On key
  Confirm that self test is OK
  The analyzer indicates that it is ready to analyze samples by displaying ‘READY’
Troubleshoot if self test FAILS
  Do not use the device for patient testing until failure is resolved

## Slide 7
Avox - Cuvettes
Single use Avox disposable cuvettes (P/N C-100B)
Stored at room temperature (15-30 C) in tightly sealed bag with desiccant.
One required per sample
DO NOT USE for patient testing if desiccant indicator is PINK!!! Patient results may be inaccurate.
The CORRECT cuvette path length must be entered into the analyzer prior to patient or QC testing.
Cuvettes and QC materials should be marked with the date they are received. When supplies are opened and put into use, they should be logged on the appropriate log. All records are kept a minimum of two years.

## Slide 8
Avox - Quality Control
If you perform patient testing on Avox, you are responsible for ensuring the required QC has been performed on the analyzer and cuvettes, PRIOR to testing patient samples.
Everyone who performs patient testing must also perform daily QC and liquid QC sometime during the year.
NEVER use an analyzer or cuvettes for patient testing if required QC checks are not within acceptable limits.
Document ALL QC values on the QC logs, including QC failures. This helps track on-going problems.
The daily QC log MUST have QC values or “Not in Use” (NIU) documented for each day.
If you document NIU, you must sign/initial next to NIU.

## Slide 9
Avox - Quality Control
Daily:
•The AVOX 1000E is checked for calibration accuracy and optics cleanliness with the use of Quality Control Filters.
Weekly:
•The instrument and cuvettes are checked for accuracy and calibration stability using 2 levels of Liquid Quality Control
Semiannually:
•The AVOX 1000E is compared against the Clinical Laboratory Co-Oximeter every 6 months and linearity is verified by testing a linearity kit.

## Slide 10
Avox - Quality Control - Filters
Orange and Yellow QC filters.
The QC filters are device specific Do NOT interchange between analyzers
QC filters must be tested each day of patient use
Filters Assesses the status of the optical system, confirms calibration and verifies performance.
If blood has contaminated the optics, the readings from the QC filters will not fall within the acceptable ranges. Please wipe the filters and retest.
All QC results should be documented on the AVOX QC log

## Slide 11
Avox - Quality Control - Filters
Verify that the serial numbers match between QC filters and AVOX analyzer.
Wipe off the cuvette-shaped yellow and orange optical filters.
Observe that the “READY” message is displayed and insert the QC filter into the AVOX.
Wait until readings appear on the display.
Record results on the provided logs and review for acceptability.
Repeat with second QC filter.
If BOTH THb and %HbO2 readings are within specified ranges, the AVOX passes the Quick Cal Check and the Test for Spilled Blood.
If any one of the four values are outside of the specified range, the AVOX fails quality tests and can NOT be used for reporting patient results until the problem has been resolved.
  Clean the filters and repeat QC.
  If unable to resolve, contact site manager and/or POCT person and take device out of service.
  Site manager can contact clinical engineering (Trimedx)

## Slide 12
Avox - Quality Control – Liquid QC
Two levels of Liquid controls are performed to verify performance of AVOX analyzer and test cuvettes.
LQC is tested:
    Weekly on all AVOX analyzers
    Whenever the cuvette path length is changed in an analyzer
    Whenever analyzer or cuvette performance is in question
    Post analyzer repair or calibration--prior to use for patient testing.
Liquid QC performance should be rotated among all users.
All LQC results should be documented on the AVOX liquid QC log
Acceptable ranges can be found on the lot specific insert that comes with each box of AVOX liquid QC.
RNA Medical QC 253 Liquid QC Levels 1 and 3
Stored at 2-8 C until expiration date
Do NOT need to equilibrate to room temperature prior to use
Call 1-800-533-6162 for order information.

## Slide 13
Avox - Quality Control – Liquid QC
The LQC should be tested immediately after removal from refrigerator.
Hold ampule at the top and bottom (with forefinger and thumb) and gently invert to mix the solution.
Restore the QC material to the bottom of the ampule and let sit for about a minute.
Using finger protection, CAREFULLY break open the ampule of control solution and SLOWLY withdraw 0.2 to 0.3 ml into a 1 ml syringe. Use a blunt needle.
EXTREME CARE SHOULD BE USED IN OPENING THE QC AMPULES. To avoid accidental sharps exposure, the ampule top should be covered with gauze, tissue, or a protective plastic sleeve when opening.
Avoid getting air bubbles in the syringe or AVOX results may be adversely affected.
Remove blunt needle from the syringe. Discard the needle in a biohazard sharps container.
Immediately place an AVOX cuvette on the syringe.
Fill the cuvette according to illustrations in the AVOX Operator’s/Service Manual.
Observe the “READY” message is displayed and insert the Control filled cuvette, with syringe attached, into the AVOX.
Record results on the AVOX Liquid QC log and review for acceptability. If all readings are within the specified ranges, the AVOX passes the liquid QC check.

## Slide 14
Avox - Quality Control – Liquid QC
DO NOT perform patient testing if any QC value is outside of acceptable ranges that’s on the RNA medical lot specific insert.
Remove the device AND cuvettes from patient use and proceed with troubleshooting procedures by verifying the following:
  Correct cuvette path length is programmed in the analyzer
  Confirm desiccant indicator is blue in the cuvette package
  Ensure liquid controls have been stored and handled appropriately
Repeat the same level of liquid control using a new ampule and a new cuvette from the same box of cuvettes.
If repeat results are within acceptable limits, then the analyzer and cuvettes may be used for patient testing.
If repeat results are still outside of acceptable limits, contact POCT person and Clinical Engineering for assistance.

## Slide 15
Avox – Cuvette Calibration
Each time a new bag of cuvettes is opened, the AVOX cuvette path length must be updated in the analyzer and/or verified.
Instructions are in the AVOX procedure and inside the cuvette box lid.
Prior to patient use and after entering a new cuvette path length in the AVOX analyzer, both levels of QC filters and liquid QC must be tested to verify analyzer/cuvette performance.

## Slide 16
Avox – Bi-annual analyzer comparison
Must be completed by user site on each AVOX analyzer:
  Each 6 months
  After analyzer repair
Two patient samples are tested to validate comparability of results with the Clinical Laboratory Co-oximeter.
Results are documented and submitted to the POCT Lab for review.
If patient comparison results do not match, remove analyzer and cuvettes from use until issue is resolved.

## Slide 17
Avox – Bi-annual linearity verification
Must be completed by user site on each AVOX analyzer:
  Each 6 months
  After analyzer repair
Linearity kit is used and is intended to challenge the full reportable range for both THb and %HbO2.
If linearity verification fails, remove analyzer and cuvettes from use until issue is resolved.
Troubleshoot by contacting POCT lab or manufacture tech support.

## Slide 18
Avox – Post repair implementation
If Avox comes back from Clinical Engineering or tech support for repair/check following must be completed PRIOR to patient use:
Perform both levels of filter QC
Perform both levels of liquid QC
Perform patient sample comparison with the Clinical Lab co-oximeter and with the other AVOX device.
Perform linearity test
All results must be within acceptable limits

## Slide 19
Avox – Sample requirements
ONLY fresh whole blood should be tested on the AVOX
Arterial or venous blood is acceptable
50 ul of fresh whole blood-analyzed within 3 minutes of collection is stated in manufacturer guide
Samples from in-dwelling lines, always adequately “clear” line prior to collecting a sample for AVOX testing
Inaccurate results will be given if samples are contaminated/diluted.

## Slide 20
Avox – Patient testing
All AVOX results should be documented in the permanent/EMR patient record.
Results should include:
  AVOX result with units of measure
  Testing personnel
  Date of testing
  Time of testing
A clear audit trail should be maintained correlating patient results with QC results for the analyzer and cuvettes that were used for patient testing.
Follow hospital policy for verification of patient identity.
Ensure sample identity throughout entire testing and reporting process.
Never force blood into the cuvettes. If the sample does not flow easily, discard the cuvette and use another one.
When filling the cuvette, do not use excessive pressure or cause the vent patch to bulge outward.

## Slide 21
Avox – Patient testing
Review QC logs to ensure instrument is ready for use.
Gloves should be worn during sample collection and analysis.
Pressing the Enter/On key and confirm that the self-test was successful. The instrument indicates that it is ready to analyze samples by displaying the “READY SCREEN”.
Verify that the correct cuvette path length is programmed in the AVOX.
Open the sealed bag to remove cuvettes. Immediately reseal bag. Do not touch the sampling area (the clear surface) of the disposable cuvettes.
Handle cuvettes by black cap.
Evacuate any air bubbles that may be in the sample.
Immediately cap the sample to prevent air contamination and keep tightly sealed.
Roll syringe between palms to keep red blood cells and plasma well mixed.
Squirt out the first few drops of sample to check for clots and to get well mixed sample.

## Slide 22
Avox – Patient testing
Attach the cuvette to tip of the syringe, point the cuvette downward at a 45º angle, and observe the sample chamber of the cuvette. Do not inject the blood into the instrument’s sample chamber.
Press the plunger gently and fill the cuvette. Stop injection of the sample as soon as blood reaches the air vent. Do not allow any air bubbles to adhere to the internal surface of the cuvette.
With syringe still attached to the cuvette, hold the cuvette firmly by its black cap, and insert the cuvette into the sample chamber. Do not depress the plunger while the syringe is in the instrument.
The filled cuvette should be inserted in the AVOX analyzer within 30 seconds of fill.
Observe the LCD display. Do not disturb the AVOX. Results should display within 10 seconds.
Record results in the patient record and remove the sample from the sample port and discard in an appropriate biohazard container.

## Slide 23
Avox – Patient results
All AVOX results should be documented in the permanent/EMR patient record.
Results should include:
  AVOX result with units of measure
  Testing personnel
  Date of testing
  Time of testing
A clear audit trail should be maintained correlating patient results with QC results for the analyzer and cuvettes that were used for patient testing.
Unexpected results should be verified by another test method.
THb results <4 or >25 will give an error and results outside of this range should be reported as “Out of Instrument Range.”

## Slide 24
Avox – Limitations
Following factors can affect result reliability:
Cuvettes exposed to moisture (desiccant indicator is pink)
Over filled cuvettes
Clotted samples
Unmixed samples
Lipid particles
Delay inserting cuvettes into the analyzer
Air bubbles
Excess heparin or saline, diluted samples
Certain medications
For more details refer to manufacture manual.
Any concerns or questions should be reported to the site manager and to the POCT Lab coordinator.

## Slide 25
Thank you for completing Avox education.
Please note this dose not replace the Avox procedure.
Point of Care Testing Department