| Influenza Antigen Rapid Detection |  |
| --- | --- |
| Purpose | This procedure provides instructions for performing the Sofia Influenza A+B FIA assay. |
| Policy Statements |  Viral rapid antigen testing is offered 24/7. - Testing is offered during the influenza season only. The season will be determined by influenza positives detected by PCR or isolation in culture. - A negative test is presumptive and it is recommended these results be confirmed by virus culture or influenza A and B PCR. - Negative results do not preclude influenza virus infections and should not be used as the sole basis for treatment or other management decisions. - If infection with a novel influenza virus is suspected, specimens should be collected and handled with appropriate infection control precautions for novel virulent influenza viruses. Antigen detection and molecular testing can be performed under BSL2 conditions. Viral culture should not be attempted until PCR results are negative. |
| Principle and Clinical Significance | This qualitative test is intended for use as an aid in the rapid differential diagnosis of acute influenza A and influenza B viral infections in symptomatic patients. The Sofia Influenza A+B FIA employs Europium-based fluorescent chemistry that is used with the Sofia Analyzer to detect influenza virus nucleoproteins. This test allows for the differential detection of influenza A and influenza B antigens. The patient specimen is placed in the Reagent Tube, during which time the virus particles in the specimen are disrupted, exposing internal viral nucleoproteins. After disruption, the specimen is dispensed into the Cassette sample well. From the sample well, the specimen migrates through a test strip containing various chemical environments including mouse monoclonal anti-influenza A and anti-influenza B antibodies and micro-particles dyed with fluorescent Europium. If influenza viral antigen is present, they will be trapped in a specific location. The Sofia Analyzer is used to scan the test strip and measures the fluorescent signal by processing the results using method-specific algorithms. Results are displayed on the Sofia Analyzer and are printed on an integrated printer. Influenza, otherwise known as the "flu," is an acute, contagious respiratory illness caused by influenza A, B, and C viruses. Of these, only influenza A and B are thought to cause significant disease, with infections due to influenza B usually being milder than infections with influenza A. Influenza A viruses are further categorized into subtypes based on the 2 major surface protein antigens: hemagglutinin (H) and neuraminidase (N). Common symptoms of influenza infection include fever, chills, sore throat, muscle pains, severe headache, weakness/fatigue, and a nonproductive cough. Certain patients, including infants, the elderly, the immunocompromised, and those with impaired lung function, are at risk for serious complications. In the United States, influenza results in approximately 36,000 deaths and more than 200,000 hospitalizations each year. |
| Test Code | DFLU |
| Special Safety Precautions | Microbiologists/virologists are subject to occupational risks associated with specimen handling. Refer to the safety policies located in the safety section of the Microbiology and Virology Policy Manual: - Biohazard Containment - Safety in the Microbiology/Virology Laboratory - Biohazardous Spills - All specimen processing and vortexing is to be performed in a Biological Safety Cabinet. - Gloves are to be worn during specimen processing and testing. - If infection with a novel influenza virus (i.e. avian, H5, H9, etc.) is suspected, specimens should be collected and handled with appropriate infection control precautions for novel virulent influenza viruses. Antigen detection and molecular assays may be performed under BSL2 conditions. Viral culture is not to be performed until PCR results are negative. - The Reagent Solution contains a salt solution (saline). If the solution contacts the skin or eye, flush with copious amounts of water. |

| Sample Materials | Acceptable specimens: Nasal washes/aspirates ONLY - NW – nasal wash - NASP – nasal aspirate In order to achieve optimal test sensitivity only nasal wash/aspirates are to be used. Unacceptable specimens: NP swabs, nasal swabs, throat swabs, sputum, bronchial washes, ETT or tracheal aspirates. Specimens contaminated with greater than a 4% concentration of blood. Specimen Collection and Transport - Refer to Lab Test Directory on StarNet Specimen Assessment - Refer to the policy Specimen Rejection Criteria. Specimen Storage Test specimens as soon as possible after collection. - nasal wash specimens (not in VTM) at RT : stable for 4 hours nasal wash specimens(not in VTM) at 2-8ºC: stable for 24 hours - Specimens in VTM stored at 2-25C : test within 72 hrs Reagents; Supplies; Equipment; Media.  Sofia Influenza A+B test cassette. - Reagent tube - Reagent Solution ampule Store the kit at room temperature, 59–86°F (15–30°C), out of direct sunlight. Kit contents are stable until the expiration date printed on the outer box. Do not freeze; - 125ul fixed volume pipet from kit - Filter pipet tip - Sterile transfer pipet;  Pipettor to deliver 250ul - Timer - BiologicalSafety Cabinet (BSC); - Viral Transport Media (VTM) (Copan(BD) Universal) M4,M4-RT,M5,M6, and Starplex MultiTrans VTMs are compatible with the Sofia Influenza A+B Assay. |
| --- | --- |

| Calibration | Sofia Calibration Check is required every 30 days. See Sofia Calibration Procedure |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Quality Control | ## Daily Quality Control: - The Sofia Influenza A+B FIA contains a built-in procedural control. Each time a test is run in the Sofia Analyzer, the procedural control zone is scanned by the Sofia Analyzer and the result is displayed on the Analyzer screen. The manufacturer's recommendation for daily control is to document the results of these built-in procedural controls for the first sample tested each day. This documentation is automatically logged into the Analyzer with each test result. A valid result obtained from the procedural control demonstrates that the test flowed correctly and the functional integrity of the Cassette was maintained. The procedural control is interpreted by the Sofia Analyzer after the Cassette has developed for fifteen (15) minutes. If the test does not flow correctly, the Sofia Analyzer will indicate that the result is invalid. Should this occur and the sample is viscous redilute the sample up to 1:4 in VTM and repeat the test with the diluted patient sample and a new test Cassette. Do not report patient test results if the control tests do not produce the expected results. - New Lot/Shipment Quality Control : - Perform QC using external positive and negative controls with each new lot or shipment before putting into service and every 30 days. Record results in the Virology QC manual(Mpls) or th Kit/Reagent QC Log (ST.Paul). - Positive and Negative External Controls are to be run once for each untrained operator. See Sofia Quality Control Procedure |  |  |  |  |  |  |
| Procedure | All clinical specimens must be at room temperature before beginning the assay. Check expiration date on each individual test package or outer box before using. Do not use any test past the expiration date on the label. |  |  |  |  |  |  |
|  | ### Specimen Processing - Vortex nasal wash sample until well mixed (30 seconds). - Place 1-2ml of specimen into labeled VTM. - Vortex for 30 seconds - Transfer ~400ul to a labeled Falcon tube. |  |  |  |  |  |  |
|  | If |  | Then |  |  |  |  |
|  | Nasal wash/aspirate < 0.5ml |  | Remove approximately 1ml of VTM before adding the nasal washing. |  |  |  |  |
|  | Nasal wash/aspirate appears acellular (very clear) |  | Remove approximately 1ml of VTM before adding the nasal washing. Add 2ml of NW to VTM. |  |  |  |  |
|  | Specimen is too viscous to pipet |  | Dilute up to 1:4 in VTM. |  |  |  |  |
|  | **Specimens in VTM can be tested immediately after vortexing. They do NOT need to incubate for 30 minutes prior to rapid antigen testing. (VRSP samples require 30 minute decontamination in VTM prior to cell culture inoculation) Test Procedure - Label a Reagent Tube with the specimen accession number. - Remove cap from the Reagent Tube. - Flick the Reagent Solution ampule and twist off the top. - Add the Reagent Solution to the Reagent Tube. Gently swirl the Reagent Tube to dissolve its contents. - Vortex specimen for 30 seconds and pipet 250 μL of the sample (NW diluted in VTM) into the Reagent Tube. Gently swirl the Reagent Tube to mix its contents. - Remove test cassette from the foil pouch just prior to testing and place specimen accession label on the top end of the cassette. 7. Fill the provided fixed volume pipette (120μL) with the patient sample from the Reagent Tube. To fill the fixed volume pipette with the patient sample: - FIRMLY squeeze the top bulb. - Still squeezing, place the pipette tip into the liquid sample - With the pipette tip still in the liquid sample, release pressure on bulb to fill the pipette. 8. Dispense the patient sample into the Cassette sample well. Firmly squeeze the top bulb to empty the contents of the fixed volume pipette into the cassette sample well (extra liquid in the overflow bulb is OK) NOTE: The fixed volume pipette is designed to collect and dispense the correct amount of liquid sample. Discard the pipette into biohazard red trash. 9. CHOOSE MODE: If using Read Now Mode for batch testing: Remove the test Cassette from the hood and place on the countertop to incubate .Set timer for 15 minutes. Read in the Sofia analyzer after 15 minutes incubation on the bench. -OR- If using Walk Away Mode: Insert into Sofia analyzer and the instrument will time the test development, scan and display the result. See Using the Sofia Analyzer section below. Note: Results remain stable for an additional 15 minutes after the 15 minute incubation time. 10. Refrigerate specimen. Using the Sofia Analyzer - Select the desired development mode: WALK AWAY or READ NOW. - The current development mode will be shown at the top of the Start Test Screen. To change modes: - Go to the Main Menu - Select CHANGE MODE - Use up or down arrows to highlight the desired mode - Select OK The Analyzer will then return to the Start Test screen displaying the selected mode. 1. Press the Down Arrow on the Sofia Analyzer keypad to go to the Patient ID field. Input patient accession number using the barcode scanner or manually enter the data using the key pad. Use the alpha button to enter letters and the numeric keypad for numbers. 2. Press Start Test and the Sofia Analyzer drawer will automatically open. 3. Verify that the correct development mode, Walk Away or Read Now, has been selected. Immediately insert the prepared patient test Cassette into the drawer of the Sofia Analyzer and gently close the drawer. 4. .Upon closing the drawer, the Sofia Analyzer will start automatically and display the progress. In the Walk Away Mode, the test results will be displayed on the screen in approximately fifteen (15) minutes after the Cassette is inserted into the Sofia Analyzer. In the Read Now Mode, the test results will be displayed on the screen within one (1) minute, after the Cassette is inserted into the Analyzer. 5. .When the test is complete, the patient results and the internal procedural control results will be displayed and automatically print. 6. To run another specimen, select Start New Test. |  |  |  |  |  |  |
| Interpretation/ Results | When the test is complete, the results will be displayed on the Sofia Analyzer screen and will automatically print on the integrated printer. When removing the cassette from the analyzer, verify accession# on cassette with accession # on printout. Place a long bar code sticker on the printout tape and transcribe results to log sheet. The Sofia Analyzer screen will display results for the procedural control as being “valid or invalid,” and will provide a positive or negative result for both influenza A and influenza B. If the procedural control is “invalid” and specimen is viscous, redilute the specimen up to 1:4 in VTM and repeat testing with a new cassette. Co-infection with influenza A and B is rare. Sofia Influenza A+B FIA “dual positive” clinical specimens (influenza A and influenza B positive) should be re-tested. Repeatable influenza A and B “dual positive” results should be confirmed by PCR before reporting results. Report as inconclusive. Call result and recommend either a RIPST and/or VRSP on the sample or recollecting a new sample for repeat testing. Invalid Result: Redilute specimen up to 1:4 in VTM depending on how thick the sample is and repeat testing with a new test cassette. If repeat testing is still invalid report as inconclusive. Call result and recommend either a RIPST and/or VRSP on the sample or recollecting a new sample for repeat testing. |  |  |  |  |  |  |
| Dilutions | Specimens are routinely run with a 1:2 dilution in VTM. If the specimen is too viscous to pipet, make up to a 1:4 dilution in VTM for testing. |  |  |  |  |  |  |
| Result Reporting | ### REPORTING: Enter results in Misys MRE Culture Entry tab in Observations by using customized keyboards or by entering a code in the result box. RESULT; KEY; CODE. Flu A Positive; S; PFLU: POSITIVE: Influenza A viral antigen detected. Flu B Positive; D; PFLUB: POSITIVE: Influenza B viral antigen detected. Flu A Negative; F; NFLU: NEGATIVE: No influenza A viral antigen detected. Flu B Negative; G; NFLUB: NEGATIVE: No influenza B viral antigen detected. Inconclusive; ICLR: Inconclusive results. CAL: Called to-;____________________. |
| Method Performance Specifications | Sofia vs. PCR or culture: In House Data 2012-2013 Influenza A – sensitivity 97%, specificity 97%,PPV 88%, NPV 99% Influenza B- sensitivity 94%, specificity 97%, PPV 77%,NPV 99% The limit of detection (LoD) for the Sofia Influenza A+B FIA was determined using a total of four (4) strains of human influenza viruses, two (2) influenza A and two (2) influenza B viruses (Table 9). Limit of Detection with Human Isolates of Influenza A and B Viral Strain; Viral Type; Sub-Type; Minimum Detectable Level (TCID50/mL). A/California/07/2009; A; 2009 H1N1; 202. A/Hong Kong/8/68; A; H1N1; 105. B/Allen/45; B; 40. B/Malaysia/2506/04; B; 24. |
| Procedure Notes | - Sofia instrument tapes and worksheets are retained for two years. - Individuals who received nasally administered influenza A vaccine may have positive test results for up to three days after vaccination. - Samples contaminated with whole blood >4% v/v or mucin >0.5% v/v may interfere in the interpretation of the test. - If differentiation of specific influenza A subtypes and strains is needed a viral respiratory culture or influenza PCR subtyping must be performed. - The test Cassette should not be left standing inside a laminar flow hood or heavily ventilated area during the 15 minute development step or drying of the sample may occur. This can retard the flow of the sample through the test strip which can cause inaccurate or invalid results. - Test undiluted nasal wash/aspirates stored at RT within 4 hours. Test undiluted nasal wash/aspirates stored at 2-8º within 24 hours. |  |  |  |  |  |  |
| Limitations | - Positive and negative predictive values are highly dependent on prevalence. False negative test results are more likely during peak activity when prevalence of disease is high. False positive test results are more likely during periods of low influenza or activity when prevalence is moderate to low. - This test detects both viable (live) and non-viable influenza A and B. Test performance depends on the amount of virus (antigen) in the specimen and may or may not correlate with viral culture results performed on the same specimen. - Positive test results do not rule out co-infections with other pathogens. - Positive test results do not identify specific influenza A virus subtypes. - Negative test results are not intended to rule in other non-influenza viral or bacterial infections. - The test is not intended to detect influenza C antigens. - The clinical performance of the Sofia Influenza A+B FIA for nasopharyngeal aspirate/wash samples has not been established in patients 60 years of age and older and may not be consistent with the clinical performance obtained with younger patients. - A negative test result may occur if the level of antigen in a sample is below the detection limit of the test or if the sample was collected or transported improperly. - Failure to follow the Test Procedure may adversely affect test performance and/or invalidate the test result. - Test results must be evaluated in conjunction with other clinical data available to the physician. - Children tend to shed virus more abundantly and for longer periods of time than adults. Therefore, testing specimens from adults will often yield lower sensitivity than testing specimens from children. - Monoclonal antibodies may fail to detect, or detect with less sensitivity, influenza A viruses that have undergone minor amino acid changes in the target epitope region. - The performance of this test has not been evaluated for use in patients without signs and symptoms of respiratory infection. |  |  |  |  |  |  |
| References | 1. Quidel Corporation, San Diego, CA 92121, Sofia Influenza A+B FIA product Insert. 2.Schweiger B, Lehmann H. Evaluation of a fluorescence-based point-of-care test for detection of influenza viruses circulating in Germany .Poster from Robert Koch-Institut, National Reference Centre for Influenza, Berlin, Germany. 3.Laboratory All Common Checklist, College of American Pathologists, 325 Waukegan Road, Northfield, IL 60093-2750, 09.25.2012 4. Centers for Disease Control and Prevention. Seasonal Influenza: http://www.cdc.gov/flu/ 4/2015. |  |  |  |  |  |  |
| Alternate Methods | RSV, Influenza A and B PCR (RIPST) could be performed if this method becomes unavailable. RIPST is performed on days with.TAT: 12-24 hours BINAX NOW FluA&B test kit may be used as alternate method if the SOFIA test kits become unavailable. |  |  |  |  |  |  |
| Appendix | Appendix A: Bench reference |  |  |  |  |  |  |
| Training Plan/ Competency Assessment | Training Plan |  |  |  | Competency Assessment |  |  |
|  | - Employee must read the procedure. - Employee will observe trainer performing the procedure. - Employee will demonstrate the ability to perform procedure, record results and document corrective action after instruction by the trainer. |  |  |  | ### Direct observation |  |  |
| Historical Record |  |  |  |  |  |  |  |
|  | Version | Written/Revised by: |  | Effective Date: |  | Summary of Revisions |  |
|  | 1 | Helen Stefan |  | 11/15/12 |  | Initial Version |  |

|  | 2 | Helen Stefan | 10/8/13 | Specimens routinely diluted 1:2 in VTM, removed using undiluted specimens ,CAP requirement to perform external QC every 30 days, NRT comment when only rapid tests are ordered, update method performance specifications. |
| --- | --- | --- | --- | --- |

|  | 3 | Helen Stefan | 11/19/13 | Removed additional comments of VNP, CTF, and NRT. |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | 4 | Helen Stefan | 2/11/14 | Repeat collection suggested removed from ICLR code translation |  |  |
|  | 5 | Helen Stefan | 4/13/15 | 260ul specimen volume changed to 250ul per Quidel procedure. |  |  |
|  | 6 | Helen Stefan | 4/27/2015 | Renumbered from V5.13 for CMS, added bench reference as appendix. |  |  |
|  | 7 | Helen Stefan | 10/15/15 | Added where to record STP QC results, incubate test Cassette on countertop not in hood. |  |  |
|  | Archived by: |  | Archived Date: |  |  |  |

Appendix A: Bench Reference Influenza Antigen Detection (Sofia)

For complete procedure see VIRO 2.12 Influenza Antigen Detection (Sofia)

Acceptable specimens: Nasal washes/aspirates ONLY

- NW – nasal wash

- NASP – nasal aspirate

| - Specimen Processing - Vortex nasal wash sample until well mixed (30 seconds). - Place 1-2ml of specimen into labeled VTM. - Vortex for 30 seconds - Transfer ~400ul to a labeled Falcon tube or cryovial. - Proceed to Test Procedure below |  |
| --- | --- |
| If | Then |
| Nasal wash/aspirate < 0.5ml | Remove approximately 1ml of VTM before adding the nasal washing. |
| Nasal wash/aspirate appears acellular (crystal clear) | Remove approximately 1ml of VTM before adding the nasal washing. Add 2ml of NW to VTM. |
| Specimen is too viscous to pipet | Dilute up to 1:4 in VTM. |

**Specimens in VTM can be tested immediately after vortexing. They do NOT need to incubate for 30 minutes prior to rapid antigen testing. (VRSP samples DO require a 30 minute decontamination in VTM prior to cell culture inoculation)

Test Procedure

- Label a Reagent Tube with the specimen accession number.

- Remove cap from the Reagent Tube.

- Flick the Reagent Solution ampule and twist off the top.

- Add the Reagent Solution to the Reagent Tube. Gently swirl the Reagent Tube to dissolve its   contents.

- Vortex specimen for 30 seconds and pipet 250 μL of the sample (NW diluted in VTM) into the Reagent Tube.  Gently swirl the Reagent Tube to mix its contents.

- Remove test cassette from the foil pouch just prior to testing and place specimen accession label on the top end of the cassette.

7.    Fill the provided fixed volume pipette (120μL) with the patient sample from the Reagent Tube.

8.    Dispense the patient sample into the Cassette sample well.

9.   CHOOSE MODE:

If using Read Now Mode for batch testing: Remove the test Cassette from the hood and place on the countertop to incubate.Set timer for 15 minutes.

Read in the Sofia analyzer after 15 minutes incubation on the bench.  Note: Results are stable for an additional 15 minutes after the 15 minute incubation time.

-OR-

If using Walk Away Mode:  Insert into Sofia analyzer immediately and the instrument will time the test development, scan and display the result.

- Refrigerate specimen after testing is completed.

Resulting

When the test is complete, the results will be displayed on the Sofia Analyzer screen and will automatically print on the integrated printer.

- When removing the cassette from the analyzer, verify accession# on cassette with accession # on printout.

- Place a long bar code sticker on the printout tape and transcribe results to log sheet. Record controls as “Valid”.

- Scan label and enter results in Misys MRE Culture Entry tab in Observations line 1 and Observation Line 2 by using customized keyboards or by entering a code in the result box.

| RESULT | KEY | CODE |
| --- | --- | --- |
| Flu A Positive | S | PFLU: POSITIVE: Influenza A viral antigen detected |
| Flu B Positive | D | PFLUB: POSITIVE: Influenza B viral antigen detected |
| Flu A Negative | F | NFLU: NEGATIVE: No influenza A viral antigen detected |
| Flu B Negative | G | NFLUB: NEGATIVE: No influenza B viral antigen detected |
| Inconclusive |  | ICLR: Inconclusive results. CAL: Called to-;____________________ |

- Verify that patient and results are correct on screen and final results. Call completed worksheet to verify result entry.