| ABL 825 Maintenance |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Purpose | This procedure provides instructions for ABL 825 MAINTENANCE. The ABL825 FLEX analyzer is to be properly operated, maintained, serviced and monitored to ensure that malfunctions of the instruments do not adversely affect the analytical results. |  |  |  |  |  |  |
| Policy Statements | - The procedure applies to chemistry staff responsible for the daily operation and verification of the Radiometer ABL825 FLEX Blood Gas Analyzer. - Required Instrument maintenance will be performed each day of patient testing - Manufacturer’s recommended schedules are adhered to - Maintenance activities are documented by the tech performing the function checks and maintained in or near the analyzer - Non-routine maintenance and repair records are kept in the analyzer’s Maintenance Binder - All maintenance records are retained for the life of the instrument, plus 2 years |  |  |  |  |  |  |
| Materials |  |  |  |  |  |  |  |
|  | Supplies |  |  | Equipment |  |  |  |
|  | - Lint free tissue - Tap water or Millipore water - Disposable disinfectant wipes |  |  | - ABL825 FLEX Blood Gas System Serial Number: 754R0427N005 “ABL 1” in Minneapolis - ABL825 FLEX Blood Gas System Serial Number: 754R0427N010 “ABL 2” in Minneapolis - ABL825 FLEX Blood Gas System Serial Number: 754R0427N006 “ABL 4” in St. Paul - ABL825 FLEX Blood Gas System Serial Number: 754R0427N007 “ABL 5” in St. Paul |  |  |  |
| Special Safety Precautions | - Refer to Laboratory Safety policy manual - Dispose of ABL 825 filled Waste Containers in Biohazardous Waste - Dispose of empty reagent containers in regular trash - Transport emptied gas cylinders to Children’s Flammable Storage Room/Hazardous Waste Storage Room |  |  |  |  |  |  |
| Records/Forms/Documents | Linearity – Calibration - Reportable Range Information Packet, Part NO. 999-719, Radiometer America Inc. |  |  |  |  |  |  |
| Procedures: |  |  |  |  |  |  |  |
| Daily Maintenance | Step | Task |  |  |  | Action |  |
|  | 1 | Check Solution Levels |  |  |  | Replace as needed. Refer to ABL Operator’s Guide section 7: Replacements |  |
|  | 2 | Check/Replace Gas |  |  |  | Refer to ABL Operator’s Guide section 7: Replacements |  |
|  | 3 | Check Paper Supply |  |  |  | Replace as needed. Refer to ABL Operator’s Guide section 7: Replacements |  |
|  | 4 | Check Autocheck QC Supply |  |  |  | Refill carousel as needed. Refer to ABL Operator’s Guide section 7 |  |
|  | 5 | Clean analyzer Touchscreen |  |  |  | Wipe clean using a lint free tissue moistened with DI water |  |
|  | 6 | Clean instrument exterior surface |  |  |  |  |  |
|  | 7 | Clean inlet gaskets. |  |  |  | Refer to ABL Operator’s Guide section 7: Replacements |  |
|  | 8 | Bleach workstation. |  |  |  | Wipe up spills using a paper towel moistened with dilute bleach solution, or a disposable disinfectant wipe |  |
|  | 9 | Document each activity as completed. |  |  |  | Analyzer Status -> Electrodes and Other -> Touch the correct Activity under User Activities -> Log Activity -> touch “Done” |  |
| Weekly Maintenance | Step | Task |  |  |  | Action |  |
|  | 1 | Run Protein Remover |  |  |  | Refer to ABL Operator’s Guide section 7: Replacements |  |
| Monthly Maintenance | Step | Task |  |  |  | Action |  |
|  | 1 | Enter the Recording Replacement Menu either by opening the lid to the electrode compartment, or via Electrodes and Other |  |  |  | After completing replacement tasks listed below, select the recorded replacement in the menu to the left. If nothing was replaced select “No Replacements done” in order to Restart the instrument. |  |
|  | 2 | Replace Reference Membrane |  |  |  | Refer to ABL Operator’s Guide section 7: Replacements |  |
|  | 3 | Replace Glu and Lac Membranes |  |  |  | Refer to ABL Operator’s Guide section 7: Replacements |  |
|  | 4 | Perform Decontamination |  |  |  | Refer to ABL Operator’s Guide section 7: Replacements |  |
|  | 5 | Clean Fan Filter |  |  |  | Refer to ABL Operator’s Guide section 7: Replacements |  |
|  | 6 | Perform Instrument Correlations |  |  |  | Refer to CH 2.05 Instrument Correlation procedure |  |
| Quarterly Maintenance | Step | Task |  |  |  | Action |  |
|  | 1 | Replace PCO2 / PO2 Membranes |  |  |  | Refer to ABL Operator’s Guide section 7: Replacements When replacing the pCO2 membrane these steps must be taken to prevent an erroneously low bias: Wait for the analyzer to complete the startup calibration plus two extra automatic calibrations. This will take approximately one hour. After the third automatic calibration perform one level 1 quality control determination. If the Quality Control is out of range, a new membrane must be installed. Do not report patient results until the newly installed membrane has been verified by acceptable quality control. |  |
|  | 2 | Replace Electrolyte Membranes |  |  |  | Refer to ABL Operator’s Guide section 7: Replacements |  |
|  | 3 | Perform tHb Calibration |  |  |  | Refer to ABL Operator’s Guide Section 6: Calibrations, tHb calibration |  |
| Half Yearly | Step | Task |  |  |  | Action |  |
|  | 1 | Replace Pump Tubing |  |  |  | Refer to ABL Operator’s Guide section 7: Replacements |  |
|  | 2 | Replace Fan Filter |  |  |  | Refer to ABL Operator’s Guide section 7: Replacements |  |
|  | 3 | Run Linearity VK kit |  |  |  | Refer to ABL Operator’s Guide section 6: Radiometer settings, Calibration verification |  |
| Yearly | Step | Task |  |  |  | Action |  |
|  | 1 | Annual Service PM |  |  |  | Contact Radiometer Service, included in warranty |  |
| Limitations | - Radiometer Tech line number during Business hours: 1.800.736.0600 - Radiometer Tech line number after Business hours: 1.800.800.7005 - Children’s Service rep’s business card is posted on the side of each instrument |  |  |  |  |  |  |
| References | - ABL800 FLEX Operator's Manual, version 6.14, Publication: 201410, Edition: E, Code number: 994-909 - Radiometer Product Notification dated 7/22/2013 |  |  |  |  |  |  |
| Historical Record | Version |  | Written/Revised by: |  | Effective Date: |  | Summary of Revisions |
|  |  |  | L. Lichty |  | May 11, 2009 |  | Initial Version |
|  |  |  | L. Lichty/D. Helfinstine |  | April 1, 2011 |  | New Format. Renumbered from CH 4.042 |
|  |  |  | Linda Lichty |  | August 1, 2013 |  | Revised pCO2 membrane replacement |
|  |  |  | Linda Lichty |  | April 18, 2016 |  | Update Operator’s Manual for v6.14 |