| Out of Control Plan |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Purpose | This procedure provides instructions for actions to take if reagent quality control results or daily equipment function checks do not conform to established limits. |  |  |  |  |
| Policy Statements | - Any daily reagent quality control testing that does not meet the established parameters as defined in TS. 18.2 Performing Daily Reagent Quality Control must be investigated and resolved prior to patient testing with the implicated reagents. - Any daily equipment function check that does not fall within established limits according to TS 18.3 Performing Daily equipment Function Verification must be investigated and resolved or backup equipment shall be used. - Follow-up investigation of reagent or equipment failures and malfunctions shall include an assessment of the effect on patient safety related to patient testing results and/or blood and blood component selection and preparation processes. |  |  |  |  |
| Procedure |  |  |  |  |  |
|  | Step | Action |  |  |  |
|  | 1 | Reagent out-of-control: If; Then. Reagent appears turbid, hemolyzed, clumped or has expired; Replace with new reagent and repeat testing.. Reagent appears OK; Repeat testing with same reagent.. Gel card appears abnormal or has expired; Replace with new gel card or lot number and repeat testing.. Testing is still out-of-control after repeat testing; - Use alternative test method if available. (E.g. tube vs. gel testing) or - Refer testing to other Children’s campus or blood center reference lab. - Quarantine or discard implicated reagents.. |
|  | 2 | Equipment out of control: If; Then. Temperature is out of control; - Check that the equipment or storage device is on. - Check thermometer integrity, replace thermometer as needed. - Adjust equipment temperature setting as needed allowing temperature to stabilize before taking an additional reading.. Other type of function verification failures; - Repeat function verification assessment - Refer to equipment manual for troubleshooting options. - Refer to TS Equipment Section for additional equipment maintenance options.. Equipment function out of control; - Use backup equipment (refer to Equipment section) or alternative testing methods. or - Refer testing or blood product preparation to other Children’s campus or blood center. - Remove implicated equipment from service. - Send mailbox to staff about removal of equipment from service - Place TSja 18.5.1 Equipment Removed from Service on piece of equipment. - Notify Children’s BioMed as needed for repairs or equipment decommission. - Document equipment malfunction on the TSf 17.10.3 Equipment Maintenance Service Record. Low level Alarm Activated on Liquid nitrogen freezer; - Security will call Blood bank for a low level alarm has been activated. - Monitor temp on the nitrogen freezer. - Place work order through St. Croix for facilities to replace nitrogen tank. Note that it is an urgent request but can wait for dayshift.. |
|  | 3 | Assess if any previously tested samples require retesting or if any blood products (issued or in inventory) may be affected. - Notify the transfusion service technical specialist or pathologist if needing assistance in this assessment. - Complete a Safety Learning Report if erroneous testing results or non-conforming products were released. - Document on TSf 17.10.3 Equipment Maintenance Service Record |  |  |  |
|  | 4 | Leave notification for the transfusion service technical specialist if: - Unable to resolve the issue. - Large volumes of reagent require quarantine or disposal. (Include reagent name, lot number, expiration date, amount) - Equipment is taken out of service. |  |  |  |
|  | 5 | Re-validate equipment per technical specialist instructions prior to use. - Document on TSf 17.10.3 Equipment Maintenance Service Record |  |  |  |
| References | AABB Standards for Blood Banks and Transfusion Service, current edition |  |  |  |  |
| Approval Workflow | Transfusion Service/Laboratory Director |  |  |  |  |
| Historical Record | Version |  | Written/Revised by: | Effective Date: | Summary of Revisions |
|  | 1 |  | J Wenzel | 12/28/2009 | Initial Version |
|  | 2 |  | J Wenzel | 4/10/2012 | CMS format |
|  | 3 |  | S. Cassidy | 8/20/13 | Added-steps to ensure equipment is removed from service, documentation on equipment maintenance service log and step 5. |
|  | 4 |  | S. Cassidy | 5/1/2016 | Added steps for low level alarm activation for nitrogen freezer. |