# *Simplexa* 3 M™ Integrated Cycler Troubleshooting Guide

**PURPOSE:**

* This procedure provides instructions for troubleshooting problems based on symptoms including a potential cause with recommendations for solutions.

## DOCUMENTATION/RECORDS

1. GAS Segment Report
2. QC and Equipment Failure Log

**PROCEDURE A:** Follow the steps in the table below for troubleshooting invalid results, warnings and error codes

General Information

| **Activity** | Step | **Action** | **Related Doc** |
| --- | --- | --- | --- |
| **Warning messages and error codes** |  1 | *Simplexa* 3M™ Integrated Cycler errors will display in a dialog box* Text will explain the cause of the problem
* The user will be provided a resolution to follow to resolve the error
 | Simplexa Operator Manual[Appendix B: Troubleshooting](file:///G%3A%5CLAB%5CMolecular%20Biology%5CA.%20Molecular%20Procedure%20Manual%5CMolecular%20Resources%5CSimplexa%20Troubleshooting%20guide%2C%20Appendix%20B.pdf) |
| 2 | Refer to *Simplexa* Operator Manual, *Appendix B: Troubleshooting* for additional recommendations * Problem scenarios with possible causes and corrective actions
* Error list with possible causes and corrective actions
 |
| **Technical support** | 3 | If unable to resolve problem, contact **Focus technical service at** **1-800-838-4548** option #3 |  |
|  | 4 | Technical Service may ask you to generate and send a Service Packet file; see Procedure B for downloading a \*.icz file |  |
| **Instrument return** | 5 | If it is determined that the instrument must be returned for service, decontaminate the Simplexa before shipping; Refer to procedure GAS 008 | [GAS 008](GAS%20008%20Equipment%20and%20Room%20Decontamination.docx) Equipment and Decontamination |
| **Records** | 6 | Document all problems and actions in the QC and Equipment Failure Log |  |
|  | 7 | File a St. Croix report if BD Field Service is scheduled for repair or preventative maintenance |  |

**PROCEDURE B:** Follow the steps to export a service packet to a USB

**Exporting a Service Packet**

| **Activity** | **Step** | **Action** | **Related Doc** |
| --- | --- | --- | --- |
| **Select run** | 1 | Open a run for analysis from the Quick Pick List or using Browse Runs |  |
| **USB** | 2 | Insert USB device |  |
| **Quick Pick** | 3 | Opening a run from the Quick Pick List

|  |  |
| --- | --- |
| Step | Action |
| 1 | Click on run located under **Analyze Completed Runs**; the 10 most-recent runs are displayed (Fig. 1) |
| 2 | If the run is not displayed use **Browse Runs** |
| 3 | Click on **Export** drop down located near the top of the Analyze Screen  |
| 4 | Select **Service** **Packet** from menu |
| 5 | Select storage location **Removable Disk (E:)**  |
| 6 | Click **OK** after export is successful |
| 7 | Email to technical services |

 | [Simplexa Operator Manual](file:///G%3A%5CLAB%5CMolecular%20Biology%5CA.%20Molecular%20Procedure%20Manual%5CMolecular%20Resources%5CSimplexa%20Operator%20Manual%20PI.MOL1101.UD_REV.F.pdf) |
| **Browse Runs** | 4 | Opening a run from Browse Runs

|  |  |
| --- | --- |
| Step | Action |
| 1 | Select **File: Browse runs** or **Browse Runs** from the Quick Pick list |
| 2 | Click on **My Runs** (Fig. 2) |
| 3 | Click on **Completed** |
| 4 | Click on the **Export** button |
| 5 | Select storage location **Removable Disk (E:)** |
| 6 | Click **OK** after export is successful |
| 7 | Email to technical services |

 |  |

 **Figure 1: Quick Pick List**

**Quick Pick: Analyze Completed**

 **Runs**

 **File: Browse Runs**

 **Quick Pick: Browse Runs**

**Figure 2: Browse Runs**

**Completed Run**

**Export button**

**PROCEDURE C:** Follow the steps to reset the barscode scanner

**Barcode Scanner**

| **Activity** | **Step** | **Action** | **Related Doc** |
| --- | --- | --- | --- |
|  | 1 | The red-laser pattern should surround the outer edge of the barcode when reading the barcode |  |
| Scanner not reading | 2 | If the barcode reader does not respond, reset by scanning the image located on page B-5, Troubleshooting guide | Simplexa Operator Manual[Appendix B: Troubleshooting](file:///G%3A%5CLAB%5CMolecular%20Biology%5CA.%20Molecular%20Procedure%20Manual%5CMolecular%20Resources%5CSimplexa%20Troubleshooting%20guide%2C%20Appendix%20B.pdf) |
|  | 3 | 4 long beeps indicate that data was not transmitted from the scanner to the computer |  |
|  | 4 | If the scanner continues to beep, disconnect the scanner for 5 sec and then reconnect |  |
|  | 5 | Contact **Focus technical service at** **1-800-838-4548** option #3 if error continues |  |

**REFERENCE**

1. Simplexa™ 3M™ Integrated Cycler Studio 5.0 , 3M™ Integrated Cycler Operator Manual Reference 34-8710-8382-9, PI.MOL1101.UD\_REV. F for use with user defined assays, Focus Diagnostics 2009-2012, Focus Diagnostics, Inc. Cypress, CA

|  |
| --- |
| **Approval** |
|  | Approved by | **Signature** | **Date** |
|  | Timothy Casey, MD |  |  |
|  | Patricia Ackerman, Technical Specialist | P. Ackerman, TS | 9.9.14 |
| **Bi-Annual Review** |
|  | Reviewed by | **Signature** | **Date** | Reviewed by | **Signature** | **Date** |
| P. Ackerman | PA | 9.9.14 |  |  |  |
| Historical Record |  |
|  | **Version** | **Written/Revised by:** | **Effective Date:** | **Summary of Revisions** |
| 1 | P. Ackerman | 6/1/2011 | Initial Version |
|  |  |  |  |
| **Distribution** |  |
|  | **Location** | **# Copies** | **Location** | **# Copies** |
|  | Molecular Diagnostics rm B422 | 1 | G: \Molecular Biology\Molecular Proc Manual\MB005.2\GAS 006 | 1 |