| FA Control Slide Preparation |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Purpose | This procedure provides directions for the preparation of quality control slides for use in fluorescent antibody (FA) staining procedures to monitor the performance of test procedures and reagents. |  |  |  |  |  |  |  |  |  |
| Policy Statements | - Positive and negative controls are run with each test run to be in compliance with CLIA regulation (Title 42 in the CFR) 493.1256 and CAP MIC. 21570 and .61370. - Control results must be acceptable before patient results are released from the laboratory |  |  |  |  |  |  |  |  |  |
| Materials |  |  |  |  |  |  |  |  |  |  |
|  | Reagents |  | Supplies |  |  | Equipment |  | Media |  |  |
|  |  Phosphate buffered saline (PBS), pH 7.4 + 0.2 |  | - Adjustable pipettor - Pipet tips, sterile - 1ml sterile pipet - 9” Pasteur pipets, sterile - 10x75 Falcon Tubes - 8 well Teflon coated microscope slides |  |  | - Vacuum aspirator - Inverted light microscope - Serofuge |  | - Cell culture tubes of appropriate cell line to support the desired virus OR RMIX plates |  |  |
| Sample Special Safety Precautions | Stock virus Or Virus controls titered for RMIX QC Microbiologists/virologists are subject to occupational risks associated with specimen handling. Refer to the safety policies located in the safety section of the Microbiology and Virology Policy Manual: - Biohazard Containment - Safety in the Microbiology/Virology Laboratory - Biohazardous Spills |  |  |  |  |  |  |  |  |  |
| Procedure | ### Respiratory Virus QC slides may be prepared by inoculating tube cell culture or using RMIX wells. ### PIV 1, 2, 3 Influenza A, B QC Slide preparation using tube cell culture: - Inoculate 2 serum free RMK’s with 0.2 ml each of the desired stock virus. - Incubate until tube exhibits 1-2+ CPE (~three days). - Perform hemadsorption (HAd) on one tube to check for growth. - If HAd positive, pass the mate to 3 serum free RMK’s. - After 2-3 days, perform HAd on one tube for each virus type. HAd should be 3-4+ (if it is not reincubate tubes for one additional day and then repeat HAd). Toss tube. - Scrape the monolayer off the side of the other culture tubes using a sterile 1 ml pipette. Vortex well and transfer to a 10 X 75 falcon tube. (Multiple tubes of one virus type can be combined in one falcon tube.) Wash one time with PBS. Resuspend cell button in 1-2 ml PBS. - Scrape cells from two uninoculated RMK’s and wash as in step 6 and mix these cells with the positive cells to dilute. - Scrape 3 uninoculated RMK’s and wash as in step 6 for use in negative control wells. - Label slides appropriately for each virus, positive and negative wells.(Example: PIV1 +/- 1/5) - Apply one drop washed cells (approximately 15-30ul) to the appropriate wells (+ or -) and check under inverted scope for adequate volume of cellular material. Adjust density as needed. Allow to dry. - Stain one slide to evaluate cell density and infectivity. #### RSV QC Slides preparation using tube cell culture: - Inoculate 1 HEp-2 tube with 0.2 ml of frozen stock RSV. - Pass to 2 new HEp- 2 tubes when tube exhibits 2-3+CPE. - Scrape monolayer from each tube wall with 1 ml pipette when tubes exhibit 2-3+ CPE. Vortex. Remove media and scraped cells and transfer to a 10 X 75 falcon tube. The contents of all the HEp-2 tubes can be combined in one falcon tube. - Wash one time with PBS. - Scrape and wash 2 uninoculated HEp-2 tubes in PBS one time for use in negative wells. - Label slides appropriately for positive and negative wells ( Example: RSV 1/5) - Apply one drop washed cells (approximately 15-30ul) to the appropriate well (+ or -) and check under inverted scope for adequate volume of cellular materials. Adjust density as needed. Let dry. - Stain one slide to evaluate cell density and infectivity. -OR- PIV 1, 2, 3 Influenza A, B ,RSV and Adenovirus QC Slide preparation using RMIX plates: #### Inoculate the appropriate stock virus into up to 5 wells following the R-Mix Inoculation Procedure in the Viral Respiratory Culture procedure. Incubate overnight at 35°C. - Remove supernatant from the positive RMIX wells and a corresponding number of uninoculated wells for the negative control. - Add 0.2 ml PBS to each well. Scrape cells into PBS using a sterile 1ml pipet and transfer to a labeled 10x75 falcon tube. - Label slides .(Example: PIV1 +/- 1/5) - Vortex cell suspension well and spot ~15ul of cell suspension to the appropriate wells (+ or -). Let dry. - Stain one slide to evaluate cell density and infectivity. - HSV QC Slides - Inoculate 1 HEp-2 tubes with 0.2 ml each frozen stock HSV 1 and HSV 2. - For each type, pass to 1 new HEp- 2 tube when tube exhibits 2-3+CPE. - Scrape monolayer from each tube wall with 1 ml pipette when tubes exhibit 2-3+ CPE. Vortex. Remove media and scraped cells and transfer to a 10 X 75 falcon tube for HSV1 and another 10x75 tube for HSV 2. - Wash one time with PBS. - Label slide wells 1/5 for HSV 1 cells and wells 2/6 for HSV 2. When staining, HSV1 antibody is stained in the top row and HSV 2 antibody is stained in the bottom so well 5 and 2 serve as negative control wells. - Apply one drop washed cells (approximately 15-30ul) to the appropriate well and check under inverted scope for adequate volume of cellular material. Adjust density as needed. Let dry. - Stain one slide to evaluate cell density and infectivity. - VZV, CMV QC Slides - Inoculate 1 SF tube with 0.2 ml frozen stock virus. - Incubate until tube exhibits 2+ CPE. - Trypsin the positive tube and an uninoculated tube for a negative control. After incubation step, add 2ml PBS to tube and vortex. Remove PBS-cell mixture and place in falcon tube. Centrifuge 1 minute to wash. Resuspend cell button in 1-2 ml PBS. - Label slides for VZV and CMV.( Example: VZV 1/5) - Apply one drop washed cells (approximately 15-30ul) to each well (+ and -) and check under inverted scope for adequate volume of cellular material. Adjust density as needed. Let dry. - Stain one slide to evaluate cell density and infectivity. - Store all slides in labeled slide boxes in the virology -70⁰ freezer. |  |  |  |  |  |  |  |  |  |
| Training Plan/ Competency Assessment | Training Plan |  |  |  | Initial Competency Assessment |  |  |  |  |  |
|  | - Employee must read the procedure - Employee will observe trainer performing the procedure. - Employee will demonstrate the ability to perform procedure, record results and document corrective action after instruction by the trainer. |  |  |  | - Direct observation |  |  |  |  |  |
| Historical Record |  |  |  |  |  |  |  |  |  |  |
|  | Version | Written/Revised by: |  | Effective Date: |  |  | Summary of Revisions |  |  |  |
|  | 1 | Char Anderson |  | March 1994 |  |  | Initial Version |  |  |  |
|  | 1.1 | Char Anderson |  | 07/17/2007 |  |  | Original put into department procedure format |  |  |  |
|  | 1.2 | Sonja Franko |  | 05/19/2010 |  |  | Updated into online format |  |  |  |
|  | 3 | Helen Stefan |  | 3/16/15 |  |  | Reformatted, changed version to whole number, added RMIX procedure, removed slide colors, reduced number of tubes to inoculate. Revised Title |  |  |  |
|  | 4 | Helen Stefan |  | 6/17/16 |  |  | Added CLIA and CAP regulations, updated logo |  |  |  |
|  | Archived by: |  |  | Archived Date: |  |  |  |  |  |  |