| ABO Grouping– Tube Testing |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Purpose | This procedure provides instructions for the tube testing of patient for A and B red cell antigens, and corresponding ABO group antibodies for the determination of ABO grouping. |  |  |  |  |  |  |  |  |  |  |
| Policy Statements | - Patients > 4 months old: ABO group shall be determined by testing the patient’s red cells with anti-A and anti-B reagents and by testing the patient’s plasma/serum for expected antibodies with A1 and B reagent cells. If a discrepancy is found between the patient’s forward and reverse typing on patient’s > 4 months old, additional testing shall be performed. Group O red cells must be selected for transfusion until the discrepancy is resolved or in the event the ABO typing is interpreted as inconclusive. - Patients < 4 months old: ABO group shall be determined by testing the patient’s red cells with anti-A and anti-B reagent. Testing the patient’s plasma/serum for antibodies with A1 or B cells shall not be performed. - Seraclone Control ABO+Rh must be run on samples testing positive with Anti-A, Anti-B and Ant-D on the patient’s first typing and as part of the first typing ABO/Rh recheck. - ABO/RH Rechecks on specimens drawn by electronic identification verification system, <4 months and/or type as group O: - Any patient that has not had a previous ABO performed by the same campus must have the ABO confirmed by a second technologist. Group O, Rh negative red cells and group AB platelets or plasma shall be selected for transfusion until confirmation of the patient’s ABO has been completed by a second technologist. - Confirmatory testing may be performed using the same sample but with a new cell suspension. - An Rh control shall be performed as part of the ABO recheck on new AB, Rh Positive patients. d. Resolution of ABO front and back typing discrepancies shall be performed as part of the recheck as needed. d. Reverse typing (A, and B cells) testing may be omitted and resulted as NT on the recheck if QNS on patient plasma. - Both technologists shall perform discrepancies resolution testing. - ABO/RH Rechecks on patients not drawn by an electronic identification verification system and are >4 months and type as non-O. - A second independent sample needs to be tested to confirm blood type. Either a previous sample can be used or a new specimen needs drawn for an ABO/Rh. - Follow the policy statements above for patients that are collected by an electronic identification system - Any patient who is being evaluated for solid organ donation cannot have A1 subgroup typing performed if they have been transfused within the last three months. - Students may only perform ABO/Rh testing on patients with a minimum of two ABO/Rh tests on record. |  |  |  |  |  |  |  |  |  |  |
| Test Codes | ABO/Rh ABO Only ABR-ABO and Rh as part of other testing battery (Type and Screen, Newborn Workup, etc) ARC-ABO and Rh Recheck |  |  |  |  |  |  |  |  |  |  |
| Related Documents | TS 4.1 Making a 3-5% Cell Suspension TS 4.8 Grading and Interpretation of Tube Reactions TS 4.35 Resolving an ABO Discrepancy |  |  |  |  |  |  |  |  |  |  |
| Materials | Equipment |  |  |  | Reagents |  |  |  |  | Supplies |  |
|  | - Centrifuge - Agglutination Viewer |  |  |  | - BioRad Anti-A - BioRad Anti-B - Biotestcell A1 and B - Seroclone Control ABO+Rh |  |  |  |  | - 10 x 75 mm test tubes - BB pipettes - Saline - Marker |  |
| Sample | Fresh patient samples of EDTA or clotted whole blood collected following general blood collection procedures are acceptable. See Collection of Patient Specimens. Citrated samples from donor unit segments or pilot tubes. The specimen should be tested as soon as possible after collection. If testing is delayed, the EDTA or clotted specimen should be stored at 2-6C and may be tested within 10 days from collection. Donor blood may be tested until the products expiration date. Specimens exhibiting gross hemolysis or contamination should not be used. |  |  |  |  |  |  |  |  |  |  |
| Quality Control | Refer to TS 18.2 Performing Daily Reagent Quality Control Reagents must be evaluated each day of use with appropriate controls. |  |  |  |  |  |  |  |  |  |  |
| Before You Begin | - Confirm sample acceptability and review patient history per procedure. - Label tubes per TS 4.6 Labeling Tubes. |  |  |  |  |  |  |  |  |  |  |
| Procedure |  |  |  |  |  |  |  |  |  |  |  |
|  | Step | Action |  |  |  |  |  |  |  |  |  |
|  | 1 | Prepare a 3-5% cell suspension of patient red cells in isotonic saline. |  |  |  |  |  |  |  |  |  |
|  | 2 | Resuspend the reagent red cells by gently inverting the vials several times. |  |  |  |  |  |  |  |  |  |
|  | 3 | Pipette 2 drops of the patient plasma/serum or 1 drop reagent antisera to the corresponding labelled tube as specified in Appendix A. |  |  |  |  |  |  |  |  |  |
|  | 4 | Add 1 drop of 3-5% suspension of the patient red cells or 1 drop of reagent red cells to tubes to the corresponding labelled tube as specified in Appendix A. |  |  |  |  |  |  |  |  |  |
|  | 5 | Mix all tubes. |  |  |  |  |  |  |  |  |  |
|  | 6 | Centrifuge for the posted optimal time in a calibrated serologic centrifuge. |  |  |  |  |  |  |  |  |  |
|  | 7 | Remove the tubes from the centrifuge and confirm that the tube ID with the request and on the computer screen or downtime worksheet. |  |  |  |  |  |  |  |  |  |
|  | 8 | Gently resuspend the cell button and examine macroscopically for hemolysis and agglutination immediately after centrifuging. Note: Positive reactions with Anti-A and Anti-B characteristically show 3+ to 4+ agglutination. Note: Incubation at RT for 20 minutes may be performed to enhance weak reactions. |  |  |  |  |  |  |  |  |  |
|  | 9 | Immediately record the graded reactions in the computer or on the downtime worksheet. |  |  |  |  |  |  |  |  |  |
|  | 10 | Use the table under Interpretation to determine the ABO. If; Then. there are any discrepancies in reaction pattern when compared to the table; - Resolve the discrepancy, per TS 4.35 Resolving an ABO Discrepancy - IF the discrepancy in the reaction can not be resolved the ABO must be interpreted as INCONLUSIVE. an infant < 4 months old types as an O Rh negative and may be a transfer from another facility. Contact the patient care unit and determine if the patient was transported in from another facility; - Contact the other facility BB regarding the infant’s original blood type. - If a discrepancy is noted between infants current blood type and the initial blood type, transfuse the infant with group O Rh Neg. red cells, AB FFP/Cryo, and AB Rh Neg. platelets - Add problem comment PRTO to the patient’s BAD file.. |
| *New AB, Rh Positive patients or new AB patients for ABO Only | 11 | Test the patient's red cell suspension with a Seraclone Control ABO+Rh - Place 1 drop of the control reagent into new tube labelled as RHC. - Add 1 drop of the patient's 3% cell suspension to the RHC tube - Spin the tube in centrifuge for the time indicated. - Gently resuspend cells completely and examine for agglutination. - Record the results in the computer or on the downtime worksheet If the RHC reaction is; Then ABO/Rh result is. Negative; Valid. Positive; Invalid-Forward to Reference lab. |
|  | 12 | Compare the current ABO results with any previous ABO results. If a previous record; And the current and previous results; Then. Exists; Agree; Finalize the results in the computer or on the worksheet. Do not agree; Resolve the discrepancy.. Does not exist and patient is collected by electronic identification verification system, <4 months of age and/or types as O; N/A; Finalize results. Add test ARC (ABO /Rh recheck) to the order if needed and forward sample for a second ABO/Rh by a second technologist. Does not exist and is not collected by electronic identification verification, >4 months of age, and/or types as Non group O; N/A; - Order an ABO/Rh (ABRH) as no charge. - Check to see if a previous sample was drawn to perform testing. (e.g. CBC) - If no previous sample, call patient’s nurse and ask if they want lab to draw new specimen or are they going to draw specimen. - Perform ABO/Rh and have second tech perform ARC. |
|  | 13 | Review the final record including a final clerical check of sample, label, request, interpretation. |  |  |  |  |  |  |  |  |  |
|  | 14 | Dispose of all tubes and pipettes used for the examinations in a biohazard waste container. |  |  |  |  |  |  |  |  |  |
| Interpretation |  |  |  |  |  |  |  |  |  |  |  |
|  | If the forward grouping reaction of patient cells with |  |  |  |  |  | And the reverse grouping reaction of patient plasma/serum tested with |  |  |  | Then interpret the ABO as |
|  | Anti-A is |  |  | Anti-B is |  |  | A1 cells is | B cells is |  |  |  |
|  | 0 |  |  | 0 |  |  | +/H or NT | +/H or NT |  |  | O |
|  | + |  |  | 0 |  |  | 0 or NT | +/H or NT |  |  | A |
|  | 0 |  |  | + |  |  | +/H or NT | 0 or NT |  |  | B |
|  | + |  |  | + |  |  | 0 or NT | 0 or NT |  |  | AB |
|  | 0 = No agglutination + = Graded agglutination H = Hemolysis NT= Not Tested All other reaction patterns must be interpreted as Inconclusive unless resolved. |  |  |  |  |  |  |  |  |  |  |
| Limitations | Tubes should be read immediately following centrifugation and delays may cause a dissociation of antigen-antibody complexes resulting in false negative or weak positive reactions. Cold agglutinins, a positive direct antiglobulin test, or rouleaux may cause false positive reactions. |  |  |  |  |  |  |  |  |  |  |
| Result Reporting | TS 5.6 Entering Results for ABO/Rh testing or for ABO/Rh Recheck TS 5.7 Entering Results for ABO Grouping Only |  |  |  |  |  |  |  |  |  |  |
| References | Product Insert, Blood Grouping Reagent, Anti-A, Anit-B, Bio-Rad Medical Diagnostics, current edition Product Insert, Biotestcell A1 and B, Bio-Rad Medical Diagnostics, current edition Product Insert, Blood Grouping Reagent, Seraclone Control ABO+RH, Bio-Rad Medical Diagnostics, current edition |  |  |  |  |  |  |  |  |  |  |
| Appendices | Appendix A: ABO Tube Labeling and Set-up |  |  |  |  |  |  |  |  |  |  |
| Approval Workflow | Transfusion Service/Laboratory Director |  |  |  |  |  |  |  |  |  |  |
| Historical Record | Version |  | Written/Revised by: |  |  | Effective Date: |  |  | Summary of Revisions |  |  |
|  | 1 |  | Hartley |  |  | 1982 |  |  | Initial Version |  |  |
|  | 2 |  | Berglund |  |  | 1985 |  |  |  |  |  |
|  | 3 |  | McGee |  |  | 1990 |  |  |  |  |  |
|  | 4 |  | ### Hansen |  |  | 1995 |  |  | Merger |  |  |
|  | 5 |  | Wenzel |  |  | 12/1996 |  |  |  |  |  |
|  | 6 |  | Wenzel |  |  | 9/1998 |  |  |  |  |  |
|  | 7 |  | Wenzel |  |  | 5/2001 |  |  |  |  |  |
|  | 8 |  | Wenzel |  |  | 6/5/2003 |  |  |  |  |  |
|  | 9 |  | Wenzel/Cassidy |  |  | 7/30/07 |  |  | ABO Discrepancy added |  |  |
|  | 10 |  | Wenzel |  |  | 3/04/08 |  |  | New format |  |  |
|  | 11 |  | J. Wenzel |  |  | 8/22/2011 |  |  | Bio Rad reagents. Limitations added. Change from 14 to 10 days specimen requirement |  |  |
|  | 12 |  | S. Cassidy |  |  | 11/18/15 |  |  | Added policy statements around second confirmatory testing of ABO |  |  |
|  | 13 |  | S. Cassidy |  |  | 08/22/16 |  |  | Added policy statement regarding organ donation and subtyping. |  |  |

Appendix A:	ABO Tube Labeling and Set-Up

| Tube labeled as: | Tests for the presence of : | Contents: |
| --- | --- | --- |
| Patient’s name/initials | N/A-test setup | 3-5% patient red cell suspension |
| Patient’s name/initials A | A antigen | 1 drop anti-A antiserum AND 1 drop patient 3-5% red cell suspension |
| Patient’s name/initials B | B antigen | 1 drop anti-B antiserum AND 1 drop 3-5% patient red cell suspension |
| Patient’s name/initials a1 | Patient’s anti-A | 2 drops patient plasma/serum AND 1 drop A1 Cells |
| Patient’s name/initials b | Patient’s anti-B | 2 drops patient plasma/serum AND 1 drop B cells |