| Capillary Blood Gas Collection |  |
| --- | --- |
| Purpose | This procedure provides specific instructions for CAPILLARY BLOOD GAS COLLECTION. |
| Policy Statements | This procedure applies to all laboratory staff responsible for collecting capillary blood gas specimens and must be used in conjunction with SCM 3.20 Capillary Blood Collection. |
| Materials and Equipment | Balanced Heparin plastic capillary tube – Radiometer safeCLINITUBES (D957P-70-125x1) (Order code 942-893), CHC# 9006 |
| Sample | Whole Blood: - Two 125 uL heparinized safeCLINITUBE (preferred) - One 125 uL heparinized safeCLINITUBES (absolute minimum) - Collected without air bubbles, or air bubbles removed prior to transport Patient Preparation: The patient should preferably be in a steady state of ventilation before and during the collection of blood sample Minimum volume: 0.125 mL Stability: Room Temperature; 10 minutes (preferred) or 30 minutes maximum Rejection criteria: - Unlabelled specimens - Clotted specimens - Air contaminated specimens |
| Special Safety Precautions | Authorized personnel should perform blood sampling. Always handle blood samples and capillaries with care. Avoid direct contact with the sample by using certified protective gloves. Do not re-use sampling devices. |
| Procedure | Follow the activities in the table below for CAPILLARY BLOOD GAS COLLECTION. Step; Action. 1; See SCM 3.20 Capillary Specimen Collection. Follow instructions until ready to collect the blood gas sample, go to Step 2.. 2; Loosely mount a capillary cap on one end of the safeCLINITUBE.. 3; Assure the safeCLINITUBE contains a mixing wire. Allow it to slide to the same end as the cap.. 4; Fill the safeCLINITUBE with freely forming drops, holding the puncture site in a downward position and allowing capillary action to pull sample into the tube. Do not allow air to enter the capillary tube.. 5; Tighten the capillary cap.. 6; Tightly mount a cap on the other end of the safeCLINITUBE.. 7; Move the mixing wire 20 times along the full length of the capillary using the magnet to mix the blood and anticoagulant together.. 8; Place one capillary tube under each Sunquest barcode label, keeping the Sunquest cap gas labels attached to each other, and assuring each sample has its own label.. 9; Transport sample tubes immediately to main lab either by walking, or placing the cap tubes in a biohazard sample bag and immediately transporting the sample via pneumatic tube.. 10; Analyze the sample within 10 minutes. If storage is unavoidable, store the sample for a maximum of 30 minutes. Before the sample is analyzed, mix it again.. |
| References | - Radiometer safeCLINITUBES Product Insert, Radiometer Medical ApS, Denmark, 989-890. - ABL 800 Flex Operator’s Guide, Radiometer Medical ApS, Denmark. - CLSI. Procedures and Devices for the Collection of Diagnostic Capillary Blood Specimens; Approved Standard – Sixth Edition, H04-06. Clinical and Laboratory Standards Institute, Wayne, PA. 2008. |