| Tissue Or Cell Transfer Technique |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Purpose | This procedure provides instructions for special technique performed on H&E-stained sections via tissue or cell-transfer procedure when specimen availability is limited and additional diagnostic workup is necessary. This transfer technique is proven and effective method for subdividing areas of a single smear or section and transferring the cells and sections to multiple slides in order to perform additional staining. Cellular and staining quality of tissues/smears will not be affected and histochemical, immunohistochemistry and cytochemical procedures can be performed on transferred samples with reliable results. |  |  |  |  |  |  |  |  |
| Scope | Histology Technical Staff |  |  |  |  |  |  |  |  |
| Materials | Supplies |  |  | Reagents |  |  | Equipment |  |  |
|  | - Scalpel blade - Forceps - Superfrost plus slides - Coplin jar - Gauze - Plastic transfer pipette |  |  | - Xylene - Graded alcohols - Mount Quick mounting media purchased through Newcomer Supply |  |  | - Oven- laboratory grade |  |  |
| Procedure |  |  |  |  |  |  |  |  |  |
|  | Step | Action |  |  |  |  |  |  | Related Document |
|  | 1 | Mark the designated tissue area(s) to be the focus of the tissue/cell transfer on underside of the original slide with a permanent marker (solvent-resistant). |  |  |  |  |  |  |  |
|  | 2 | For previously stained sections: - Soak the original slide in xylene until coverslip is easily removed without damage to tissue section. - If the slides were coverslipped with tape, soak slides in acetone until coverslip is easily removed. - Run slide through several changes of xylene to ensure removal of residual mounting media. - DO NOT proceed through alcohol steps. |  |  |  |  |  |  |  |
|  | 3 | For unstained paraffin sections or smears: - Soak slides in several changes of xylene, a minimum of 10 minutes. - DO NOT proceed through alcohol steps. |  |  |  |  |  |  |  |
|  | 4 | Spread Mount-quick mounting media on xylene coated slides with an applicator stick or pipette tip. Use sufficient Mount-Quick to form a meniscus over the entire tissue section or cell preparation. Slide must be xylene coated and not allowed to dry out. |  |  |  |  |  |  |  |
|  | 5 | Place slide in a flat horizontal position in a 60 degree C oven for 90 minutes to 2 hours or in a 37 degree C oven overnight until meniscus hardness. |  |  |  |  |  |  |  |
|  | 6 | Mark the surface of the hardened mounting media meniscus to corresponding areas marked on the underside of the slide with permanent marker to maintain tissue orientation in the transfer process. |  |  |  |  |  |  |  |
|  | 7 | Immerse the slide in a coplin jar of water warmed in a 45 degree C-60 degree C water bath; soak for 90 minutes to 3 hours. |  |  |  |  |  |  |  |
|  | 8 | Slowly pry meniscus off at the edges with a scalpel blade. If the mounting media does not peel off easily, return the slide to warm coplin jar and continue soaking. |  |  |  |  |  |  |  |
|  | 9 | Pre-label, and moisten the charged slides that will receive the transferred segment(s) with water. |  |  |  |  |  |  |  |
|  | 10 | Divide transferred segment into multiple section: once media section is removed from the original slide, cut section with water moistened scalpel blade or surgical scissors into desired segment. |  |  |  |  |  |  |  |
|  | 11 | To transfer one intact segment of tissue; place on the moistened charged slide, meniscus side up. |  |  |  |  |  |  |  |
|  | 12 | Soak a piece of gauze in warm water; place gauze over the tissue section. Using thumb, press gauze flat transferring water to the section, flattening edges and sealing transferred section to slide. Note: Remember the orientation of the correct side up. If a transferred section is inadvertently placed upside down, the specimen will not adhere to the slide and will be lost in further processing steps. |  |  |  |  |  |  |  |
|  | 13 | Place slide(s) in 37 degree C - 60 degree C oven in a flat horizontal position for 1 hour or longer to securely adhere transferred section to slide. |  |  |  |  |  |  |  |
|  | 14 | Soak slide(s) in four changes of xylene for 10 minutes each or longer until all the Mount-Quick mounting media is completely dissolved. |  |  |  |  |  |  |  |
|  | 15 | Re-hydrate through two changes each of 100% and 95% alcohols, 10 dips each. Wash well with distilled water. |  |  |  |  |  |  |  |
| Procedure Notes | - If any residual mounting medium remains on the slide, tissue transfer and staining may be compromised. - The meniscus of Mount-Quick mounting media must be thoroughly hardened on the slide before proceeding. It is crucial that the meniscus hardens but still remains pliable. - To avoid brittleness do not exceed temperatures above 60 degree C during the hardening process. - If Mount-Quick is not completely dissolved from the tissue section, subsequent staining may be compromised. - If the specimen is already on positively charged slides, additional time may be needed for removal from the original slide, and success at transfer may be poor. |  |  |  |  |  |  |  |  |
| References | - Gill, Gary. Cytopreparation: Principles and Practice, Essentials in Cytopathology. New York: springer Science and Business Media, 2013. 393-395 - Kubier, Patty and Rodney Miller. "Tissue Protection Immunohistochemistry." American Journal of Clinical Pathology 117 (2002): 194-198. - Luna, Lee G. Manual of Histologic Staining Methods of the Armed Forces Institute of Pathology. 3rd Ed. New York: Blakiston Division, McGraw-Hill, 1968. 67-69. - Miller, Rodney and Patty Kubier. "Immunohistochemistry on Cytologic Specimens and Previously Stained Slides (When No Paraffin is Available)." The Journal of Histotechnology 25.4 (2002): 251-257. - Sherman ME, D Jimenez-Joseph, MD Gangi and RR Rojas-corona. "Immunostaining of Small Cytologic Specimens. Facilitation with Cell Transfer." Acta Cytologic 38.1 (1994): 18-22. - Weibel, E. "The Preparation of Serial Microscopic Sections in From of Plastic Films." Lecture, Annual meeting of the International Academy of Pathology, Boston, April 1959. - Zu, Youli, Maryann Gangi, and Grace Yang, "Ultrafast Papanicolaou Stain and Cell-Transfer Technique Enhance Cytologic Diagnosis of Hodgkin Lymphoma." Diagnostic Cytopathology 27.5 (2002): 308-311. - Modifications developed by Newcomer Supply Laboratory. |  |  |  |  |  |  |  |  |
| Authorization |  |  |  |  | Signature |  |  |  | Date |
|  | Medical Director |  |  |  |  |  |  |  |  |
| Annual Review | Designee |  | Written/Revised by: |  |  | Effective Date: |  | Summary of Revisions |  |
|  | Histology Supervisor |  | Prabha Chintapalli |  |  | 11/21/2016 |  | Initial Version |  |