| Levetiracetam |  |  |
| --- | --- | --- |
| Purpose | This procedure provides instructions for PERFORMING LEVETIRACETAM PROCEDURE on serum in Children’s Laboratory. The ARK Levetiracetam Assay is a homogeneous enzyme immunoassay intended for the quantitative determination of levetiracetam in human serum or plasma on automated clinical chemistry analyzers. Levetiracetam concentrations can be used as an aid in management of patients treated with levetiracetam. |  |
| Policy Statements | - This procedure applies to all personnel who run the Dimension RxL MAX and RxL. - All components of individual ARK Levetiracetam Assay reagent kits must be used together, and not shared between kits. |  |
| Principle | ARK Levetiracetam Assay is a homogeneous immunoassay based on competition between drug in the specimen and levetiracetam labeled with the enzyme glucose-6-phosphate dehydrogenase (G6PDH) for binding to the antibody reagent. As the latter binds antibody, enzyme activity decreases. In the presence of drug from the specimen, enzyme activity increases and is directly proportional to the drug concentration. Active enzyme converts the coenzyme nicotinamide adenine dinucleotide (NAD) to NADH that is measured spectrophotometrically as a rate of change in absorbance. Endogenous serum G6PDH does not interfere with the results because the coenyzme NAD functions only with the bacterial enzyme used in the assay. |  |
| Clinical Significance | Levetiracetam (KEPPRA®, (S)-α-ethyl-2-oxo-1-pyrrolidine acetamide) is an anti-convulsant drug approved for use as adjunctive therapy in the treatment of epilepsy. KEPPRA is indicated as adjunctive therapy in the treatment of partial onset seizures in adults and children 1 month of age and older with epilepsy. KEPPRA is indicated as adjunctive therapy in the treatment of myoclonic seizures in adults and adolescents 12 years of age and older with juvenile myoclonic epilepsy. KEPPRA is indicated as adjunctive therapy in the treatment of primary generalized tonic-clonic seizures in adults and children 6 years of age and older with idiopathic generalized epilepsy. In a study published in the journal Epilepsia, nine out of fifteen patients (60%) showed a good response to levetiracetam with reduction in seizure frequency. Four out of fifteen patients (26.7%) became seizure free and 5/15 (33.3%) patients had fewer seizure frequency with therapy. Five out of fifteen patients (33.3%) had no response to therapy. Of those with generalized tonic clonic epilepsy, 4/11 (36.4%) had become seizure free, 3/11 (27.2%) had a decrease in seizure frequency, and 3/11 (27.2%) showed no response. Adverse effects occurred in four children (26.6%) including behavioral change (1/15), tiredness (1/15), decrease appetite (1/15), and reason unknown (1/15) (child was seen as an in-patient and followed at an outlying hospital). Treatment with levetiracetam resulted in reduction of seizure frequency in 60% of the patients included in this study. Levetiracetam was efficacious and safe for use in children with idiopathic generalized epilepsy with few, mild side effects reported. |  |
| Analyzer | Dimension RxL MAX, Siemens Healthcare Diagnostics Inc. Backup method is Medtox. |  |
| Sunquest Test Codes | LEVET: Levetiracetam (Keppra) |  |
| Specimen | Serum (no gel) (60 μL). Refer to specimen collection procedures for collection of diagnostic blood specimens. Plasma (lithium heparin) is an acceptable specimen type, however, using the same specimen matrix for individual patients is preferred, therefore serum is the preferred specimen type. Patient Preparation: A steady state, trough (pre-dose) sample is generally accepted as most consistent for therapeutic drug monitoring of Levetiracetam. Time of blood draw since last dose should be noted. Minimum volume: 0.2 mL Stability: 2-8 C / 7 days, -10 °C / 1 month. Rejection criteria: Unlabelled specimens, other than serum or heparinized plasma. Specimens collected using gel separators. Preparation: Process the blood as soon as possible after collection, since hydrolysis of levetiracetam may occur in the prolonged presence of whole blood. Complete clot formation should take place before centrifugation. Serum or plasma should be physically separated from cells as soon as possible with a maximum limit of two hours from the time of collection. Specimens should be free of particulate matter. Whole blood specimens should be centrifuged according to Specimen Processing procedures prior to analysis. See Processing Procedure Manual. Transfer serum or plasma to a properly labeled RXL sample cup or tube. Minimum labeling includes sample accession ID, and/ or patient name, medical record number, collection date and time. Do not use gel separators. Do not induce foaming and avoid repeated freezing and thawing to preserve the integrity of the specimen from the time it is collected until the time it is assayed. |  |
| Materials | Supplies; Reagent; Preparation; Stability. - Empty Flex™ Reagent Cartridge available from Siemens PN DF99 - Automatic Pipettes and tips - Sample Cups - Sample Segments - 6 mL syringe - 2 mL syringe - 21 ga needles; ARK Levetiracetam Assay – 5024-0001-00 Reagent – Antibody/Substrate 1 X 28 mL Reagent – Enzyme 1 X 14 mL; ARK Levetiracetam Assay reagents are provided liquid, ready to use and may be used directly from the refrigerator Do not freeze. Do not expose to temperatures above 32°C.; 2–8°C/ expiration date printed on the label when stored upright and tightly capped.. ARK Levetiracetam Calibrator – 5024-0002-00; Calibrators are ready to use. Mix each level by gentle inversion before dispensing. Squeeze 2 - 4 drops into individual sample cups for each level.; 2–8°C/ expiration date printed on the label when stored tightly capped. - Improper storage of reagents can affect assay performance - Do not mix reagents for ARK Levetiracetam between different kits. Use all reagents within the same kit lot together, or discard. |
| Risk and Safety | For In Vitro Diagnostic Use. For prescription use only. Reagents and are provided as a matched set and should not be interchanged with reagents from different lot numbers. Handle all patient specimens as if they were potentially infectious. |  |
| Preparing the Reagent Flex | STEP | ACTION |
|  |  | Label the outside of the ARK Diagnostics Levetiracetam kit with the following information - Expiration date of the kit - Flex Lot number (see step # 2 for explanation) - Sequence numbers of the Flexes 1, 2, 3, and 4 as they are prepared (see step # 3 for explanation) |
|  |  | Lot number is determined by the expiration date and month of the unopened kit using the Julian Calendar. There is a copy of a Julian calendar hanging from the left side of the MAX. - For example, the ARK kit for Levetiracetam expires April 30, 2013. The lot number would be LE3120 - LE is the mnemonic for Levetiracetam - First number, 3 is the expiration year (2013) - Last three numbers, 120 is the Julian calendar date (April 30th) 31days in Jan, 28 days in Feb 2013, 31 days in March, and 30 days in April |
|  |  | The number of Flexes per kit determines sequence number and the expiration date of each newly prepared Flex. There are 4 Flexes per kit (108 tests). The Flex is stable for 7 days. - For example, the first Flex was prepared 01/13/2012. The sequence number will be 10120 - The first number (1) is the first Flex of the kit. Second Flex will start with 2, etc. - The last 4 numbers are the month and day the prepared Flex expires |
|  |  | Obtain an empty Flex reagent cartridge (see attachment A) |
|  |  | Label the empty Flex package with the Lot number and Sequence number |
|  |  | Fill a 6 mL syringe with 6 mLs of R1 (component A), and use the tip of the needle to pierce the corner of the cellophane in well 1. |
|  |  | Add 2 mLs of R1 in wells 1, 2, and 3 of the empty compartment Flex |
|  |  | Fill a 3 mL syringe with 3 mLs of R2 (component B), and use the tip of the needle to pierce the corner of the cellophane in each well. |
|  |  | Add 1 mL of R2 (component B), in wells 4, 5, and 6 of the empty compartment Flex |
|  |  | Cover the vent and fill holes and the prepared flex may be stored for up to 24 hours at 2 – 8 C with the Levetiracetam reagents. Note on the flex the QC needs. Do NOT load flex on instrument if any tests remain in the previous flex. This is to assure quality control is run on the newly prepared flex before use with patient samples. |
|  |  | Load Flex (see procedure below) |
|  |  | Attachment A: Empty Flex Reagent Cartridge Well # 1 2 3 4 5 6 (2mL) (2mL) (2mL) (1mL) (1mL) (1mL) B A A A B B Narrow end of flex EMPTY GB2209057220 |
| Calibration | Assay Range:; 2.0 to 100.0 μg/mL. Reference Material:; ARK Levetiracetam Calibrators B, C, D, E, and F. Suggested Calibration Levels:; B 5.0 μg/mL 1 X 2 mL C 12.5 μg/mL 1 X 2 mL D 25.0 μg/mL 1 X 2 mL E 50.0 μg/mL 1 X 2 mL F 100.0 μg/mL 1 X 2 mL. Calibration Scheme:; Five levels in duplicate. Calibration Frequency:; - Whenever a new lot number of reagents is used - Whenever indicated by quality control results - Whenever required by standard laboratory protocols - Once every 40 days. Assigned Coefficients:; C0 0.000 C1 1.000. Analytical Measuring Range; 2.0 to 100.0 μg/mL The AMR is verified with each calibration using 5 levels of calibrator that span the full reportable range. Further studies are not necessary.. STEP; ACTION. Prepare Calibrators. See Materials section of this procedure.. Program the calibration from the Operating Menu, press F5: Process Control. Select F1: Calibration and press the “enter” key for the password. Press F2: Setup and Run. Select XLEV by pressing the Shift and AST keys Press F1: Other lot if necessary to select the appropriate lot number to be calibrated. Complete the information on the screen: - Operator ID: enter your tech ID - Calibrator Name/Lot #: enter the Standard Lot #. Do not change Name. - Segment Position: enter the letter of the segment and the start position for the Calibration (Example: A1) - Calibrator Bottle Values: enter 5, 12.5, 25.0, 50.0, and 100.0. Press F8: QC yes/no to add QC. Press F4: Assign Cups, then F7: Load/Run. Verify that there is adequate volume, load segment, and press F4: Run. To review data after calibration is completed: - From the Operating Menu, press: F5: Process Control - Press F1: Calibration, then hit the “enter” key for the password - Press F3: Review Data Select XLEV. |
| Calibration (cont) | Scan for outliers within the replicates for each level of calibrator NOTE: “Assay range” codes on the print out does not mean that the calibration needs to be rejected. IF; THEN. There is an outlier within the replicates; Press F8: Reject Data to reject the calibration Repeat the calibration. The precision is good with no outliers; Press F6: See QC to review the QC, Slope (m), Intercept (b), and Correlation Coefficient (r) with the calculated calibration. IF; THEN. Slope (m): is <0.97 or >1.03 Correlation coefficient is <0.990 or >1.000; Press F8: Reject Data to reject the calibration Verify that the programmed calibrator lot number and the bottle values are correct Repeat the calibration. Slope (m): is 0.97-1.003 QC is within the established range Intercept (b): close to zero or clinically insignificant Correlation coefficient(r): 0.990-1.000; Press F2: Accept Data to accept the calibration. |
| Quality Control | ARK Levetiracetam Control Dropper vials - LOW (7.5 μg/mL) 1 X 4 mL - HIGH (75.0 μg/mL) 1 X 4 mL Use each lot as a set Frequency: - Two levels of controls must be run every 24 hours - After loading a new Flex™ reagent cartridge - After calibration - After any major maintenance/ repairs have been performed on the analyzer - When indicated by QC results Storage and Stability: Unopened: 2°- 8°C. Open: until expiration date on label when stored tightly capped at 2°- 8°C Procedure Controls are ready to use. Mix each level by gentle inversion before dispensing. Waste 1 drop of control before dispensing. Squeeze sufficient volume (~40μL/drop) into individual sample cups for each level. Sunquest Control names: Level LOW = C-LEV1, Level HIGH = C-LEV2 Acceptable Ranges: Ranges are current in Sunquest and the instrument. Refer to the Quality Control Procedure for QC exception codes. |  |

| Loading of Flex™ Reagent Cartridge | Step | Action |
| --- | --- | --- |
|  | 1. | Press ALT I (Inventory Screen) |
|  | 2. | Press F4: Add Reagent |
|  | 3. | Place the prepared Levetiracetam flex in the automatic loader NOTE: Do not use the RMS autoloader. |
|  | 4. | When prompted, use the keyboard to enter the Flex cartridge information. |
|  | 5. | At METHOD: Press the “right delete” (\) to remove the default method. Enter the XLEV method by pressing Shift and AST (X indicates it is an open channel method) then press Enter. |
|  | 6. | At LOT NUMBER: Press “right delete” to remove the default lot number. Enter the lot number as specified in Procedure B above, then press Enter. |
|  | 7. | At SEQUENCE NUMBER: Press “right delete” to remove the default sequence number. Enter the sequence number as specified in Procedure B above then press Enter. Ig nore the last three numbers that follow the sequence number. |
|  | 8. | Press F1: ACCEPT |
|  | 9. | Allow the reagent to equilibrate for 15 – 30 minutes on the analyzer before use. |
| Analysis Procedure | Step; Action. 1; Check analyzer daily maintenance. IF; THEN. Routine maintenance has been performed; Proceed to step # 2. Routine maintenance has not been performed; Perform maintenance as described in the Siemens Dimension® Operation Guide.. 2; Check reagent inventory. IF; THEN. Reagent inventory sufficient; Proceed to step # 3. Reagents are insufficient; Prepare a new Flex™ reagent cartridge NOTE: Calibration is required when opening a new reagent kit.. 3; Check Quality Control status. IF; THEN. Quality Control has been performed within the last 24 hours; Proceed to step # 5. Quality Control has not been performed within the last 24 hours; Proceed to step # 4.. 4; Run Quality Control. IF; THEN. Quality Control within range; Proceed to step # 5. Quality Control out of acceptable range; Follow Troubleshooting controls.. 5; Load sample rack. 6; Press RUN on the keyboard. 7; Determine if the result is within the Analytical Measurement Range (AMR) and the Clinically Reportable Range (CRR). IF; THEN. Result is < AMR; Report as < AMR. Result is > AMR; Analyzer automatically dilutes the sample and calculates the results, or Make appropriate dilution when result is above the auto dilution range to a maximum dilution of 3... 8; Result is > CRR; Result as > CRR. |
| Interferences | Interference studies were conducted using CLSI/NCCLS Protocol EP7-A2 as a guideline. Drug Interference Levetiracetam-selective antibody did not cross-react with other anti-epileptic or co-administered drugs tested. A high concentration of each compound was spiked into normal human serum with known levels of Levetiracetam. Measurement of levetiracetam resulted in ≤10% error in the presence of drug compounds at the levels tested. Specificity Metabolites Levetiracetam is hydrolyzed to its major metabolite 2-pyrrolidone-N-butyric acid (ucb L057) and two minor metabolites. The metabolite ucb L057 was tested for cross-reactivity and demonstrated <10% interference and cross reactivity. See the product insert for a list of substances for which no significant interference was noted. |  |
| Reference Range | 5 – 45 mcg/mL Levetiracetam drug concentrations should be used in conjunction with information available from clinical evaluations. Circulating levels of levetiracetam may be affected by compliance, renal function, pregnancy, drug-drug interactions and timing of the sample draw. |  |
| Critical Values | None defined. |  |
| Limitations | Linear range of detection: 2.0 to 100.0 mcg/mL The instrument reporting system contains flags and comments to provide the user with information regarding instrument processing errors, instrument status information and potential errors in open channel method results. Refer to your Dimension Operator’s Guide for the meaning of report flags and comments. Any report containing flags and/or comments must be resolved prior to reporting. |  |
| Dilutions | - Above 100.0 mcg/mL: - Dilute results with “assay range” appended. - Make appropriate dilution with the zero calibrator (CAL A), to obtain results within the assay range. Determine optimum dilution by dividing RxL result obtained by 100 (the measuring range) and rounding up to the next whole number. The maximum dilution is 1:3. - Label diluted sample with “label foot” or Accession number, and dilution factor. - Program dilution factor in RxL. Reassay. Resulting readout is corrected for dilution. - Document dilutions and calculations, and have results checked prior to reporting. |  |
| Result Reporting | - Results between 2.0 to 100.0 mcg/mL without error messages are released - Results below 2.0 mcg/mL: report as < 2.0 mcg/mL instead of the numerical value. - Results >100.0 mcg/mL are reported as the numerical result following a maximum dilution of 1:3 - Results that exceed the assay range following the maximum dilution are reported as >300.0 mcg/mL. - To convert results from mcg/mL levetiracetam to μmol/L levetiracetam, multiply mcg/mL by 5.88 |  |
| Specimen Storage | Promptly stopper tested specimen and store upright in specimen rack. Every 8 hours remove specimens to refrigerated/freezer storage. Samples are retained 7 days in specimen storage freezer. |  |
| References | - ARK Diagnostics, Inc. package insert for the ARK Levetiracetam Assay, ARK Diagnostics, Inc., Sunnyvale, CA 94089 USA, 1600-0169-00 Rev 01, Printed in USA, Revised November 2009 - ARK Diagnostics, Inc. package insert for the ARK Levetiracetam Calibrator, ARK Diagnostics, Inc., Sunnyvale, CA 94089 USA, 1600-0170-00 Rev 01, Printed in USA, Revised November 2009 - ARK Diagnostics, Inc. package insert for the ARK Levetiracetam Control, ARK Diagnostics, Inc., Sunnyvale, CA 94089 USA, 1600-0171-00 Rev 01, Printed in USA, Revised November 2009 - ARK  Antiepileptic therapeutic Drug Monitoring Assays for Siemens Dimension RXL System, UDP 11-018 Rev 02 Siemens RXL, ARK Diagnostics, Inc. February 2012 - Epilepsia, Sunday, December 3, 2006 Poster Session II 7:30 a.m.–4:30 p.m. Article first published online: 30 OCT 2006 - Medtox Laboratories Test Catalogue, MEDTOX Scientific, Inc. 402 West County Road D, St. Paul, MN 55112, Phone: (800) 832-3244, (651) 636-7466 |  |
| Appendices | User Defined Method Specifications The following specifications are programmed into the Siemens Dimension RxL in the specified field under Diagnostics (F7) and Open Channels (F8). Refer to Dimension Operator’s Guide for more help. Channel # 3 Name: XLEV Measurement Mode: Absorbance Standard Curve: Logit Delivery; Time; Component 1; Component 2; Component 3; Chase; Mix. 1st R; -57.6; (A) 200 μL; 0; 0; 0 μL; NONE. 2nd R; 119.2; (B) 100 μL; 0; 0; 25 μL; MODERATE. 3rd R; ***; ( ) 0 μL; 0; 0; 0 μL; NONE. S1; 0.0; 3 μL; 0; 0; 25 μL; GENTLE. Photometry; Time; Cartridge Configuration; Well 1; Well 2; Well 3; Well 4; Well 5; Well 6. P1:; 220.0; Component:; (A); (A); (A); (B); (B); (B). P2:; 320.0; Aliquots:; 9; 9; 9; 9; 9; 9. P3:; ***; #Well Life [hours]:; 168; 168; 168; 168; 168; 168. P4:; ***; #On Board Life: 168 hrs (7 days); #Calibration: 960 hrs (40 days). Sample Volume:; 3 μL. Method Calculation The following specifications are programmed into the Dade Dimension in the specified field under Diagnostics (F7), Open Channels (F8), and Calculations (F7). Type all characters. { C = RATE(P1,P2, 340NM, 452NM); RETURN C; } |

| Historical Record | Version | Written/Revised by: | Effective Date: | Summary of Revisions |
| --- | --- | --- | --- | --- |
|  |  | Linda Lichty | 7/16/2012 | New test |
|  |  | Linda Lichty | 10/8/2012 | Set reference range, defined QC |
|  |  | Linda Lichty | 6/28/13 | Add policy, do not mix kit components |
|  |  | Erin Bartos | 12/16/2016 | Added #10 to preparing the flex |