| Topiramate |  |  |
| --- | --- | --- |
| Purpose | This procedure provides instructions for performing the Topiramate procedure on serum in Children’s Laboratory. The ARK Topiramate Assay is a homogeneous enzyme immunoassay intended for the quantitative determination of topiramate in human serum or plasma on automated clinical chemistry analyzers. The results obtained are used in the diagnosis and treatment of topiramate overdose and in monitoring levels of topiramate to help ensure appropriate therapy. |  |
| Policy Statements | This procedure applies to all personnel who run the Dimension RxL MAX |  |
| Principle | ARK Topiramate Assay is a homogeneous enzyme immunoassay based on competition between drug in the specimen and topiramate epitope labeled with the enzyme glucose-6-phosphate dehydrogenase (G6PDH) for binding to the antibody reagent. As the latter binds antibody, enzyme activity decreases. In the presence of drug from the specimen, enzyme activity increases and is directly proportional to the drug concentration. Active enzyme converts the coenzyme nicotinamide adenine dinucleotide (NAD) to NADH that is measured spectrophotometrically as a rate of change in absorbance. Endogenous serum G6PDH does not interfere with the results because the coenyzme NAD functions only with the bacterial enzyme used in the assay. |  |
| Clinical Significance | Topiramate (2,3:4,5-Di-O-isopropylidene--D-fructopyranose sulfamate) is an anti-convulsant drug approved for use as adjunctive therapy in the treatment of epilepsy and is often prescribed as mono-therapy or as one component of a multiple anti-epileptic drug therapy. Topiramate is used alone or with other medications to treat certain types of seizures in people who have epilepsy. Topiramate is also used with other medications to control seizures in people who have Lennox-Gastaut syndrome (a disorder that causes seizures and developmental delays). Topiramate is used to treat patients who continue to have seizures even when they take other antiseizure medications. Topiramate is also used to prevent migraine headaches but not to relieve the pain of migraine headaches when they occur. Topiramate is in a class of medications called anticonvulsants. It works by decreasing abnormal excitement in the brain. Hyperchloremic, non-anion gap, metabolic acidosis (i.e., decreased serum bicarbonate below the normal reference range in the absence of chronic respiratory alkalosis) is associated with Topiramate treatment. This metabolic acidosis is caused by renal bicarbonate loss due to the inhibitory effect of Topiramate on carbonic anhydrase. Changes in several clinical laboratory values (increased creatinine, BUN, alkaline phosphatase, total protein, total eosinophil count and decreased potassium) have been observed in a clinical investigational program in very young (<2 years) pediatric patients who were treated with adjunctive Topiramate for partial onset seizures. Topiramate treatment produced a dose-related increased shift in serum creatinine from normal at baseline to an increased value at the end of 4 months treatment in adolescent patients (ages 12-16 years) who were treated for another indication in a double-blind, placebo-controlled study. Topiramate treatment with or without concomitant valproic acid (VPA) can cause hyperammonemia with or without encephalopathy. |  |
| Analyzer | Dimension RxL MAX, St Paul, MN, Siemens Healthcare Diagnostics Inc. |  |
| Sunquest Test Codes | TOPX Topiramate (Topamax, Topiragen) |  |
| Sample | Plasma (lithium heparin) NO GEL Refer to specimen collection procedures for collection of diagnostic blood specimens. Serum NO GEL is an acceptable specimen type, however, using the same specimen matrix for individual patients is preferred. Patient Preparation: A steady state, trough (pre-dose) sample is generally accepted as most consistent for therapeutic drug monitoring of Topiramate. Time of blood draw since last dose should be noted. Minimum volume: 0.2 mL Actual Test volume: 4 uL Stability: 2-8 C / 7 days, -10 °C / 1 month. Rejection criteria: Unlabelled specimens, other than serum or heparinized plasma. Specimens collected using gel separators. Preparation: Serum or plasma should be physically separated from cells as soon as possible with a maximum limit of two hours from the time of collection. Specimens should be free of particulate matter. Whole blood specimens should be centrifuged according to Specimen Processing procedures prior to analysis. See Processing Procedure Manual. Transfer serum or plasma to a properly labeled RXL sample cup or tube. Minimum labeling includes sample accession ID, and/ or patient name, medical record number, collection date and time. - Do not use gel separators. - Do not induce foaming and avoid repeated freezing and thawing - Fibrin, red blood cells, and other particulate matter may cause an erroneous result. Ensure adequate centrifugation |  |
| Materials | Supplies; Reagent; Preparation; Stability. - Empty Flex™ Reagent Cartridge available from Siemens PN DF99 - Automatic Pipettes and tips - Sample Cups - Sample Segments - 6 mL syringe - 2 mL syringe - 21 ga needles; ARK Topiramate Assay – 5015-0001-00 Reagent – Antibody/Substrate 1 X 28 mL Reagent – Enzyme 1 X 14 mL; - ARK Topiramate Assay reagents are provided liquid, ready to use and may be used directly from the refrigerator - Reagents R1 and R2 will be transferred to analyzer-specific reagent containers prior to use. - Avoid cross-contamination of R1 and R2 - Do not freeze. - Do not expose to temperatures above 32°C.; 2–8°C/ expiration date printed on the label when stored upright and tightly capped. Improper storage of reagents can affect assay performance. ARK Topiramate Calibrator – 5015-0002-00; - Calibrators are ready to use. - Mix each level by gentle inversion before dispensing. - Squeeze 2 - 4 drops of each level into individual sample cups.; 2–8°C/ expiration date printed on the label when stored tightly capped. |
| Risk and Safety | For In Vitro Diagnostic Use. For prescription use only. Reagents R1 and R2 are provided as a matched set and should not be interchanged with reagents from different lot numbers. Handle all patient specimens as if they were potentially infectious. |  |
| Preparing the Reagent Flex | STEP | ACTION |
|  |  | Label the outside of the ARK Diagnostics Topiramate kit with the following information - Expiration date of the kit - Flex Lot number (see step # 2 for explanation) - Sequence numbers of the Flexes 1, 2, 3, and 4 (see step # 3 for explanation) |
|  |  | Lot number is determined by the expiration date and month of the unopened kit using the Julian Calendar. There is a copy of a Julian calendar hanging from the left side of the MAX. - For example, the ARK kit for Topiramate expires April 30, 2014. The lot number would be TO4120 - TO is the mnemonic for Topiramate - First number, 4 is the expiration year (2014) - Last three numbers, 120 is the Julian calendar date (April 30th) 31days in Jan, 28 days in Feb 2014, 31 days in March, and 30 days in April |
|  |  | The number of Flexes per kit determines sequence number and the expiration date of each newly prepared Flex. There are 4 Flexes per kit (108 tests). The Flex is stable for 7 days. - For example, the first Flex was prepared 01/13/2013. The sequence number will be 10120 - The first number (1) is the first Flex of the kit. Second Flex will start with 2, etc. - The last 4 numbers are the month and day the prepared Flex expires |
|  |  | Obtain an empty Flex reagent cartridge (see attachment A) |
|  |  | Label the empty Flex package with the Lot number and Sequence number |
|  |  | Fill a 6 mL syringe with 6 mLs of R1 (component A), and use the tip of the needle to pierce the corner of the cellophane in well 1. |
|  |  | Add 2 mLs of R1 in wells 1, 2, and 3 of the empty compartment Flex |
|  |  | Fill a 3 mL syringe with 3 mLs of R2 (component B), and use the tip of the needle to pierce the corner of the cellophane in each well. |
|  |  | Add 1 mL of R2 (component B), in wells 4, 5, and 6 of the empty compartment Flex |
|  |  | Cover the vent and fill holes and the prepared flex may be stored for up to 24 hours at 2 – 8 C with the Topiramate reagents. Note on the flex the QC needs. Do NOT load flex on instrument if any tests remain in the previous flex. This is to assure quality control is run on the newly prepared flex before use with patient samples. |
|  |  | Load Flex (see procedure below) |
|  |  | Attachment A: Empty Flex Reagent Cartridge Well # 1 2 3 4 5 6 (2mL) (2mL) (2mL) (1mL) (1mL) (1mL) B A A A B B Narrow end of flex EMPTY GB2209057220 |
| Calibration | Assay Range:; 1.5 to 54.0 mcg/mL. Reference Material:; ARK Topiramate Calibrators B, C, D, E, and F. Suggested Calibration Levels:; B 2.0 mcg/mL C 4.0 mcg/mL D 8.0 mcg/mL E 24.0 mcg/mL F 60.0 mcg/mL. Calibration Scheme:; Five levels in duplicate. Calibration Frequency:; - Whenever a new lot number of reagents is used - Whenever indicated by quality control results - Whenever required by standard laboratory protocols - Once every 40 days. Assigned Coefficients:; C0 0.000 C1 1.000. Analytical Measuring Range; 1.5 – 54.0 mcg/mL The AMR is verified with each calibration using 5 levels of calibrator that span the full reportable range. Further studies are not necessary.. STEP; ACTION. Prepare Calibrators. See Materials section of this procedure.. Program the calibration from the Operating Menu, press F5: Process Control. Select F1: Calibration and press the “enter” key for the password. Press F2: Setup and Run. Select XTOP by pressing the Shift and CA keys Press F1: Other lot if necessary to select the appropriate lot number to be calibrated. Complete the information on the screen: - Operator ID: enter your tech ID - Calibrator Name/Lot #: enter the Standard Lot #. Do not change Name. - Segment Position: enter the letter of the segment and the start position for the Calibration (Example: A1) - Calibrator Bottle Values: enter 2, 4, 8, 24, and 60.. Press F8: QC yes/no to add QC. Press F4: Assign Cups, then F7: Load/Run. Verify that there is adequate volume, load segment, and press F4: Run. To review data after calibration is completed: - From the Operating Menu, press: F5: Process Control - Press F1: Calibration, then hit the “enter” key for the password - Press F3: Review Data Select XTOP. |
| Calibration (cont) | Scan for outliers within the replicates for each level of calibrator NOTE: “Assay range” codes on the print out does not mean that the calibration needs to be rejected. IF; THEN. There is an outlier within the replicates; Press F8: Reject Data to reject the calibration Repeat the calibration. The precision is good with no outliers; Press F6: See QC to review the QC, Slope (m), Intercept (b), and Correlation Coefficient (r) with the calculated calibration. IF; THEN. Slope (m): is <0.97 or >1.03 Correlation coefficient is <0.990 or >1.000; Press F8: Reject Data to reject the calibration Verify that the programmed calibrator lot number and the bottle values are correct Repeat the calibration. Slope (m): is 0.97-1.003 QC is within the established range Intercept (b): close to zero or clinically insignificant Correlation coefficient(r): 0.990-1.000; Press F2: Accept Data to accept the calibration. |
| Quality Control | ARK Topiramate Control Dropper vials - LOW (2.5 mcg/mL) 1 X 4 mL - MID (10.0 mcg/mL) 1 X 4 mL - Use each lot as a set Frequency: - Two levels of controls must be run every 24 hours - After loading a new Flex™ reagent cartridge - After calibration - After any major maintenance/ repairs have been performed on the analyzer - When indicated by QC results Storage and Stability: Unopened: 2°- 8°C. Use prior to expiration date on container Open: until expiration date on label when stored tightly capped at 2°- 8°C Procedure Controls are ready to use. Mix each level by gentle inversion before dispensing. Waste 1 drop and then squeeze sufficient volume (~40μL/drop) into individual sample cups. Sunquest Control names: Level LOW = C-TOP1, Level MID = C-TOP2 Acceptable Ranges: Ranges are current in Sunquest and the instrument. Refer to the Quality Control Procedure for QC exception codes. |  |

| Loading of Flex™ Reagent Cartridge | Step | Action |
| --- | --- | --- |
|  | 1. | Press ALT I (Inventory Screen) |
|  | 2. | Press F4: Add Reagent |
|  | 3. | Place the prepared Topiramate flex in the automatic loader NOTE: Do not use the RMS autoloader. |
|  | 4. | When prompted, use the keyboard to enter the Flex cartridge information. |
|  | 5. | At METHOD: Press the “right delete” (\) to remove the default method. Enter the XTOP method by pressing Shift and CA (X indicates it is an open channel method) then press Enter. |
|  | 6. | At LOT NUMBER: Press “right delete” to remove the default lot number. Enter the lot number as specified in Procedure B above, then press Enter. |
|  | 7. | At SEQUENCE NUMBER: Press “right delete” to remove the default sequence number. Enter the sequence number as specified in Procedure B above then press Enter. Ignore the last three numbers that follow the sequence number. |
|  | 8. | Press F1: ACCEPT |
|  | 9. | Allow the reagent to equilibrate for 15 – 30 minutes on the analyzer before use. |
| Analysis Procedure | Step; Action. 1; Check analyzer daily maintenance. IF; THEN. Routine maintenance has been performed; Proceed to step # 2. Routine maintenance has not been performed; Perform maintenance as described in the Siemens Dimension® Operation Guide.. 2; Check reagent inventory. IF; THEN. Reagent inventory sufficient; Proceed to step # 3. Reagents are insufficient; Prepare a new Flex™ reagent cartridge NOTE: Calibration is required when opening a new reagent kit.. 3; Check Quality Control status. IF; THEN. Quality Control has been performed within the last 24 hours; Proceed to step # 5. Quality Control has not been performed within the last 24 hours; Proceed to step # 4.. 4; Run Quality Control. IF; THEN. Quality Control within range; Proceed to step # 5. Quality Control out of acceptable range; Follow Troubleshooting controls.. 5; Load sample rack. 6; Press RUN on the keyboard. 7; Determine if the result is within the Analytical Measurement Range (AMR) and the Clinically Reportable Range (CRR). IF; THEN. Result is < AMR; Report as < AMR. Result is > AMR; Analyzer automatically dilutes the sample and calculates the results, or Make appropriate dilution when result is above the auto dilution range to a maximum dilution of 2... 8; Result is > CRR; Result as > CRR. |
| Interferences | Interference studies were conducted by ARK, Inc. using CLSI/NCCLS Protocol EP7-A2 as a guideline. Drug Interference Topiramate-selective antibody did not cross-react with other anti-epileptic or co-administered drugs tested. A high concentration of each compound was spiked into normal human serum with known levels of topiramate (approximately 5 and 20 μg/mL) and assayed along with a serum control of topiramate. Measurement of topiramate resulted in ≤10% error in the presence of drug compounds at the levels tested. Clinically high concentrations of potentially interfering substances in serum with known levels of topiramate (approximately 5 and 20 mcg/mL) were evaluated. Each sample was assayed using the ARK Topiramate Assay, along with a serum control of topiramate. Measurement of topiramate resulted in ≤10% error in the presence of interfering substances at the levels tested. Refer to the product insert for additional information. Specificity Metabolites Metabolites of topiramate are found primarily in urine of patients being administered topiramate therapy. ARK Topiramate Assay serum and plasma results are unlikely to be affected by metabolism of topiramate drug, since plasma levels of metabolites are usually not clinically significant. See the product insert for a list of substances for which no significant interference was noted. |  |
| Reference Range | 2 to 25 mcg/mL A therapeutic range for topiramate has not been well established. The therapeutic range on a trough sample for seizure control is 2 to 25 mcg/mL. An inconsistent correlation exists between levels of circulating topiramate to toxicity, adverse affect or clinical efficacy. Topiramate drug concentrations should be used in conjunction with information available from clinical evaluations and other diagnostic procedures. Clinicians should carefully monitor patients during therapy and dosage adjustments. Pharmacokinetics may vary widely, particularly with co-medications, age, and/or compromised renal function. Multiple samples over time may be needed to determine steady-state concentrations for individual patients. |  |
| Critical Values | None defined. |  |
| Limitations | Linear range of detection: 1.5 to 54.0 mcg/mL The instrument reporting system contains flags and comments to provide the user with information regarding instrument processing errors, instrument status information and potential errors in open channel method results. Refer to your Dimension Operator’s Guide for the meaning of report flags and comments. Any report containing flags and/or comments must be resolved prior to reporting. |  |
| Dilutions | - Above 54.0 mcg/mL: - Dilute results with “assay range” appended. - Prepare a 1:2 maximum dilution with the zero calibrator (CAL A), to obtain results within the assay range. - Label diluted sample with “label foot” or Accession number, and dilution factor. - Program dilution factor in RxL. Reassay. Resulting readout is corrected for dilution. - Document dilutions and calculations, and have results checked prior to reporting. |  |
| Result Reporting | - Results between 1.5 to 54.0 mcg/mL without error messages are released - Results below 1.5 mcg/mL: report as < 1.5 mcg/mL instead of the numerical value. - Results >54.0 mcg/mL are reported as the numerical result following a maximum dilution of 1:2 - Results that exceed the assay range following the maximum dilution are reported as >108.0 mcg/mL. - To convert results from mcg/mL topiramate to μmol/L topiramate, multiply mcg/mL by 2.95. |  |
| Specimen Storage | Promptly stopper tested specimen and store upright in specimen rack. Every 8 hours remove specimens to refrigerator/freezer storage. Samples are retained 7 days in specimen storage freezer. |  |
| Alternate Method | Refer samples to Medtox Laboratory: - When the Children’s Hospital method fails quality parameters, and cannot be used - Twice yearly for Alternate Proficiency assessment, after testing has been completed in house. - Order test TOPI in Sunquest |  |
| References | - ARK Diagnostics, Inc. package insert for the ARK Topiramate Assay, ARK Diagnostics, Inc., 48089 Fremont Blvd, Fremont, CA 94538, USA, 1600-0105-00 Rev 04, Printed in USA, Revised August 2012 - ARK Diagnostics, Inc. package insert for the ARK Topiramate Calibrator, ARK Diagnostics, Inc., 48089 Fremont Blvd, Fremont, CA 94538, USA, 1600-0137-00 Rev 03, Printed in USA, Revised August 2012 - ARK Diagnostics, Inc. package insert for the ARK Topiramate Control, ARK Diagnostics, Inc., 48089 Fremont Blvd, Fremont, CA 94538, USA, 1600-0138-00 Rev 03, Printed in USA, Revised August 2012 - ARK  Antiepileptic therapeutic Drug Monitoring Assays for Siemens Dimension RXL System, UDP 11-018 Rev 02 Siemens RXL, ARK Diagnostics, Inc. February 2012 - Epilepsia, Sunday, December 3, 2006 Poster Session II 7:30 a.m.–4:30 p.m. Article first published online: 30 OCT 2006 - http://www.nlm.nih.gov/medlineplus/druginfo/meds/a697012.html - Torrent Pharma Inc., 5380 Holiday Terrace, Suite 40, Kalamazoo, Michigan 49009. 8029987 Revised August 2011 |  |
| Appendices | User Defined Method Specifications The following specifications are programmed into the Siemens Dimension RxL in the specified field under Diagnostics (F7) and Open Channels (F8). Refer to Dimension Operator’s Guide for more help. Channel # 4 Name: XTOP Measurement Mode: Absorbance Standard Curve: Logit Delivery; Time; Component 1; Component 2; Component 3; Chase; Mix. 1st R; -57.6; (A) 200 μL; 0; 0; 0 μL; NONE. 2nd R; 119.2; (B) 100 μL; 0; 0; 25 μL; MODERATE. 3rd R; ***; ( ) 0 μL; 0; 0; 0 μL; NONE. S1; 0.0; 4 μL; 0; 0; 25 μL; GENTLE. Photometry; Time; Cartridge Configuration; Well 1; Well 2; Well 3; Well 4; Well 5; Well 6. P1:; 220.0; Component:; (A); (A); (A); (B); (B); (B). P2:; 320.0; Aliquots:; 9; 9; 9; 9; 9; 9. P3:; ***; #Well Life [hours]:; 168; 168; 168; 168; 168; 168. P4:; ***; #On Board Life: 168 hrs (7 days); #Calibration: 960 hrs (40 days). Sample Volume:; 4 μL. Method Calculation The following specifications are programmed into the Siemens Dimension in the specified field under Diagnostics (F7), Open Channels (F8), and Calculations (F7). Type all characters. { C = RATE(P1,P2, 340NM, 452NM); RETURN C; } |

| Historical Record | Version | Written/Revised by: | Effective Date: | Summary of Revisions |
| --- | --- | --- | --- | --- |
|  | 1 | Linda Lichty | 10/29/2013 | New test |
|  | 2 | Erin Bartos | 12/16/2016 | Added #10 to preparing the flex |