| Performing BCS-XP Daily Maintenance |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Purpose | This procedure provides instructions for PERFORMING BCS-XP DAILY MAINTENANCE. |  |  |  |  |  |  |
| Policy Statements | - This procedure applies to all laboratory technologists who operate the BCS-XP, and to the section supervisor. |  |  |  |  |  |  |
| Materials | Supplies |  |  |  | Equipment |  |  |
|  | - Distilled water - Clean towel - Barbicide disinfectant solution. King Research chc# 31111. Prepare working solution by diluting one 125ml bottle of concentrate to 2.0 L with deionized water. Working Barbicide solution is stable for 8 weeks. Do not use this product for cleaning surfaces, lanes or racks on the analyzer. |  |  |  | - N/A |  |  |
| Procedure | Follow the activities in the table below for PERFORMING BCS-XP DAILY MAINTENANCE. |  |  |  |  |  |  |
|  | Step | Action |  |  |  |  | Related Document |
|  | 1 | Print Joblist icon: - Click JOBLIST. - Click PRINT. |  |  |  |  |  |
|  | 2 | Print Control Journal icon: - Click CONTROL JOURNAL icon. - Click PRINT. |  |  |  |  |  |
|  | 3 | Check temperature on the Analyzer screen, note failures/corrective action in the Action Log |  |  |  |  |  |
|  | 4 | Empty rotor waste (if the cardboard bin is removed, rotors must be dumped out). |  |  |  |  |  |
|  | 5 | Empty fluid waste container; refill DH2O container with distilled water. |  |  |  |  |  |
|  | 6 | Rinse tubing system: - Click System, then on Syringe. - Click Continue. - Click both Left and Right syringe boxes. - Select Dist. Water in the “Rinse with…” box. - Enter 1 in the No. of rinsing cycles field. - Click Rinse. - Click Close when finished. |  |  |  |  |  |
|  | 7 | Wipe pipettors with alcohol wipe: - Click System and then on Park. - Click Login. - Open plexiglass shield when pipetters have stopped moving. - Wipe both pipetters with downward motion. - Close the shield. - Click Done. - Click Close. |  |  |  |  |  |
|  | 8 | Check pipetters delivery stream: - Click System and select Probe. - Click Continue. - Both the Left and Right boxes should have an X in them. - Click the triangle in the Washing intensity box. - Select Check probe from the pop-up menu. - Click Start. - View the streams through the shield as DH2O is dispensed into the wash stations. - If both pipettors are dispensing without a spray or as a divided stream, click Close. - Stream divided or sprays? Click the triangle in the Washing intensity box, select Intensive, click Start. - Recheck delivery stream. - Click Close when finished |  |  |  |  |  |
|  | 9 | Perform Disinfect (shift change): - Click System. - Select Shift change. - Click Options: - Select Delete samples and sample measurements. - Click Close. - Click Delete. |  |  |  |  |  |
|  | 10 | Replace Washing Solution (every 48 hours): - Click System. - Select Washing solution. - Click Left bottle and Center bottle. - Click Change. - Open plexiglas shield. - Change bottles. - Close the shield. - Click Done. - Click Close. |  |  |  |  |  |
|  | 11 | Add fresh Saline to Saline bottle. |  |  |  |  |  |
| References | - BCS®XP System Instruction Manual, 1000 767.0506, Manual Version 1.0, Siemens Diagnostics Inc., Marburg Germany, copyright 2006. |  |  |  |  |  |  |
| Historical Record | Version |  | Written/Revised by: | Effective Date: |  | Summary of Revisions |  |
|  | 1 |  | Al Quigley | 06/01/11 |  | Updated, Reformatted |  |
|  | 2 |  | Al Quigley | 02/06/17 |  | Implemented the use of Barbicide as the onboard disinfectant |  |