| Performing BCS-XP Monthly Maintenance |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Purpose | This procedure provides instructions for PERFORMING BCS-XP MONTHLY MAINTENANCE. |  |  |  |  |  |  |
| Policy Statements | - This procedure applies to all laboratory technologists who operate the BCS-XP, and to the section supervisor. |  |  |  |  |  |  |
| Materials | Supplies |  |  |  | Equipment |  |  |
|  | - Deionized water - Clean towel - 70% Isopropyl Alcohol (2-Propanol) CHC# 15609, Ricca Chemical, Vendor Item # 4210-1, Cardinal Health. Gallon Jugs Available. Do not use this product for cleaning surfaces, lanes or racks on the analyzer. - Barbicide disinfectant solution. King Research chc# 31111. Prepare working solution by diluting one 125ml bottle of concentrate to 2.0 L with deionized water. Working Barbicide solution is stable for 8 weeks. Do not use this product for cleaning surfaces, lanes or racks on the analyzer. - 5% Bleach Solution - BCS-XP Validation Kit (part# OQSC 11) |  |  |  | - N/A |  |  |
| Procedure | Follow the activities in the table below for PERFORMING BCS-XP MONTHLY MAINTENANCE. |  |  |  |  |  |  |
|  | Step | Action |  |  |  |  | Related Document |
|  | 1 | Complete the daily maintenance. MAI 1.5 Performing BCS-XP Daily Maintenance |  |  |  |  |  |
|  | 2 | Complete the weekly maintenance MAI 1.7 Performing BCS-XP Weekly Maintenance Then add the following: - Clean the barcode reader with 5% Bleach. - Clean the keyboard with 5% Bleach. - Clean the mouse pad with 5% Bleach. - Oil syringe tips. Refer to Instruction Manual, Section 11.3.1, pp. 13-15 for syringe removal instructions, and to page 23 for oiling instructions. - Data Maintenance Backup: Refer to Instruction Manual, Section 10.16, pp. 118-121. |  |  |  |  |  |
|  | 3 | Turn on BCS-XP according to the weekly maintenance procedure. MAI 1.7 Performing BCS-XP Weekly Maintenance |  |  |  |  |  |
|  | 4 | Use the BCS-XP Validation Kit (part# OQSC 11) to verify the BCS-XP meets pipetting and optics precision specifications: - Open the Validation Kit. - Fill the tube labeled Empty Tube 0101 with deionized water. - Fill the bottle labeled Empty bottle with deionized water. - Ensure that the BCS-XP has a minimum of 5 rotors. - Remove the cap from the bottle labeled Test Solution R. - Place it and the deionized water bottle in the reagent rack, load in lane 4. - Remove the caps from sample tubes 105, 106, 107, and 108. Place the tubes in a sample rack. - Load the sample rack on the analyzer. |  |  |  |  |  |
|  | 5 | Perform the following tests separately: - Request the Precision Sample test for sample 105: DO NOT SELECT ANY OTHER TESTS UNTIL THE RESULT OF SAMPLE 105 IS DISPLAYED. The result must be < 5.0. This is the precision of the sample pipetter. If the result is > 5.0, perform an intensive wash, then repeat the precision sample test. When the result is correct, continue with the next sub-step. - Request the Precision Reagent test for sample 101: DO NOT SELECT ANY OTHER TESTS UNTIL THE RESULT OF SAMPLE 101 IS DISPLAYED. The result must be < 5.0. This is the precision of the sample pipetter. If the result is > 5.0, perform an intensive wash, then repeat the precision sample test. When the result is correct, continue with the next sub-step. - Request the Precision SampleRep test for samples 106, 107, and 108: Each sample must have a value between 0.95 and 1.05. If any sample is out of range, perform pipetter maintenance and repeat the test. |  |  |  |  |  |
|  | 6 | Log the results on the BCS-XP Validation Result Form. |  |  |  |  |  |
| References | - BCS®XP System Instruction Manual, 1000 767.0506, Manual Version 1.0, Siemens Diagnostics Inc., Marburg Germany, copyright 2006. |  |  |  |  |  |  |
| Historical Record | Version |  | Written/Revised by: | Effective Date: |  | Summary of Revisions |  |
|  | 1 |  | Al Quigley | 06/01/11 |  | Updated, Reformatted |  |
|  | 2 |  | Al Quigley | 02/06/17 |  | Implemented the use of Barbicide as the onboard disinfectant. |  |