| Performing BCS-XP Weekly Maintenance |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Purpose | This procedure provides instructions for PERFORMING BCS-XP WEEKLY MAINTENANCE. |  |  |  |  |  |  |
| Policy Statements | - This procedure applies to all laboratory technologists who operate the BCS-XP, and to the section supervisor. |  |  |  |  |  |  |
| Materials | Supplies |  |  |  | Equipment |  |  |
|  | - Deionized water - Clean towel - 70% Isopropyl Alcohol (2-Propanol) CHC# 15609, Ricca Chemical, Vendor Item # 4210-1, Cardinal Health. Gallon Jugs Available. Do not use this product for cleaning surfaces, lanes or racks on the analyzer. - Barbicide disinfectant solution. King Research chc# 31111. Prepare working solution by diluting one 125ml bottle of concentrate to 2.0 L with deionized water. Working Barbicide solution is stable for 8 weeks. Do not use this product for cleaning surfaces, lanes or racks on the analyzer. - 5% Bleach Solution |  |  |  | • N/A |  |  |
| Special Safety Precautions | Dispose of working Barbicide solution into a waste container located in the Flammable cabinet. When this container is ¾ full it should be taken to the hospital disposal site for hazardous waste and replaced with an empty container in the Flammable cabinet. Dispose of 70% Isopropyl alcohol into a waste container in the flammable cabinet. When this container is ¾ full it should be taken to the hospital disposal site for hazardous waste and replaced with an empty container in the Flammable cabinet. |  |  |  |  |  |  |
| Procedure | Follow the activities in the table below for PERFORMING BCS-XP WEEKLY MAINTENANCE. |  |  |  |  |  |  |
|  | Step | Action |  |  |  |  | Related Document |
|  | 1 | - Disinfect the distilled water container with 70% Isopropyl alcohol. - Allow 70% Isopropyl alcohol to sit for at least five (5) minutes. - Shake and swirl occasionally. - Empty the container. |  |  |  |  |  |
|  | 2 | Rinse the distilled water container with copious amounts of distilled water before refilling. |  |  |  |  |  |
|  | 3 | Disinfect the supply tube (canula) for the distilled water container: - Remove the canula from the disinfectant canister and place it on a clean towel. - Insert the distilled water canula into the disinfectant canister. - Click System. - Click Tube Disinfection. - Click Continue, the dialogue Disinfect tubes is displayed, click Continue. The system will now fill the tubing with Barbicide disinfectant. - Wait for the procedure to finish. When the filling of the tubes is finished and the incubation period of three minutes is over, a new message is displayed in the dialogue Disinfect tubes, asking you to remove the intake nozzle from the container with disinfectant and place it into an extra container with distilled water. DO NOT REMOVE THE NOZZLE, SET A TIMER AND ALLOW THE DISINFECTANTSOLUTION TO REMAIN IN THE TUBES FOR A MINIMUM OF SEVEN MORE MINUTES. - When completed, remove the distilled water canula from the container of Barbicide and place in an extra container of distilled water. The intake nozzle must be submerged deeper in the distilled water than it had previously been in the disinfectant solution. h. Click Continue, the system is now rinsed with water. Wait for the procedure to finish, remove the distilled water canula from the extra container and place it back in the water container. Click Quit. i. Put disinfectant canula back into the disinfectant. |  |  |  |  |  |
|  | 4 | - Click File, Quit, OK. - Remove the check from the Disinfect Instrument box. - Click Quit now to close the BCS-XP program. |  |  |  |  |  |
|  | 5 | Turn off the BCS-XP analyzer (the switch is located on the back left side of the instrument). |  |  |  |  |  |
|  | 6 | Clean and disinfect the surface of the BCS-XP with 5% Bleach solution. |  |  |  |  |  |
|  | 7 | Clean and disinfect the racks with a lint-free cloth soaked with 5% Bleach solution. |  |  |  |  |  |
|  | 8 | Clean and disinfect the lanes in the BCS-XP with a lint-free cloth soaked with 5%Bleach solution. |  |  |  |  |  |
|  | 9 | Power up the BCS-XP analyzer. |  |  |  |  |  |
|  | 10 | Click BCS-XP on the Windows Desktop. |  |  |  |  |  |
|  | 11 | Clean the rotor handler’s rubber grippers: - Open the waste lid and remove the waste container. - Reach upwards through the waste depository and clean the grippers with a Kim-wipe covered, cotton-tipped applicator stick soaked in water. - Replace the waste container and close the lid. |  |  |  |  |  |
| References | - BCS®XP System Instruction Manual, 1000 767.0506, Manual Version 1.0, Siemens Diagnostics Inc., Marburg Germany, copyright 2006. |  |  |  |  |  |  |
| Historical Record | Version |  | Written/Revised by: | Effective Date: |  | Summary of Revisions |  |
|  | 1 |  | Al Quigley | 06/01/11 |  | Updated, Reformatted |  |
|  | 2 |  | Al Quigley | 02/06/17 |  | Implemented the use of Barbicide as the onboard disinfectant. |  |