# Simplexa™ RSV & Influenza A/B Direct Procedure

PURPOSE

- This procedure provides instructions for preparing samples, setting up the PCR reaction and running the Simplexa™ RSV & Influenza A/B Direct for the detection of RSV & influenza A/B from nasal specimens

#### POLICY STATEMENT

- PCR testing is performed daily, 0700 –1530

ABBREVIATIONS

| - ABC : Analyzer Before Computer - BSC: BioSafety Cabinet - BSL: BioSafety level - CBA: Computer Before Analyzer - Ct : crossing threshold - DAD : Direct Amplification Disc - F/T : freeze/thaw - IC : internal control - LIS: laboratory information system - MM : master mix - NA : Nucleic Acid - NEGC : negative control | - NP: nasopharyngeal swab - NW: nasal wash specimen - PCR: polymerase chain reaction - POSC: positive control - PPE: personal protective equipment - RIP: Simplexa RSV & Influenza A/B PCR - UNAC: Specimen unacceptable, please recollect - UTM: universal viral transport media - Area/Room 1: Clean room - Area/Room 2: Processing room - Area/Room 3: Amplification room |
| --- | --- |

## DOCUMENTATION/RECORDS

- Simplexa Flu A/B & RSV Direct Segment Report - run-specific

- LIS Incomplete and worksheets

- Pending Log

- Daily Maintenance Log

## SAFETY CONSIDERATIONS

- Standard precautions for infectious agents. Refer to MB 2.02, Biohazard containment

- Use of engineering controls: Refer to MB 3.01 Engineering Controls to Prevent Nucleic Acid Contamination

- General Safety: MB 2.01 Safe Work Practices

- Caution: PPE including protective eyewear must be worn when working with concentrated Extran

#### MATERIALS REQUIRED

| Equipment | Reagents | Supplies |
| --- | --- | --- |
| Room 1: Clean room - -10 to -30° C freezer - Laminar flow Hood Room 2: Processing - Refrigerator 2 – 8° C - BSC BSL-2 - -70⁰ C freezer - 100 or 200 µl pipette Room 3: Amplification - Liaison MDX | Simplexa Flu A/B & RSV Direct kit MOL2651 - Reaction Mix (24) 50 µl | 2.0 mL cryovials |
|  | Simplexa Flu A/B & RSV Control Pack MOL1455 - 10 tubes, 100 µl | Nitrile gloves (powder-free) |
|  | Negative control – UTM | Filtered pipette tips, 100 or 200 µl |
|  | Sani-Cloth Bleach wipes | Gripper rack |
|  | 70% alcohol | Cryovial storage box |
|  | 5% Extran | Sharps disposal container |
|  | Universal viral transport media (UTM) | Replacement Foil wedge |

## QUALITY CONTROL

- Assay Controls; refer to  MB 9.03

  - POSC and NEGC: run daily, first run of the day

- POSC – Simplexa Flu A/B & RSV Positive Control Pack

- NEGC – UTM

  - An IC is incorporated into each reaction mixture

- QC Monitors:

| Control | Control Monitor |
| --- | --- |
| Positive Control (POSC) | Reagent failure and primer-probe integrity |
| Negative Control (NEGC) | Reagent and/or environmental contamination, cumulative effect |
| Internal Control (IC) | PCR inhibition in specimen, reagent failure or process error |

- Before reporting patient results, all controls must yield valid results. Refer to MB 9.05, Procedures G, Evaluating and Interpreting Results.

PROCEDURE A: Follow the steps in the table below to prepare specimens for testing

Testing Preparation

| Activity | ###### Step | Action | Related Doc |
| --- | --- | --- | --- |
| RIP | 1 | Call worksheet RIP; use this worksheet for sample identification throughout testing. | MB 1.01 Specimen Management |
| Sample Order Room 2 | 2 | Position samples and controls in first disc as follows: Sample; Position. POSC; Position 1. NEGC; Position 2. Patient samples; 3 – nn (max. 24 per run). | MB 3.01 Engineering Controls MB 2.01 Safe Work Practices |
| Organizing run Room 2 | 3 | Using the RIP worksheet as a layout, organize patient specimens and labels Step; Action. a; Color code worksheets and labels per run. b; Number patients on worksheet in consecutive order. c; Number corresponding patient labels according to assigned numbers on worksheet, color coded by run. d; Number each primary patient specimen according to worksheet. |  |
| Transfer NP swabs, nasal washes/aspirates | 4 | Number and label a 2.0 ml cryovial for each nasal wash/aspirate and NP swab specimen to be tested Step; Action. a; Number cap of each cryovial according to assigned number on worksheet. b; Properly label the cryovial with patient bar-coded label matching the number on the cap to the number on the label. c; Vortex specimen in original container until well mixed. d; Verify number on primary and secondary container before transfer. e; Transfer specimen into cryovial with corresponding number on cap - Only one tube can be open at a time. |  |
| Change gloves | 5 | Change gloves |  |

PROCEDURE B: Follow the steps in the table below for setting up the computer

Computer set-up

| Activity | ###### Step | Action | Related Doc |
| --- | --- | --- | --- |
| Computer Set-up Room 3 | 1 | Set up Liaison; take run specific patient labels into room 3 Step; Prompt; Action/Entry. a; ------; Turn on the Liaison MDX (ABC). b; ------; Turn on the Liaison computer. c; ------; Log on computer. d; User name; administrator. e; Password; focusIC#1. f; -----; Double-click on Integrated Cycler DX icon to open program. g; User name; Enter personal user code. h; Password; Enter personal password code. i; Assay definition; From the main screen, scan the reagent lot barcode, small data matrix located on the lower left corner of the REF card. j; Disc ID; Scan the disc barcode on the DAD to show disc layout - Used wedges are shown in black and unavailable for use - Available wedges are shown in gray Fig. 1 Figure 1. k; -----; Number available wedges according to worksheet layout. l; -----; Enter sample IDs: scan barcode ID from each label consecutively - Type drop down box: : select Unknown (default). m; -----; Enter controls according to layout - POSC – select PC-FABR from the Type drop down box - NEGC – select NTC from the Type drop down box. n; -----; Load DAD into instrument. o; -----; Select the instrument from the drop down box (lower right). p; -----; Click Run to begin processing the disc ! Once run is started, it cannot be cancelled; canceling will require reloading new samples into unused wedges. ! Users cannot be changed while running. q; -----; Recycle labels when run is complete; do not take back to room 2. |  |
| New user | 2 | To switch users: Select File: Switch Users Note: Change users before starting instrument(s) |  |
| Change PPE | 3 | Remove lab coat and change gloves before leaving area |  |

PROCEDURE C: Follow the steps in the table below for reagent handling

Reagent Handling

| Activity | ###### Step | Action | Related Doc |
| --- | --- | --- | --- |
| Room 1 | 1 | Remove one MM tube for each sample to be tested from - 20⁰ C freezer and thaw at room temperature - Use MM within 30 min | MB 9.04 Reagent and Control Prep |
|  | 2 | Remove lab coat; move the MM tubes to room 2 |  |
|  | 3 | Remove POSC from -70⁰ C freezer to thaw at room temperature |  |
| Room 2 | 4 | When thawed, gently flick POSC and MM tubes to mix; briefly centrifuge - Do not vortex - Do not refreeze | MB 9.03 Storage and Stability |
|  | 5 | Proceed to PCR set-up |  |

PROCEDURE D: Follow the steps in the table below for PCR set-up and amplification

PCR set-up and amplification

| Activity | ###### Step | Action |  | Related Doc |
| --- | --- | --- | --- | --- |
| Vortex samples Room 2 | 1 | Vortex specimen tubes prior to set-up if they have been sitting for more than 30 min |  |  |
|  | 2 | Remove DAD from package and set on disc cold block |  |  |
|  | 3 | Number wedges according to worksheet layout |  |  |
| Load MM Room 2 | 4 | Peel back the foil cover, one at a time, to expose the SAMPLE and Reaction (R) wells. ! Do not touch underside of foil to prevent cross contamination |  | Simplexa Operator's Manual IVD |
|  |  | Figure 2: | Figure 3: SAMPLE well Reagent (R) well |  |
|  | 5 | Pipette 50 µl of MM into the Reaction (R) well first before sample Tip. - To prevent aerosols and possible contamination, hold the pipette at a 30-degree angle inserting the tip under the roof of the well to reduce possible contamination - Caution: Avoid placing pipette tip at the bottom of the well to prevent possible punctures in the foil that may lead to instrument contamination. |  |
| Load samples | 6 | Pipette 50 µl of sample/control into the SAMPLE well - Caution: Pipette leakage outside of well may lead to external disc contamination when resealing wedge |  |  |
| Seal disc remove tabs | 7 | Seal the foil wedge before opening the next foil cover |  |  |
|  | 8 | After all wedges are filled, carefully remove the perforated foil tab - If foil is torn, it must be replaced with a replacement foil wedge to prevent carryover contamination |  |  |
|  | 9 | Use the disc applicator to seal the foil firmly on all wedges |  |  |
| Change gloves | 10 | Remove lab coat and change gloves |  |  |
|  | 11 | Move to room 3 |  |  |
| Room 3 | 12 | Place disc into the instrument; close lid |  |  |
| Start Run | 13 | Select test instrument from drop down box |  |  |
|  | 14 | Start run |  |  |
| Change PPE | 15 | Remove lab coat |  |  |
|  | 16 | Change gloves before leaving room 3 |  |  |
| Run time | 17 | Approximate run time: 1 h 15 min |  |  |
|  | 18 | On the screen, a progress bar indicates time to completion; refer to Fig. 4 |  |  |
|  | 19 | When run is complete, remove disc from instrument; check well volumes |  |  |
| Run completion | 20 | Place in bio-bag |  |  |
|  | 21 | If disc is completely used, discard in red biohazard container |  |  |
|  | 22 | If there are unused wedges, retain disc in a sealed bio-bag in room 2 |  |  |

Figure 4: Progress in Real-Time

Progress bar shows estimated end time

PROCEDURE E: Follow the steps in the table below for reviewing data and sample failures

Reviewing and printing Completed Runs

| Activity | Step | Action | Related doc |
| --- | --- | --- | --- |
| Analysis Complete | 1 | ### When the run is complete, the results will display on the screen; positive results appear red Figure 5: Analysis Complete |  |
| Print Review amplification curves | 2 | ### Click the Print button to print a full report of the results, Fig. 5 Step; Action. a; √ Include Ct values. b; √ Include graphs. c; Scroll through the report , reviewing comments, failures and amplification curves. d; A valid curve shows a smooth, exponential increase, Fig. 6. E; Invalid curve may be linear or a curve with data “spikes” where the curve crosses the threshold. f; Click Print. g; Export results to LIS; refer to procedure F. Figure 6: Valid and invalid amplification curves Valid Valid Invalid |  |
| Detailed analysis | 3 | ### For a detailed analysis of the completed run, click the Details button to open the Analysis Window |  |
| Analyzing Runs Detail tab | 4 | ### Click on the run Details tab to display a summary of the run, fluid checks, Ct values and any sample failures that are highlighted in yellow Figure 7: Details Screen |  |
| Analyzing Runs To view graphs by dye, click on the dye checkbox Data / Detail tabs Data tab | 5 | ### Click Data tab to Select or Deselect samples to be exported to LIS ### Select or deselect samples to view graphs (optional); reviewed in Fig. 5 ### Select or deselect samples to export to LIS - Export results to LIS; refer to procedure F Export drop down Figure 8: Data Screen Select and Deselect buttons | Refer to procedure G for evaluating QC and patient results Refer to procedure F for Exporting results to LIS |

PROCEDURE F: Follow the steps in the table below for exporting data to LIS from the analysis screen

Exporting Data to LIS

| Activity | ###### Step | Action | Related Doc |
| --- | --- | --- | --- |
|  | 1 | Confirm daily POSC and NEGC are valid before reporting patient results |  |
|  | 2 | Positive patient results: Confirm name, accession number and disc location of primary sample before releasing results |  |
| Select data | 3 | If all test results were valid upon review, select √ results to be exported on the Data tab, refer to Fig.8 | MB 9.07 Reporting and Archiving Results |
|  | 4 | Do not send invalid patient results or POSC and NEGC. Deselect by clicking on individual box(es) |  |
| Export | 5 | From the Export drop down box, select LIS and then LIS folder; click OK Figure 9: Export to LIS |  |
|  | 6 | A message that the run exported successfully will appear. Click OK |  |

PROCEDURE G: Follow the activities below for evaluating QC and patient results

QC and Patient Results

| Activity | Step | Action |  |  | Related doc |
| --- | --- | --- | --- | --- | --- |
| LIS interps | 1 | ### Patient results will be translated in LIS as Positive or Negative for Flu A, Flu B and/or RSV. If the sample is interpreted as “Invalid” by Simplexa, results will need to entered manually as Equivocal or Unresolved after review |  |  | MB 9.07 Resulting and Archiving Results |
| Review | 2 | ### Review patient and QC amplification curves before releasing results ### Check for exponential growth and data spikes ### Check for possible inhibition or low target signal ### Review “QC statement/Note” on the Segment Report for failures and error messages |  |  |  |
| Simplexa software interps |  | ### QC and patient results are interpreted by the software |  |  |  |
|  | 3 | If | Then |  |  |
|  |  | ### Detected; LIS positive | ### Flu A, Flu B and /or RSV are present in the sample |  |  |
|  |  | ### Not Detected; LIS negative | ### Flu A, Flu B and /or RSV are absent in the sample |  |  |
|  |  | ### Invalid result | ### Unable to determine the presence or absence of Flu A, Flu B and/ or RSV - Possible IC failure - Insufficient sample volume - Retest sample with new mm from the same kit or from a new kit |  |  |
|  |  | ### EC500 Data Quality Errors | ### Indicates a data quality error; weak or late amplification; Extract sample and repeat testing |  | ### Refer to Proc. I for EC500 errors |
| QC conditions not met Invalid assay | 4 | ### Failure indications will be highlighted in yellow on the Details tab, Fig. 7 Step; Action. a; ### Click the Print Preview button to review the “Data Quality message” and error code on the Segment report under QC Notes, Fig. 7. b; ### Review sample graph for amplification and Ct values. c; ### Refer to Troubleshooting Guide for additional information. d; Click the Print button to generate a report; place in molecular office review box. e; Record corrective action on QC and Equipment Failure Log. f; Record number of failed samples on Failed Run log and a brief explanation. Figure 10: Review Review curves QC Statements/Notes | Simplexa Operator's Manual IVD Appendix B: Troubleshooting MB 9.06 Troubleshooting Guide |
|  |  | If |  | Then |  |
|  | 5 | Valid assay: Controls as expected |  | - Report patient results |  |
|  |  | Invalid assay conditions: POSC/ NEGC failure |  | - Do not report patient results; invalid run - Failure caused by reagent or system failure - Review graphs for amplification - Notify technical director or designee for review - Repeat testing | Simplexa Operator's Manual IVD Appendix B: Troubleshooting |
|  |  | NEGC positive |  | - Do not report patient results; invalid run - Possible contamination of samples - Review graphs for amplification - Review the specimen handling/ preparation technique - Notify technical director or designee for review - Repeat testing | MB 9.06 Troubleshooting Guide Early IC Ct < 25: |
|  |  | Internal Control failure |  | - Possible sample inhibition - F/T sample; avoid pipetting mucus - Quick spin if large amount of mucus present - Repeat testing - IC did not amplify due to pipetting error - MM and sample reversed; placed in wrong wells - Early Ct < 25: Check DAD volumes | - Possible contamination - Bad DAD - Improper mixing - Reagent not completely thawed |
|  |  | Problem unresolved |  | - Refer to Troubleshooting section of the Operator’s Manual and MB 9.06 Troubleshooting guide - Call DiaSorin technical service, 1-800-838-4548, option 3 - Notify section technical director or designee |  |
| Problem Log | 6 | Do not report patient results until problem is resolved |  |  |  |
|  | 7 | Record problem and corrective action in the QC and Equipment Failure Log |  |  |  |

PROCEDURE H: Follow the activities below for repeat testing

Repeat Testing

| Activity | Step | Action | Related doc |
| --- | --- | --- | --- |
| Timeframe | 1 | Perform repeat testing from original specimen aliquot |  |
|  | 2 | Repeat within 3 days if stored at 2 – 8⁰ C |  |
| Vortex | 3 | Vortex the specimen tubes prior to retesting |  |
| Type of Failure | 4 | Review type of failure and any error messages containing the cause of the problem and possible solutions; refer to available troubleshooting guides for additional information Failure; Action. System error; - Read error dialog box containing software messages regarding the cause of the problem and possible solutions - Review amplification curves for exponential growth - Follow recommended actions - Repeat run including a POSC/NEGC - Contact technical service if problem does not resolve. Reagent failure; - Review proper storage conditions - Use MM within 30 min after thaw - MM subjected to 1 F/T only - Repeat testing. IC failure; - IC did not amplify due to sample inhibition - F/T sample; avoid pipetting mucus if present - Quick spin if large amount of mucus present - Repeat testing - If sample remains unresolved, call caregiver for new collection - IC did not amplify due to pipetting error - MM and sample reversed; placed in wrong wells - Repeat testing - Early Ct < 25: Check DAD volumes. Insufficient volume; - Not enough sample reached the detection chamber for testing - Check sample for mucus - F/T or quick spin to remove mucus - Repeat testing. POSC failure; - Target not detected - System/reagent failure - Repeat run including POSC and NEGC; vortex patient samples prior to testing - Flick POSC to mix before repeat testing - If POSC fails on repeat, thaw new POSC - Target and IC not detected - Review pipetting, possible sample and MM reversed - Repeat run including POSC and NEGC. NEGC; - NEGC contaminated - Repeat run including POSC and NEGC - Review patient graphs for low level contamination - Review specimen handling/processing technique - IC not detected - System/reagent failure - Possible pipetting error, sample and MM reversed - Repeat run including POSC and NEGC. Failure unresolved; - Call DiaSorin technical service - Notify section technical director or designee. | Simplexa Operator's Manual IVD Appendix B: Troubleshooting MB 9.06 Troubleshooting Guide MB 9.03 Storage and Stability DiaSorin technical service, 1-800-838-4548, option 3 Early IC Ct < 25: - Possible contamination - Bad DAD - Improper mixing - Reagent not completely thawed |

PROCEDURE I: Follow the steps in the table below to Resolve EC500 Data Quality errors/unresolved results

EC500 Data Quality Errors

| Activity | Step | Action |
| --- | --- | --- |
|  | 1 | Perform a F/T cycle on the specimen |
| Extraction | 2 | Extract 200 µl of specimen on the easyMAG using the RVP protocol - Include an extraction control (RSV or Flu) if not extracting with the RVP samples - If extracting with RVP samples, a valid RVP POSC will serve to document a successful extraction - It is not necessary to add IC since IC is incorporated into the MM |
| Dilute eluate | 3 | Add 150 µl of UTM to the patient eluate and control eluate, total volume ~200 µl |
|  | 4 | Retest according to the RIP protocol using 50 µl of the diluted eluate in UTM |
|  | 5 | If the error does not resolve, report as UNR and request a new sample for testing |

PROCEDURE J: Follow the steps in the table below for Liaison MDX instrument shutdown in room 3

Computer and Instrument Shutdown

| Activity | Step | Action |
| --- | --- | --- |
| CBA | 1 | CBA: Shut down computer and then the analyzers when all runs are completed (Computer before analyzer) |
|  | 2 | Click on the Close button or “X” out of the program |
| Shutdown menu | 3 | Click on the Start button (Windows icon) |
|  | 4 | Next to Restart, click on |
|  | 5 | Select Shutdown from the drop down menu |
| CBA | 6 | After the computer has shutdown, turn off the analyzers |
| Clean | 7 | Decontaminate work area; refer to MB 9.08 |

PROCEDURE J: Follow the steps in the table below for storing test specimens

Storage and Retention of test specimens

| Activity | Step | Action |
| --- | --- | --- |
| Storage | 1 | Store test samples in -70⁰ C freezer, shelf 4, for 3 - 6 mo. |
|  | 2 | Write date range on cryo-storage box including month, day and year |
| Disposal | 3 | Discard samples after elapsed time in red biohazard container |

PROFICIENCY TESTING

- CAP ID3, 3 shipments per year, 5 challenges each

#### METHOD PERFORMANCE

- Clinical Sensitivity/Specificity 2, 4:

- Flu A: 97.1% / 97.9%

- Flu B: 100% / 99.9%

- RSV: 98.6% / 89.5%

- Analytical Sensitivity 2, 4:

- Flu A: 0.1 – 0.05 TCID50/ml

- Flu B: 1 – 20 TCID50/ml

- RSV: 2 TCID50/ml

#### ALTERNATE METHOD

- Viral Respiratory Culture

- Sunquest Order code: VRSP

- Specimen container

- 2 NP Swabs: BBL CultureSwab with Liquid Stuart’s

- Nasal wash/aspirate (0.5 – 1 ml):  sterile screw top container

- Logistics

- Transport at RT or refrigerated

- Laboratory: Transfer 1 ml of wash/aspirate into UTM or cut 2 NP swabs into UTM

- Analytic time: 1-3 days

- Testing : Daily

## LIMITATIONS

- Results must be considered in conjunction with the clinical history, epidemiology data and other information available to the clinician evaluating the patient.

- If a novel influenza A is suspected based on clinical and epidemiological data, specimens should be collected and sent to Minnesota Dept. of Health for testing.

- This assay does not differentiate between influenza A subtypes, H1, H3 and 2009 H1.

- This assay does not differentiate between types RSV A and RSV B.

- Negative results do not rule out influenza A, influenza B or RSV.

- PCR detection of influenza A, influenza B or RSV does not distinguish between viable and non-viable organism.

- Test performance is not established for monitoring treatment for influenza A, influenza B or RSV

- False-negative results can occur if the viruses are below the level of the analytical sensitivity and if performed very early in the course of the illness.

- False-negative results may occur if the virus has genomic mutations, insertions, deletions or rearrangements

- Specimen integrity is dependent on the proper collection, transport, handling and storage of the specimen. Failure to meet set criteria can result in falsely negative results.

- When very high levels of influenza A are present with very low levels of RSV and influenza B, the signal from the RSV and FluB may not be detected due to competitive interference.

- False positive results can occur if proper handling and processing protocols are not followed.

REFERENCES

- Simplexa™ 3M™  Integrated Cycler Studio 5.0 , 3M™ Integrated Cycler Operator Manual Reference 34-8710-8239-1, PI.MOL1101.IVD_REV. F for use with IVD assays, Focus Diagnostics 2009-2012, Focus Diagnostics, Inc. Cypress, CA

- Simplexa™ Flu A/B & RSV Direct Circular PI.MOL2650.IVD, Rev. F, 18-September-2015, Focus Diagnostics, Cypress, CA 90630

- Children’s Hospitals and Clinics of MN,  Simplexa Flu A/B & RSV Direct  Verification/Validation Study, 2016

- 510(k) Substantial Equivalence Determination, Decision Summary, number K120413, July 13, 2012, Focus diagnostics, Inc., 11331 Valley view St, Cypress, CA, 90630

- Mitchell W. Woodberry, Shankar R, Cent A, Jerome KR, Kuypers J, Comparison of the Simplexa Flu A/B & RSV Direct Assay and Laboratory-Developed Real-Time PCR Assays for Detection of Respiratory Virus, JCM 2013

- Influenza viruses:  http://www.cdc.gov/flu/about/viruses/index.htm

- RSV: http://www.cdc.gov/rsv/about/index.html

| Historical Record |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Version | Written/Revised by: | Effective Date: | Summary of Revisions |
|  | 1 | P. Ackerman | 11.30.2016 | Initial Version |
|  | 2 | P. Ackerman | 3.30.17 | Added Proc. I EC500 Data Quality Errors; added early Ct < 25 to IC failures |