| Glassware Washer Operation |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Purpose | This procedure provides instructions for operating the ThermoForma Glassware Washer Model 8845 |  |  |  |  |  |  |
| Policy Statements | This procedure applies to all staff who operate the glassware washer The manufacturer’s operating manual is located in the cabinet in the dishroom |  |  |  |  |  |  |
| Materials | - Supplies - Alkaline Detergent (CST-100 Alkaline Detergent Concentrate) - Acid Neutralizer (CST-200 Neutralizer Concentrate) - Salt (Salt Crystals specified for water softeners) - Equipment - Two-level loading baskets - General purpose compartment racks |  |  |  |  |  |  |
| Quality Control | Testing for detergent residue is performed monthly to ensure proper operation of the glassware washer. See procedure MCVI 6.91 Detergent Residue Testing |  |  |  |  |  |  |
| Procedure | - The main power switch is located in the side compartment and is left in the “ON” position. - Open the washer door by rotating the door release to” O”. - Load glassware to be washed into compartment racks. - Place compartment racks in dishwasher. This dishwasher has two-level loading baskets; the top basket can be removed to accommodate taller glassware if needed. - Close the door and rotate the door release to “I” - For standard wash select program 5. - Select “Dry On” if drying is desired. - When program has completed, remove glassware from dishwasher. |  |  |  |  |  |  |
| Maintenance/ AddingSalt | The machine is equipped with a water softener. The ionic exchange resins of the water softener are regenerated using sodium chloride. The only routine maintenance required is to introduce salt into the tank when the alarm message “BEFORE STARTING ADD:SALT” appears on the display: - The salt container is located on the interior bottom of the machine. - Unscrew the tank filler cap counterclockwise to open. - Pour in salt until the container is full. - Screw the cap clockwise to close the container making sure that salt residues do not block the cap threads. The cap must be tightened completely without applying excessive pressure. If the cap resists turning check that it is not blocked by salt residue. |  |  |  |  |  |  |
| References | ThermoForma Operating and Maintenance Manual for Undercounter Laboratory Glassware Washer Models: 8837, 8845, and 8855. Manual No. 7008837 Rev. 1 |  |  |  |  |  |  |
| Training Plan/ Competency Assessment | Training Plan |  |  | Initial Competency Assessment |  |  |  |
|  | - Employee must read the procedure - Employee will observe trainer performing the procedure. - Employee will demonstrate the ability to perform procedure, record results and document corrective action after instruction by the trainer. |  |  | Direct Observation |  |  |  |
| Historical Record |  |  |  |  |  |  |  |
|  | Version | Written/Revised by: | Effective Date: |  | Summary of Revisions |  |  |
|  | 1 | Sonja Franko | 8/10/2009 |  | Initial Version |  |  |
|  | 2 | Helen Stefan | 8/27/15 |  | Reformatted and numbered for CMS, Added instructions for adding salt |  |  |
|  | Archived by: |  | Archived Date: |  |  |  |  |