RIP Positive Control

New Reagent Verification QC Worksheet

# Reagent/Kit: \_Simplexa Flu A/B & RSV Positive Control Pack\_Catalog #: \_MOL2660\_

New Lot/Shipment # (box)\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_Exp. Date \_\_\_\_\_\_\_\_\_\_Receive Date \_\_\_\_\_\_\_\_\_\_\_\_\_

New Lot/Shipment # (vial)\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_Exp. Date \_\_\_\_\_\_\_\_\_\_

Old Lot/Shipment #­­ (box)\_\_\_\_\_\_\_\_\_­­­­­­­­\_\_\_\_\_\_\_\_\_\_\_ Exp. Date \_\_\_\_\_\_\_\_\_\_Receive Date \_\_\_\_\_\_\_\_\_\_\_\_

Old Lot/Shipment #­­ (vial)\_\_\_\_\_\_\_\_\_­­­­­­­­\_\_\_\_\_\_\_\_\_\_\_ Exp. Date \_\_\_\_\_\_\_\_\_\_

If a QA failure occurs, document observation, record corrective action and notify technical specialist/designee. Contact DiaSorin Molecular Technical Service: 1.800.838.4548 (Option 3) if problem cannot be resolved.

**Expected Results**

|  |  |  |  |
| --- | --- | --- | --- |
| Flu ADetected (Ct 28 – 34) | Flu BDetected (Ct 30 – 36) | RSVDetected (Ct 28 – 33) | NEGCNot Detected (Ct 27 – 32) |

**Results**

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **No.** | **Control** | **Result (Ct value)** | **Pass/Fail** | **Comments/Tech**  |
|  Test against *in use* MM reagents  |
| 1 | POSC: Flu A  | **Old lot** | **New lot** | **P or F** |  |
|  |  |
| 2 | POSC: Flu B  | **Old lot** | **New lot** | **P or F** |  |
|  |  |
| 3 | POSC: RSV | **Old lot** | **New lot** | **P or F** |  |
|  |  |
| 4 | NEGC/IC |  | **P or F** |  |

Problem: \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Corrective Action: \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_

Reviewed by: \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_ Date: \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_