| Irradiation of Blood Products-MPLS |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Purpose | This procedure provides instruction for operation of the CIS-US IBL 437C Irradiator. |  |  |  |  |  |
| Policy Statements | - Absolute indications for irradiation of cellular blood component (RBCs, WBCs and platelets): - Documented or suspected, acquired or congenital immunodeficiency disorders. - Infants <4 months old throughout that admission period. - Therapy induced immunosuppression, aggressive chemotherapy, immunotherapy, or extensive radiation therapy. - Granulocyte concentrates/buffy coats. - Directed donation form biologic relative. - HLA-matched apheresis platelets/crossmatch compatible platelets. - Exchange transfusion in newborns. - Irradiation is not required for acellular components such as frozen/thawed plasma or cryoprecipitate. - Leukocyte reduction is NOT an acceptable alternative to irradiation. - Irradiated products may be used for any patient. Patient’s not requiring irradiation will be billed for a non-irradiated product charge. - Positive HIV status is not an indication for irradiation. - Solid organ transplant “candidate” does not require irradiation. - The ordering provider shall indicate on the blood product order the need for irradiated products. All subsequent transfusions must be irradiated until the physician or pathologist determines irradiation is unnecessary. - Information regarding the need for irradiation will be reflected tin the patient’s BAD file with PROBLEM comments UIRD or IR4 and the ATTRIBUTE of IRD. Refer to TS 5.3 Making change to a blood administration recorded. - IR4 may be removed when a patient >4 months old is discharged or on re-admission. - Gamma irradiation of 25 Gy (2500cGy) will be targeted to the central portion of the container to render the T-lymphocyte in the blood component incapable of replication thus prevent GVHD. A minimum dose of 15GY (1500cGY) shall be delivered to all other parts of the component. - The expiration date of red cells products must be corrected to 28 days from the date of irradiation if the original outdate >28 days. - Irradiate just prior to transfusion if possible due to the increased levels of potassium in stored irradiated blood. - RBCs that have irradiated >72 hours should not be used for large transfusions (>20 mls per Kg) - Irradiated products MAY NOT be returned to the Blood Center. - In an emergency do not withhold blood products due to lack of irradiation. Notify the attending practitioner. - The Blood Bank technologist should consult a pathologist if it is unclear why irradiated products have been ordered for a patient. |  |  |  |  |  |
| Related Documents | - TSf 10.8 Irradiation Logsheet - TS 17.27 Using Radsure Irradiation Indicators. |  |  |  |  |  |
| Materials | Equipment |  |  | Supplies |  |  |
|  | - IBL 437C irradiator |  |  | - RAD-SURE® irradiation indicators 25 Gy, FDA registration #BK920035 (stored in small BB refrigerator) |  |  |
| Quality Control | TS 17.21 IBL 437C Blood Irradiator function verification and maintenance |  |  |  |  |  |
| Procedure |  |  |  |  |  |  |
|  | Step | Action |  |  |  |  |
|  | 1 | Open the chamber door and remove the canister by aligning the large opening on the canister with the blue lines on the shelf. Always turn the canister clockwise. |  |  |  |  |
|  | 2 | Place a RAD-SURE® Indicator on the selected unit(s). |  |  |  |  |
|  | 3 | Document the required information on the irradiation logsheet. |  |  |  |  |
|  | 4 | Place the product(s) to be irradiated in the canister and replace the top.The number of units that can be irradiated in one batch is limited to: - 4 packed red cells - 8-10 random platelets - 1-2 platelet pheresis - 2 syringes - Refrigerated and room temperature storage products MAY NOT be irradiated in the same batch. |  |  |  |  |
|  | 5 | Place the canister in the machine by aligning the propeller, then sliding the canister into place by lining up the large opening on the canister with the blue lines on the shelf. |  |  |  |  |
|  | 6 | Rotate the canister a quarter turn clockwise and lock into place by aligning the red indicator line on the outside wall of the canister with the red indicator line on the shelf. |  |  |  |  |
|  | 7 | Close the door. The NO CANISTER indicator light (1) will go off and the door will mechanically lock. |  |  |  |  |
|  | 8 | Verify or adjust the irradiator timer as to the posted 2500 cGy time. |  |  |  |  |
|  | 9 | Press the CYCLE START switch (4) to begin the irradiation cycle. The drum will start rotating. |  |  |  |  |
|  | 10 | Verify that the CANISTER ROTATION indicator light flashes to signal the exposure of the canister to the source and the start of the rotation of the canister. Document this on the irradiation logsheet. |  |  |  |  |
|  | 11 | Verify that the machine is operating properly by checking the indicator lights. The CYCLE START light will go off as the process begins. |  |  |  |  |
|  | 12 | The irradiation cycle is automatically terminated when the set time has elapsed. The IRRADIATION light will go off when the blood product(s) is no longer exposed to the source. |  |  |  |  |
|  | 13 | When the canister is at the door, the door will mechanically unlock. The CANISTER ROTATION indicator light will go off. |  |  |  |  |
|  | 14 | Open the door to the chamber. Rotate the canister a quarter turn clockwise to release it from its holder. Align the large opening in the canister with the blue lines and remove the canister from the chamber. |  |  |  |  |
|  | 15 | Remove products from canister and observe RAD-SURE indicator. The window must be black and opaque, covering up the word "NOT". If the word “NOT” is clearly visible, notify the section technical specialist or lead. Note: A blood product may receive 2 cycles of irradiation for a maximum exposure of 5000cGY if the Radsure indicator was omitted on the first cycle or if the indicator/blood product was inadvertently placed in a position outside of the dose mapping field. |  |  |  |  |
|  | 16 | Complete the remaining required information on the irradiation logsheet. |  |  |  |  |
|  | 17 | Remove the blue “Not Irradiated” tag from the unit if attached. |  |  |  |  |
|  | 18 | Update the component type in the computer in function BCP. Reminder: The expiration date of red cells products must be corrected to 28 days from the date of irradiation if the original outdate >28 days. |  |  |  |  |
|  | 19 | Apply product over-label. |  |  |  |  |
| References | - AABB Technical Manual, current edition - AABB Standards for Blood Banks and Transfusion Services, current edition - IBL 437C Irradiator for Blood Products Operator’s Manual, Compagnie ORIS Industries, France 1987. |  |  |  |  |  |
| Approval Workflow | Transfusion Service/Medical Director |  |  |  |  |  |
| Historical Record | Version |  | Written/Revised by: |  | Effective Date: | Summary of Revisions |
|  | 1 |  | J. Jaimez; |  | 8/13/93 | Initial Version |
|  | 2 |  | D. Hansen |  | 8/5/96 | New format |
|  | 3 |  | J. Wenzel |  | 11/17/1998 |  |
|  | 4 |  | D. Hansen |  | 9/1999 |  |
|  | 5 |  | J. Wenzel |  | 9/2001 |  |
|  | 6 |  | J. Wenzel |  | 5/01/2003 |  |
|  | 7 |  | J. Wenzel |  | 6/30/06 |  |
|  | 8 |  | J. Wenzel |  | 10/20/08 | New format |
|  | 9 |  | S. Cassidy |  | 04/10/2012 | CMS format |
|  | 10 |  | S. Cassidy |  | 04/09/2018 | Added statement on how long a RBC unit can be irradiated for large transfusions |